## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 12 January 2017 – 28 November 2017

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 8 February 2017 – 12 February 2017

Number of original Hex files: 74 	Number of casts processed: 68

INSTRUMENT SUMMARY

Six Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1449, 1451, 1452 and 1513. Each CTD had an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

Four different CTDs were used. The calibration information could not be confirmed. The CTDs for 1 had been updated in 2016. The others have calibration coefficients from 2014.

To assist in analysis, event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project based on the date. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than one cast on one day, as frequently occurred.

The soak period for the CTD varied greatly from about 25s to 2 minutes, but was usually about 1minute. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, but it is believed that local time was used. All times were converted to UTC by adding 7 or 8 hours depending on whether local time was PDT or PST.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

74 hex and con configuration files and 72 hdr files were delivered.

The file names are non-standard and contain the opening N followed by the date in YYMMDD format and 2 digits to indicate the sequential casts on a single day. There are no event numbers as this is a time-series rather than a cruise.

The non-standard names are in the usual format for this project and will be kept.

- Conversion of Raw Data

- The raw data were converted by matching the configuration file to the input file name.

- There was one con file with no matching hex file.

- There was one hex file with no matching con file, but since it was the 3rd cast of the day the con file from the previous cast was used to create the CNV file.

- (Note A few files were renamed later because they had been given the wrong month in the raw file names.)

- This produced 74 CNV files.

- A spreadsheet “NanooseCtdStations2017.xls” was provided with positions for the casts but there were only 57 stations. For 12 of the 18 casts missing positions there were repeat occupations on a single day with only 1 position given. The 1 position was used for the missing positions on the same day.

- One of the other 5 casts (n1711601) with no positions had surface data only and it was full of spikes, so it won’t be processed further.

- For the other 4 casts a standard position was: 49° 20’N, 124° 5’ W.

- For events 1025, 1103, 1108 and 1116 a note should be placed in the header to state that the positions are estimates.

- The spreadsheet provided was adjusted by adding the missing positions and saved as 2016-99-merge-hdr.csv. Some editing was required:

- The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this.

- A column LOC:Station was added with every row having “Nanoose” entered.

- A column of event numbers was adding to the CSV file using MMDD part of the file name as event number.

- The soak period ranged from 25s to 2 minutes, with the most common wait was for 1 minute. A 2-minute soak is generally advised for the best salinity data, but in most cases the conductivity seems to have reached equilibrium after 1 minute.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom but the CTD did not always get within 15m, in which case the bottom value did not get entered into the IOS Headers.

- WILDEDIT

It did not appear that this step was needed as no spikes were found, but some pressure spikes were later found in file n17071901.cnv. So WILDEDIT was run on that file and other steps were then repeated.

- FILTER

As usual for SBE 19 CTDs the conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.  Tests were run on a few casts to ensure this was the best choice.

The pressure was not filtered since it is quite smooth already. In previous years tests showed that filtering made things worse in sections where pressure was constant.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Temperature was advanced by 0.73s relative to pressure since that has worked well for data from this project in previous years. Alignment can be fine-tuned later using SHIFT if further adjustment is found useful.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.
The program MERGE CSV FILES TO HEADERS was then used to add positions, station name, event numbers and project name to the headers.

8. Checking Headers

Track Plot was used to ensure positions look normal. There are no obvious outliers though, as usual, there is quite a scatter in locations.

Header Check was run and the speed check turned up some files that were misnamed:

- Two casts that were given file names based on March 4 were actually collected on April 4 according to the SBE headers.

- One cast collected on August 29th as named as being from Sept. 29.

The file names were changed to match the dates and affected file lists and spreadsheets were also updated.

Header Check was rerun and the only potential problems noted are some slightly negative pressures and unbelievable salinity. The pressure problems appear to only affect the first record.

A cross-reference list was produced. The times appear to be in local time. The times of casts are not consistent enough to notice if clocks might have changed, but there is no obvious trend towards later times in summer, so it is likely that PST and PDT were used. So ADD TIME CHANNEL was used to add 8 hours to casts before March 12th and from November 5 to the end of the year. ADD TIME CHANNEL was used to add 7 hour to all other casts.

Surface check was run and values ranged from -0.2 to +8db, except for 4 casts which started very deep at ~300db. Those casts contain upcast data.

The average surface pressure for casts that started in the top 10db was 1.2db.

Profile plots were examined and problems were found with some casts:

- n17031701 - shallow – have another cast from that day. NO FURTHER PROCESSING

- n17040401 – odd profile and shallow – have 2 other casts from that day. NO FURTHER PROCESSING

- n17060101 – may need editing near surface

- n17061302 – may need editing near surface

- n17062701 – only surface data  NO FURTHER PROCESSING

- n17062702 – may need editing of spikes

- n17071901 – may editing of spikes

- n17090602 – only deep data and there are 2 other casts from that day. NO FURTHER PROCESSING

- n17102002 – upcast only – put through REVERSE

- n17102003 - upcast only – put through REVERSE

- n17102402 - upcast only – put through REVERSE

The 4 cases that are not to be processed further were dropped from the file list.

The cases of “may need editing” will be reexamined after DELETE.

- The altimetry header information from each file was exported to a spreadsheet. 70% of the casts got within 15m of the bottom with none closer than 4.5m of bottom.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features. Three casts with features that might benefit were chosen to test. The results were not consistent with only slight improvements with varying settings. Since 4 different CTDs were used, it is likely that the setting would not work well for all casts and the unstable features may best be handled through editing.

SHIFT was not run.

10. CLIP and REVERSE

CLIP was used to remove the first 50 scans; by eliminating the early soak period editing will be simpler. Plots were examined to ensure no good data had been removed. Two of the upcast files were put through CLIP again to remove only 40 scans so bottom data were not lost.

3 casts had to be put through REVERSE. After that step the CLIP files were copied to CLIP2 and the REVERSE files were copied to CLIP2 so there are a complete set of files named CLIP2.

11. DELETE

Then DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were warnings for 3 events:

- #317 - All warnings pertain to jumps in pressure present in the original files.

- # 719 – Most warnings look like those of #317 except that there is a patch of clearly bad pressure and conductivity data at about 192db.  This is found in the original file but got spread by the filter. WILDEDIT was run on this file and all other steps then repeated. DELETE still issued many warnings but they do represent real jumps and the data conductivity and salinity around 192db looks much improved.

- # 906 – There is a large jump after pressure =332db in the original file with just 1 record after that at 374db. The bottom values compare reasonably well with another cast on the same day but it is hard to judge quality with just 1 record and it is not very useful given no intermediate values so it is probably best to remove the last record.

Using WILDEDIT, adjusting the filter applied earlier or using the pressure filter within DELETE won’t help with the pressure jumps. DELETE appears to have work properly.

12. CTDEDIT

The DEL files were copied to EDT.

Plots were made to identify casts that may need editing.

DEL files were opened in CTDEDIT. 31 of the casts were edited, mostly to smooth cases where unstable features appeared to be due to minor misalignments of T and C in the presence of temperature gradients. A few records were removed from the top and bottom of some casts.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen. File n17101901 looked bad as it was well outside the climatology and looked very different from another cast a few minutes later using a different CTD.  The fact that the repeat was done so quickly using a different instrument probably indicates that the CTD operator could tell something was wrong. That particular CTD was not used again for 2017 Nanoose sampling.

T-S plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – As mentioned in the previous section file n17101901 was well outside the climatology for temperature and salinity and very different from another cast run minutes later. The only other excursions from the local climatology were:

- Temperatures were a little high in the top 40m for 2 casts in May and around 15m for one cast in July.

- Salinity was slightly low around 40m for one January cast and near the surface down to about 30 for 2 casts in May.

These excursions do not look like evidence of calibration problems.

Repeat casts – Casts were all close in space, and on a few occasions there were 2 or more casts on one day but they are not necessarily very close together. A few casts that were close in space and time were examined.

- As mentioned earlier cast n17101901 looked very different from n17101902 and was well outside the climatology throughout the profile, so it will not be archived.

- For casts n17031702 and n17031703 which were only a few minutes apart the temperature and salinity differed by about 0.02C° and 0.01psu along lines of constant σt but the differences did vary quite a lot. The same CTD was used for both.

- For casts n17060601 and n17060602 which were about 35 minutes apart the temperature and salinity differed by about 0.05C° and 0.02psu along lines of constant σt.. The same CTD was used for both.

- For casts n17102401 and n17102402 which used the same CTD, there is a difference in temperature of about 0.05C° and in salinity of about 0.03psu along lines of constant σt above 300m but they look close below that.

- For casts n17071901 and n17071902 which used different CTDs and occurred about 25 minutes apart, the differences are about 0.05psu along lines of constant σt with salinity consistently higher for one cast while temperature differences were larger but varied in sign.

- For casts n17080901 and n17080902 in an area with smooth gradients the differences were about 0.015C° and 0.004psu along lines of constant σt,

- For casts n17090601 and n17090603 in an area with smooth gradients the differences were about 0.010C° and 0.003psu along lines of constant σt,.

These differences are no doubt due to a combination of the instruments’ resolution limits and real change. The late summer results, based on simpler profiles, are probably the best measure of the instruments’ repeatability.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The CTDs are serviced annually.

- For events 1025, 1103, 1108 and 1116 the following note was added to the header note:

WARNING: There was no position information available for this cast, so an

estimated position was entered.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing and header check were produced. No errors were found.

Particulars

n17031701 - shallow – have another cast from that day. NOT PREPARED FOR ARCHIVE.

n17030401 –misnamed – should be 17040401 – odd profile and shallow – have 2 other casts from that day. NOT PREPARED FOR ARCHIVE.

n17030402 –misnamed – should be 17040402

n17062701 – only surface data   NOT PREPARED FOR ARCHIVE.

n17090602 – only deep data and there are 2 other casts from that day. NOT PREPARED FOR ARCHIVE.

n17092901 –misnamed – should be 17082901

n17101901 – Deleted as outside climatology and very different from another cast on same day which was within the climatology. Last use of that CTD during this projecct. NOT TO BE ARCHIVED.

n17102002 – upcast only – put through REVERSE

n17102003 - upcast only – put through REVERSE – appears to be the same cast as n17102002 but with fewer records and looks slightly less stable. Not to be archived.

n17102402 - upcast only – put through REVERSE

n17102501 – No position available – used standard position: 49° 20’N, 124° 5’ W

n17111601 – Surface data only, no positions, lots of spikes. NOT PROCESSED.

n17111602 – No position available – used standard position: 49° 20’N, 124° 5’ W

n17110301 – No position available – used standard position: 49° 20’N, 124° 5’ W

n17110801 – No position available – used standard position: 49° 20’N, 124° 5’ W

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	12 January 2017	    End:	28 November 2017

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

NOTE: Four SEACATs were used for this cruise.

CTDs

| CTD# | Make | Model | Serial# | Altimeter | Calibration data – Pressure, Temperature, Conductivity |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 19 | 1449 | 1094 | October 2014 |
| 2 | SEABIRD | 19 | 1451 | 945 | May 2016 |
| 3 | SEABIRD | 19 | 1452 | 1093 | July 2014 |
| 4 | SEABIRD | 19 | 1513 | 1096 | July 2014 |

The calibration information used in conversion can be found in the headers of the CTD files and a summary of which CTDs were used for which casts can be found in document 2017-99-ctd-final.doc which is in the archive folder Cruise_Data_Processing\2017-99\Archive\DOC.