## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | ### Updated channel names & formats in TOB & loop files. GG & SH |
| ### 18 July 2023 | ### Corrections flags/comments for CHE files:19, 20, 36, 37, 43, 53, 60, 64, 69, 72, 76, 79, 82. |
| ### 26 March 2023 | ### Added HPLC S.H. |
| 12 May 2021 | Added DOC/TOC/TDN/TN to 10 CHE files. S. H. Corrected SBE:DO format & channel name for Conductivity in all CHE files. |
| 7Aug2018 | Corrections to CHL precision statement in CHE files & new QF file for CHL |

## PROCESSING NOTES

Cruise: 2018-01

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: Sir Wilfrid Laurier

Date: 18 February 2018 – 8 March 2018

Processed by: Germaine Gatien

Date of Processing: 13 April 201 –  2018

Number of original HEX files:  	48 (+1 test file)	Number of CTD files:

Number of bottle files: 	48			Number of bottle casts processed: 45

Number of original TSG files: 6			Number of processed TSG files:	 2

# INSTRUMENT SUMMARY

Two SeaBird Model SBE911+CTDs were used for this cruise, #1222 and 443.

CTD #1222 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#3038) on the primary pump, a SeaPoint Fluorometer (#3642) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#62355).

CTD #0443 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the primary pump, a SeaPoint Fluorometer (#3641) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#62355).

A thermosalinograph (Seacat 21 S/N 3274) was mounted with a Wet Labs WETstar fluorometer (S/N SCF3654).

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial #60572.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order with detailed comments on problems. Notes from the Chief scientist were very helpful as was the log of loop sampling.

The poor quality of sample labels led to many errors in recording sample numbers.

For both CTD systems the secondary salinity was selected for archiving. For CTD #1222 the primary salinity was closest to the bottles but the secondary channels looked more stable in T-S space. For CTD #0443 the secondary salinity was closest to the bottle salinity and the secondary channels look much more stable in T-S space.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore. When the data are divided geographically it is seen that for CHL values <2ug/L, the CTD fluorescence is fairly close to the CHL when nearer to shore but much higher well offshore.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 300db

±0.2 mL/L from 300db to 500db

±0.05 mL/L from 500db to 2000db

Low by up to 0.03mL/L below 2000db

There were 6 Thermosalinograph files, but only 2 were prepared for the archive. The first 2 files had uncertainties in both time and positions with neither reliable enough to attempt to patch data from other ship records. For file #3 there were excellent time/position matches with 4 CTD casts, so the clock and NMEA data appear to be fine, though this sort of check would not detect a small time error. The 4th and 5th files were short, lasting only 2 minutes and 13 minutes respectively, as changes were made to time acquisition; the ship was stopped so the files do not include useful data; they were not prepared for the archive. The 6th file was large and the data look good. There were only minor spikes that mostly contained only one point and they were well separated; they were removed easily using WILDEDIT.

The TSG fluorometer readings were slightly lower than those from the CTD fluorometer. TSG fluorescence read significantly lower than chlorophyll samples from the loop and rosette, which is typical of offshore sampling.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Sampling notes from the chief scientist noted a number of issues. Based on these notes a few corrections were made in the headers of the files before conversion.

CTD #1222 was used for casts 1 to 33.

CTD #0443 was used for casts 36 to 100.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

- Since it was last serviced in late 2015 CTD #1222 and the DO and conductivity sensors mounted on it had been used for one other cruise.

- CTD 0443 and the conductivity and DO sensors mounted on it had not been used since the latest factory servicing in August 2017.

The configuration changed after cast #33 when the CTD was switched.

Two configuration files were created and parameters were checked. No errors were found** but factor E for the DO sensor on CTD 0443 should be tested as the default value may not be appropriate.

2018-01-ctd1.xmlcon – for casts 1-33

2018-01-ctd1.xmlcon – for casts 36 to 100

** An error in the pressure offset for CTD #1222 was found later and is discussed in section12.

##### Hysteresis Tests

CTD #1222 was not tested as there was only 2 bottles from deeper than 1500m. Most of the casts using this CTD were too shallow for hysteresis to be significant.

For casts 37, 53, 59, 72, 79 which used CTD #0443 and included deep sampling, tests were done using factor E in the DO calibration set to 0.0036, 0.0037 and 0.0038. The best choice looked to be 0.0036, the default value. Judging by the changes caused by increasing the E value, it is clear that reducing it will have a poor effect but a small increase might help. A value of 0.00362 did improve the results and it is clear that a further very small increase would likely be ideal. So a choice of E=0.0363 was made.

Small changes in H1 or H3 were tested to see if they were more or less useful than adjusting E. As usual changing H factors was not helpful.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-01-ctd1.xmlcon or 2018-01-ctd2.xmlcon. Depth was included.

Cast #1 had no sampling so will not be processed further.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Outliers were found in salinity in 4 casts and CTDEDIT was used to clean salinity as follows:

- Cast #19 - Salinity:T0:C0 at about 10db.

- Cast #20 - Salinity:T0:C0 at about 10db.

- Cast #62 - Salinity:T0:C0 at about 75db.

- Cast #64 - Salinity:T0:C0 at about 75db and Salinity:T1:C1 at about 180db.

The output files from CTDEDIT were copied to the BOT files.

A preliminary header check and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. An adjustment had to be made to the sample numbers for event #60 because sample number 295 was used twice – during events 59 and 60; #9295 was used for cast 60.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

File #60 was reordered on sample order.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-01-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-01chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-01chl.csv. Loop data were saved in a separate file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-01oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2018-01oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2018-01SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 13 to 30 days after collection.

The files were simplified and saved as 2018-01sal.csv. File 2018-01sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-01_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-01-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2018-01).xls. Values given as < were changed to 0 and those given as – were replaced with pad values; the comments that will go into the header will explain that 0 means below minimum detectable level which is 0.1nM. The file was then saved as 2018-01DMS.csv and converted to individual DMS files. There was a separate report on analysis techniques and problems.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. This step is important to avoid missing some samples in the creation of CHE files.

Problems found:

- Event #18/19 - Salinity sample given as 34, but comment says it was labelled 35. There was no sample on rosette sheets for either 34 or 35, but there is an entry for sample 33 and no such sample found. So this sample is presumed to be the missing #33.

- Event #19 - DMS sample numbers have some errors. There should be samples 35, 37, 38 & 39 but the file has 35, 37, 36 & 39. They are all 0 values. So 36 has no sample indicated and 38 does, so 36 was renamed 38.

- Event 37 -Salinity sample #158 is missing and there is a sample 150 that is not on the rosette sheet. The CTD salinity for #158 looks close to the sample so it looks like this was mislabelled.

- Event 44/53 – Salinity sample 225 listed as from cast 53, should be 44.

- Event #56 – CHL sample #262 lost between raw and QF page

- Event #59/60 – Duplicate sample #s 295. Use sample 9295 for cast 60 and 295 for cast 59.

- Event #60 – 4 lines were removed from the SAMAVG file since there was no sampling.

- Problems with the nutrient spreadsheet starting with event 46 were fixed by the analyst.

- Event #69 - There were 2 DMS samples called #383 and none labelled #381.  The Niskin number makes it clear which one should be renamed #381.

- Event #81/82- Nutrient sample #501 from event 82 is missing. There is a sample 500 from cast 81 but that is not on the rosette sheet. Could 500 really be 501? The CHL sample labels were amended in pencil to read 500 but entered in the spreadsheet as 501. The CHL analyst changed 500 to 501 and added a flag 2.

- Cast #88 – There is an extra salinity sample not mentioned on the sampling log but the value looks reasonable.

There are loop data in the salinity, nutrient and chlorophyll spreadsheets, so those data were moved to file 2018-01-loop-data.xlsx.

The chief scientist provided a log of loop sampling. There were loops taken while underway and others were taken at the end of CTD casts while the 5m rosette was being fired.

Times were added to the loop file based on the end time of CTD casts or the loop time if there was no CTD cast at the time.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The data were studied in 2 groups, one for each CTD system.

For CTD #1222 there are only 17 samples, of which only 9 are below the surface. Of those, 2 are somewhat out of line and were excluded. The primary salinity was found to be high by 0.0019 and the secondary high by 0.0038; the standard deviation in both was 0.0005. The primary has slightly less pressure dependence in the fit against pressure. There are insufficient data to make a fit against time worth studying.

The largest outlier is only out of line by ~0.004. There is no obvious source of error in the CTD data or in other samples from the same bottle. Given the scatter in the fit, no flag is justified for this outlier.

For CTD #0443 there are many deep samples. First, bottles were excluded if the standard deviation in the CTD data was >0.0008 and if pressure was <100db; one other outlier had been flagged 4 and was excluded.  The average differences showed the primary salinity to be high by 0.0018 and the secondary by 0.0011. There seemed to be considerable pressure dependence in the secondary salinity due to bottles between 100 and 500db from a single cast. When bottles with pressure <500db were excluded the opposite picture emerged with the primary salinity having very larege pressure dependence. This is unusual and may be related to the observations of unusual temperature differences between the two sensors as noted in section 10. The temperature differences appear to have both temporal and pressure dependence but that is not completely clear.

The only bottles between 100 and 500db come from a single cast. A close examination of the downcast data at those depths shows some reversals in temperature and a lot of noise in both temperature channels. The primary temperature has more spikes but the secondary bottles from this cast look more out of line in the comparison, though if there is some slight inefficiency in bottle flushing the secondary would be the more believable. Given the complications, it is wise to drop these bottles from the comparison.

The pressure dependence from 800m down is much higher in the primary than in the secondary and much higher than expected. Combined with the observations reported in section 10 it appears that the primary temperature sensor may be performing poorly. The CTD technician was informed. The secondary sensor has a little more time-dependence than the primary but it is very flat with pressure below 500db, so should be chosen for archiving, if possible.

The 23 bottles fired at 2000m were studied to compare the noise in the bottle and CTD data. A smaller window was chosen for the CTD data selected from the SAM files to avoid the overlap in data selected by Compare when many bottles are fired in a short time. Some observations:

- The CTD data show little variation with a standard deviation of 0.0001

- The bottle data show much more variation with a standard deviation of 0.0025.

- The variability in the differences comes almost totally from the bottle data variability.

- The CTD moved vertically through a 5m while bottles fired; the range of individual (i.e. not averaged) salinity values recorded through that time ranged from 34.5817 to 34.5831.

- The precision study shows the Sp value of 0.0018.

The variations in differences during this stop at 2000m are larger than the variations from all the other bottles below 500m. Examination of CTD data during the stop shows that there was an adequate wait for values to settle down and CTD values were not varying much. The variability is due to bottle variations which could be explained by analysis error variability, but that does not explain why the average difference is higher than found from other bottles. The explanation for that may be incomplete flushing of Niskin bottles. Given that these bottles were fired at the bottom of the cast, the bottles would tend to have salinity that is slightly lower than in situ values.

A few more outliers were removed based on differences. The primary salinity was then found to be high by 0.0017 and the secondary by 0.0007 with a standard deviation of ~0.002 for both. The standard deviation is similar to the pooled standard deviation reported by the analyst in the precision study.

For details see file 2018-01-sal-cast61-study.xls.

There were no major outliers that were not already flagged.

For full details for the COMPARE run see file 2018-01-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The data were examined in 4 groups – CTD 1222, CTD 0443 all, CTD 0443 offshore and CTD 0443 inshore.

CTD1222: There is a lot of scatter, mostly coming from the cast at P2. The P2 bottles were all flagged due to analysis problems. When only the P1, P4 and P8 casts are included and a few outliers are identified associated with complex reversals in the DO profiles or cases with the bottle reading lower than the CTD, the fit was:

CTD DO Corrected = CTD DO * 1.0327 + 0.0627 (R2 = 0.97)

The first 2 casts were fairly shallow with quiet descent rates while the latter 2 were deeper and had noisy descent rates.

CTD 0443:ALL - A plot of all casts together has huge scatter. Choosing outliers is very subjective.

CTD 0443: Offshore – The near-surface data are mostly out of line with the trendline from other bottles. An examination of the profiles shows that DO was quite well-mixed in the top 75m. This will reduce any errors in CTD DO due to slow response but not the calibration errors.

A fit using data from below 60m and excluding 2 outliers based on residuals is:

CTD DO Corrected = CTD DO * 1.0283 + 0.0491 (R2 = 0.93)		(Fit 1)

A fit using only surface bottles forced through the origin is:

CTD DO Corrected = CTD DO * 1.0285                (R2 = 0.05)		(Fit 2)

If we force the offset to be the same as for a fit using only the offshore casts the fit is:

CTD DO Corrected = CTD DO * 1.0211 + 0.0491 (R2 = 0.10)		(Fit 3)

The scatter in the surface-only fits makes them unreliable, but gives at least a hint of how much of the correction derived from deep-water comparisons is due to calibration drift. The difference between applying Fit 1 and Fit 3 is 0.007*DO in mL/L. So the general fit overcorrects by about 0.05mL/L at the surface. Assuming that the fit through so much scatter is reliable, we could conclude that the correction due to slow response is about 0.05mL/L and the correction due to drift is about 0.20mL/L, or 80% of our observed error is due to calibration drift.

This study leads to 2 conclusions:

- Corrections in well-mixed layers may be too high, but by how much is in some doubt.

- Where surface waters are very well mixed, excluding surface bottles will lead to a better fit to other levels.

- Doing a DO-dependent fit may improve the fits somewhat for offshore waters where the well-mixed surface layers are the only areas with high DO values. However, the choice of DO value for a cut-off will miss some well-mixed waters or extend into higher gradient waters as the depth varies considerably. Whether that compromise is justified would depend on how good the near-surface comparison looks. For these data there is far too much scatter to consider such a step.

CTD 0443: Inshore – The fit when cases where bottle DO > CTD DO are excluded plus a few outliers based on residuals is:

CTD DO Corrected = CTD DO * 1.0152 + 0.0696 (R2 = 0.62)

The fit is poor and the slope of the fit much lower than for the offshore. This is likely due to poor flushing of Niskin bottles so that the CTD DO is being compared with samples from deeper in the water column. Another factor may be the dominance of well-mixed DO profiles reducing the error due to poor response.

The choice of fit to apply is subjective. The near-shore fits tend to be unreliable so we rely on offshore fits. Those likely overcorrect the near-surface DO and under-correct in higher gradients. A pressure-dependent fit is not practical as it would have to vary from cast to cast even at the same site and there are usually few bottles in the high-gradient zones. A general fit for all levels is a reasonable compromise.

Outliers:

Most outliers were from the near-surface area:

Event 13 –Niskin #1out of line but not enough to justify a flag given it was fired close to bottom.

Event 14 – All bottles were flagged by analyst, and all but the 10m bottle were significant outliers, while the 10m bottle was somewhat out of line. The 3 flags were changed to 4 by the analyst.

Event 20 – One bottle was excluded as an outlier but the CTD data look noisy. 3 other bottles were slightly out of line but came from a section of rapid reversals in DO. No flags are justified as the mismatches are likely due to either slow response of the CTD or mismatches in depths of CTD and bottles.

Event 32 – The only outlier was only slightly out of line and the CTD data were noisy.

Events 37-100 – The outliers were either from the surface layer or only slightly out of line so that flags are not justified.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The data for each CTD were studied separately.

CTD #1222:  Fluorometer #3642

The ratio FL/CHL versus Extracted CHL looks typical of these sensors. It starts at about 3.5 for very low CHL and drops rapidly to about 1as CHL values rise. The ratio continues to drop slowly but is close to 1 for 0.6< CHL<1.1ug/L. There are no CHL>1.3ug/L.

CTD #0443: Fluorometer #3641

The ratio FL/CHL versus Extracted CHL looks typical of these sensors. It starts at about 10 for extremely low CHL and drops rapidly to about 1as CHL values rise. The ratio continues to drop slowly but is close to 1 for 0.6< CHL<1.1ug/L. It drops to about 0.75 in the range 2< CHL<9ug/L.

For full details of the comparison see file 2018-01-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-01-ctd1.xmlcon or 2018-01-ctd2.xmlcon.

The hysteresis and Tau functions were selected. Depth was included in the conversion.

Separate cast lists were prepared for each CTD and for casts with and without PAR sensor.

A few casts were examined and all expected channels are present.

- The transmissivity looks a little odd with higher values on the upcast than the downcast.

- PAR, Fluorescence, Dissolved Oxygen and Altimetry have normal profiles.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined from each of the CTD systems used during this cruise; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. One cast for each CTD system was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts for each CTD to examine differences between sensor pairs. The highlighted entries are for CTD1222.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0026 | 1000 | -0.0005 | +0.0001 | +0.0012 | High, Mod |
|  | 1900 | -0.0009 | +0.0001 | +0.0020 | “ |
| 2018-01-0027 | 1000 | -0.0002 | +0.0001 | +0.0013 | High, Mod |
|  | 1900 | -0.0005 | +0.0001 | +0.0020 | High, Noisy |
| 2018-01-0032 | 1000 | -0.0002 | +0.0001 | +0.0011 | High, Noisy |
|  | 1900 | -0.0006 | +0.0001 | +0.0018 | High, VNoisy |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
|  | 1900 | -0.0007 | -0.0001 | -0.0012 | High, VNoisy |
|  | 2900 | -0.0012 | -0.0001 | -0.0001 | “ |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
|  | 1900 | -0.0007 | -0.0001 | -0.0010 | “ |
|  | 2900 | -0.0013 | -0.0001 | -0.0000 | “ |
|  | 3900 | -0.0019 | -0.0001 | +0.0011 | “ |
| 2018-01-0072 | 1000 | -0.0005 | -0.0002 | -0.0023 | High, XNoisy |
|  | 1900 | -0.0008 | -0.0001 | -0.0012 | “ |
|  | 2900 | -0.0015 | -0.0001 | +0.0000 | “ |
|  | 3900 | -0.0020 | +0.0000 | +0.0010 | “ |

- CTD 1222: All differences are small with some pressure dependence in the salinity and temperature. This CTD was not used for any casts deeper than 2000m.

- CTD 0443: The differences are not large for any sensor but the temperature and salinity pressure dependence below 2000m is larger than usual.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. There was an error in station name for cast #95. That was corrected.

The header check was run and the only problem found was negative DO during cast #33; this was at the bottom of the cast at the point the cast was aborted.

The cruise track was plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Water depths were compared with the log book entries. A few changes were made to the headers:

- The depth entries were adjusted for casts #13, 14, 18, 43, 53, 88 and 90.

- The depth for #19 was changed to 1315 based on Chief Scientist recommendation.

- For cast #24 it looks like the water depth and bottom depth were reversed, so 1317 was entered for the water depth in the file.

- For cast #62 the file depth entry is better than the log, but too high. Based on previous years’ observations plus the altimetry it was changed to 3990m.

- The altimetry headers were removed from casts 27 and 61 as they came from spikes.

- For the bottle files, the same changes were made as above.

A surface check was run and shows an average surface pressure for the cruise was 3.2db. The deepest values were from casts from the offshore areas where seas were likely rough.

**The pressure offset for sensor #1222 was entered incorrectly at the time of conversion. During cruise 2018-05 deck pressure measurements showed the pressure to be within ±0.04db when the correct offset was used. The offset will be applied during calibration of events #1-33 and will reduce the pressures by 0.2366db.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the setting worked well in making the fluorescence offset reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts for CTD #1 and 3 for CTD #2 to see what shifts are appropriate for each sensor.

For CTD #1222 the best setting was -0.7 records for the primary conductivity and -0.5 for the secondary conductivity.

For CTD #0443 the best setting was -0.3records for the primary conductivity and -0.5records for the secondary conductivity. The primary alignment made little difference as the T-S plots were very noisy with any setting.

SHIFT was run twice applying the settings above to the primary and secondary channels. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The sensors on CTD #1222 were used for one other cruise since it was last serviced at the factory.  The flushing of Niskin bottles was poor so recalibration was based on this cruise.

The sensors on CTD #0443 have not been used since it was last serviced.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.

CTD #1 - The temperature was slightly high at about 500m at station P5. The salinity was high at about 75m for stations P4, P5, P6 and P8. This looks real as there is a good correspondence to bottles. The climatology for the temperature does not cover that level, likely due to very high variability.

CTD #2 - The only excursions were at stations P18, P19 and P20 where the temperature was slightly low at the surface and the salinity was slightly high at the base of the halocline.

None of these excursions look like cases of calibration problems.

Repeat Casts –

CTD #1 - There were repeat casts.

- Casts #20 and #24 occurred about 5 hours apart at P8. Differences along lines of constant σt were <0.002º for the temperature channels and <0.0002psu for salinity at σt= 27.47 (near 1300m). This shows excellent repeatability.

CTD #2 - There were repeat casts.

- Casts #53 and #58 occurred about 8 hours apart at P26. Differences along lines of constant σt were <0.002º for the temperature channels and <0.0005psu for salinity at σt= 27.62 (near 1900m). This shows excellent repeatability.

Repeat Casts with different CTDs – Casts #33 and 36 were run at the same site before and after switching CTDs. The second cast only went to 300m where short-term variability is likely to be fairly high. The differences along lines of constant sigma-t are small at 175m at ≤0.03C° and ≤0.006psu. This is good repeatability for this level.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

For CTD #1 the primary salinity was closest to the bottles but the secondary channels look more stable in T-S space.

For CTD #2 the secondary salinity was closest to the bottle salinity and the secondary channels look much more stable in T-S space.

So the secondary channels were selected for editing for both CTDs.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing. The descent rate was generally noisy for the offshore casts and for some casts there were many complete reversals of direction resulting in considerable corruption by shed wakes. However the corruption was less than seen during many other winter Line P cruises.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; further editing was applied to 1 file. Some unstable features remain but they are in areas of active mixing and upcast data have similar features.

The chief scientist noted that there were many stops during the downcast for event #41. Comparison plots of files #41 and #42 at the same site show that the data from event #41 are worth archiving.

##### Recalibration

File 2018-01-recal1 was prepared to apply the following corrections based on the study in section 5:

- For CTD #1

Pressure and Depth: subtract 0.2366 to fix error in conversion

Secondary salinity: subtract 0.0038

Dissolved oxygen: apply linear correction with slope 1.0327 and offset=0.0627

- For CTD #2

Secondary Salinity: subtract 0.0007

Dissolved oxygen: apply linear correction with slope 1.0283 and offset 0.0491

This correction was first applied to the SAMAVG and MRGCLN2 files. COMPARE was run using the recalibrated CTD data. The results (using the same data as in the original comparison) from each CTD were examined separately.

CTD #1 – The average salinity difference was <0.0001psu after recalibration

CTD #1 – The CTD DO values were higher than the bottle values by an average of 0.0028mL/L. There were only deep bottles included.

CTD #2 – The average salinity difference was 0.0004psu after recalibration

CTD #2 – The CTD DO values were higher than the bottle values by an average of 0.024mL/L with the near-surface values tending to highest, likely due to incomplete flushing of bottles. For DO<4mL/L the average difference is +0.0042mL/L.

These show that the recalibrations were appropriate, but that near-surface dissolved oxygen values from the CTD are higher than bottles, most likely due to inefficient flushing.

See file 2018-01-sal-comp2.xls and file 2018-01-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.027mL/L (standard deviation of 0.0432L/L). A display of the differences versus pressure shows that the DO sensor reads higher than bottles above 1000m and lower below that. This is likely due to incomplete flushing of bottles which has the opposite effect above and below the Oxygen Minimum Zone. This comparison is used to make a rough estimate of accuracy for the DO sensor, assuming that most of the larger outliers are due to poor flushing of Niskin bottles.

The data from the 2 sensors were also examined separately, but little could be deduced from that as there were so many fewer and shallower bottles from CTD #1. No obvious problems were found.

No further recalibration is justified. See 2018-01-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. The Dissolved Oxygen data from the bottom of cast #33 were found to be bad and it is likely other channels are affected. The CTD communication failed shortly after that record. All records after scan 54211 were removed from the FIL file using CLIP.

Bin-average was rerun. This removed the bad data and the DO profile looks fine.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 25, 26, 27, 33, 37, 43, 53, 62, 64, 72, 73, 74 and 79.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to remove the header entry TIME UNITS, fix the vessel name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that some

records were removed in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 300db

±0.2 mL/L from 300db to 500db

±0.05 mL/L from 500db to 2000db

Low by up to 0.03mL/L below 2000db

For details on the processing see document: 2018-01_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~78% to 115% with the lowest value in Saanich Inlet and other slightly low values in Haro Strait. The highest values were in the Strait of Georgia. All values from P1 to P26 plus station CPF1 were between 96% and 104%. Of those stations those westward of P12 were all>100%. The results are consistent with good dissolved oxygen calibration since we usually see values ~103% offshore. There was no obvious difference between the two sensors.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 47, 50, 52, 53, 56, 57, 59, 67, 69, 71, 74, 82, 99.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, remove the header entry TIME UNITS, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-01-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. One problem was found and corrected:

- File #60 should have had 4 records removed where there was no sampling. A return was made to the MRG file to fix that and then later steps were corrected.

After discussions with the chief scientist a few changes were made to flags and comments and to nutrient values for sample 547.

Standards check and a header check were run on all files. A problem was found in file #60 and steps from CALIBRATE had to be redone. No further problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 6 hex files.

There was no flow meter and no external temperature sensor.

The sampling interval was 30s for the first file and 5s thereafter.

There are notes from the chief scientist indicating possible problems with time and/or positions for the first file.

a.) Checking calibrations

The configuration file changed several times through the cruise. The changes included scan length, sampling interval and whether the scan times were added. The temperature and conductivity parameters did not change. Those were checked and are correct.

b.) Conversion of Files

The 6 files were converted to CNV files using the configuration file used in acquisition for each.

As noted in the sampling notes from the chief scientist, the intake thermistor did not work, so the data were reconverted without that channel.

##### c.) WILDEDIT

Plots show discrete single-point spikes in the salinity channel, so tests were run using a variety of settings to remove them using WILDEDIT. The best results were found by using parameters:

Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

Keep data within this distance of the mean = 0

The routine was run on all files.

##### c.) Conversion to IOS Header Format and Initial Quality Checks

The CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots showed that the salinity signal had some one-sided spikes likely to be due to bubbles, but they are well separated so may be amenable to editing.

A track plot shows problems with files 1 and 2. File 3 seems ok. Files 4 and 5 contain almost no data. File 6 is large and seems fine. Checks were made by trying to match times in various records to see if positions were close.

- File #1 – This file overlaps in time with file #2.  Time/locations do not match the CTD log book. Positions vary little –look like all from Saanich Inlet. The record ends at 17:05 but file #2 starts at 15:12. Even if we subtract 2 hours from the file #1 times, we do not come close to matching the log book entries. So this is a bad file. Neither time nor positions can be trusted. There is also no match with the ship track records. Not be processed.

- File #2 – Track plot looks odd. At 16:17 (P1-event 13) the time/location matches the CTD log book. But then instantly the position changes to one that the ship was at a few hours later. So there are some doubts about time and/or positions. At 18:19 (P2-event 14) the time/position again matches the log book but the match is poor for event 15 at the same site. This is mysterious. The temperature and salinity change suddenly between those two stops, so it looks like the records between P1 and P2 are unreliable. There are no loop samples that intersect with the questionable sections. At P3 the match is poor, but at P4 the time/positions look fine. The file then ends. This file does not look reliable enough to archive. Not to be processing.

- File #3 – Track plot looks ok. Time/location match CTD log book. To be processed.

- File #4 – 3 minutes only. Not worth processing

- File #5 – 13 minutes only. Time/location match CTD log book. Don’t know why it was stopped. Very short and during a ship stop so not worth processing.

- File #6 – Large file and track plot looks good. Time/place entries match log file. Looks good. To be processed.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 46 cases of overlap.

Similarly, TSG data were found to match the loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. When all casts are included the average differences between latitude and longitude are larger than we normally see, but the median values are both 0°. If we remove the first 4 cases for which there appears to be a poor match in positions, the average is the same as the median. There is one latitude difference during file #2 that does look high at 0.001°, but examination of the full file shows that 11 seconds later there is a better match to the CTD position. So there may remain some problem with positions, it is not a major difference during the stop, but it is possible that it starts again when the ship is in motion. File #2 is not reliable.

The differences in latitude and longitude from event #25 (station P5) to the end of the cruise were all ≤0.0002° and the median differences were 0.0000° for both. This shows both the times and positions are reliable for both systems from cast #25 onwards (at least while the ship was stopped).

d.) Comparison of T, S and Fuor from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 No intake temperature available.

- Flow Rate There was no flow meter.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature from the CTD and the TSG. Positive indicates that the TSG is reading higher than the CTD.

First the data were examined using all available matches:

|  | Tdiff | SALdiff | FLtsg/FLctd |
| --- | --- | --- | --- |
| min | -0.7692 | -3.1339 | 0.18 |
| max | 0.3403 | 1.0881 | 1.61 |
| median | 0.2234 | 0.0166 | 0.88 |
| average | 0.0106 | -0.1529 | 0.02 |
| std dev | 0.2971 | 1.3173 | 0.50 |

Then the first four matches which are clearly bad were removed.

|  | Tdiff | SALdiff | FLtsg/FLctd |
| --- | --- | --- | --- |
| min | -0.4070 | -3.1339 | 0.34 |
| max | 0.3147 | 0.0250 | 1.61 |
| median | 0.2270 | 0.0166 | 0.88 |
| average | 0.1981 | -0.1887 | 0.94 |
| std dev | 0.1280 | 0.6629 | 0.23 |

This brought the temperature average difference closer to the median, but the salinity differences look even worse. The TSG Fluorescence/ CTD Fluorescence ratio looks better.

Finally only offshore casts were included. (Stations P2 to P26)

|  | Tdiff | SALdiff | FLtsg/FLctd |
| --- | --- | --- | --- |
| min | 0.1736 | 0.0097 | 0.69 |
| max | 0.2789 | 0.0250 | 1.61 |
| median | 0.2308 | 0.0173 | 0.87 |
| average | 0.2317 | 0.0177 | 0.95 |
| std dev | 0.0224 | 0.0039 | 0.21 |

For both the temperature and salinity differences the median and average are much closer while that for the fluorescence ratio is slightly farther apart. The standard deviations are much lower for temperature and salinity and slightly lower for fluorescence.

The TSG temperature is expected to be higher than the CTD because of warming in the loop. For the Tully a difference on the order of 0.2C° is expected. With no experience of the Laurier TSG system it is not known how much of the temperature difference is due to warming, but it is likely that most of it is.

The TSG salinity is higher than the CTD salinity by a median value of 0.017psu and an average of 0.0177psu with a standard deviation of 0.004psu.  When only the 10 cases with the lowest standard deviation are included the average difference is 0.0177ps, and the standard deviation is 0.0042psu. We are used to seeing the TSG salinity reading lower than the CTD due to bubbles, so this is a change. This is likely a better reflection of calibration than we usually get.

The ratio of the TSG fluorescence to CTD fluorescence has a median value of 0.88 and standard deviation of ~0.20. When only the 10 cases with the lowest standard deviations are included, the median ratio was 0.99 but the standard deviation was larger at 0.34.

A second comparison was made with CTD data from 5m and the results were virtually the same, showing that the water is well-mixed vertically around 4 to 5m.

For more details see 2018-01-ctd-tsg-comp.xls.

- Loop vs Rosette Bottles

Differences were calculated between samples from the loop and from the rosette samples closest to 5m. There were only 4 salinity samples and 5 extracted chlorophyll samples available.

There are only 4 points of comparison for salinity with all being within 0.0012psu with loops samples lower by a median of -0.0003psu, so the samples are very close.

For extracted CHL the loop samples were lower than the rosette by a median of 0.02ug/L but for 3 of the 5 cases they were identical. So these are good results.

The loop samples were lower than the CTD values in the bottle files by a median of 0.002psu, which is not significant with so few samples.

For more detail see 2018-01-loops-tsg-rosette-comp.xls.

- Loop Bottle - TSG Comparisons

There were 17 loop samples for both salinity and extracted chlorophyll. When the data from February 20th are included the differences are quite high for salinity, though there is little difference in fluorescence comparisons. When those 4 loop samples are excluded, the TSG Salinity is found to be high by a median of 0.011, standard deviation of 0.003psu. The 4 loops taken during stops indicate that the TSG is higher by an average of 0.0136psu while those when moving by an average of 0.0105psu (excluding Feb. 20). The TSG fluorescence is found to be higher than extracted CHL by a factor of about 1.9, standard deviation 0.3.

For more detail see 2018-01-loops-tsg-rosette-comp.xls.

- Calibration History

The temperature and conductivity sensors were recalibrated in December 2017.  No history was found for this equipment.

Conclusions

- The time recorded for files #1 and #2 overlap.

- Either time or positions or both are unreliable for files #1 and 2.

- The TSG temperature is higher than the CTD temperature by about 0.023 C° which is likely to be primarily due to heating in the loop. A proxy for intake temperature will be prepared by subtracting 0.023 from the lab temperature.

- The TSG salinity is higher than the CTD salinity by about 0.017psu between P5 and P25.

- The TSG salinity is higher than loop samples by an average of 0.011psu with a standard deviation of 0.003psu. When underway it is higher than loops by 0.0105psu while when stopped it is higher than loops by about 0.014psu. This might indicate more bubbles in the water while underway.

- The loop salinity was very close to the samples from the 5m rosette bottles.

- There were some single-point one-sided spikes towards lower salinity values but far fewer than seen in the Tully system. These are likely due to bubbles.

- The TSG salinity is reading a little higher than the CTD, which could be partly because it is drawing water from deeper than the CTD samples. However, the water is quite well mixed at 5m so this shouldn’t make a big difference and the TSG salinity is slightly higher than the loops. So it is likely that TSG salinity is high by roughly 0.01psu. Recalibration seems justified.

- The TSG fluorometer is reading a little lower than the CTD fluorometer. It reads much higher than the loop CHL samples, but that is typical of these sensors in offshore waters.

- Only files #3 and #6 will be archived.

e.)Editing

Each file was opened in CTDEDIT to check that there are no obvious bad data. There remain small spikes in salinity that could be real and there is one section of odd looking fluorescence, but again that could be real. No editing was applied.

f.) Recalibration

Since the comparisons of TSG salinity with the offshore CTD data and the loop samples are consistent, and since vertical salinity gradients are very low near the surface in the offshore area so that the need for an exact depth match between TSG and CTD is not as important as usual, there is adequate justification to recalibrate salinity. So 0.01psu will be subtracted from Salinity:T0:C0.

Since there was no external temperature available a proxy was created. First, ADD CHANNEL was run to add channel Temperature:Lab with the channel values set equal to Temperature:Primary. In the next step 0.23C degrees will be subtracted from Temperature:Primary based on the comparison with offshore CTD casts described above.

CALIBRATE was applied to files #3 and 6 using file 2018-01-recal-tsg.ccf  to subtract 0.01psu from channel Salinity:T0:C0 and to subtract 0.23C degrees from channel Temperature:Primary.

g.) Preparing Final Files

The files were bin-averaged over 6 scans, thus reducing the interval to 30s.

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference,  Flag and Position:New channels.

CLEAN was run to fix the channel limits in the headers.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats. Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found except that there is a station name in the files. That indicates where the file started, but is meaningless for these data, so it was removed.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The Chief Scientist provided file 2018-01 Loop log.xlsx which included event numbers, sample numbers and what was sampled. Times and dates, latitude and longitude (using negative sign for longitude) were added to this file based on the log entries, selecting the end time of samples taken during rosette casts.

The loop data provided by analysts were combined with the information from the loop log and saved as 2018-01 Loop_Data.xlsx.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2018-01-che-surface.csv.

Data from below 7m were removed.

A sample method column was added. ROS was entered for the method.

The Start Time was copied into a second column and the first was formatted for date and the second for time. Columns were rearranged to fit a model 6-line header.

A 6-line header was added to the file and it was saved as 2018-01-surface-6linehdr.csv

The data were sorted on event number, then pressure and added to the 6-line header file.

There were no loop DO samples so there was no need for the Temperature:Draw to derive mass units, so they were removed from the spreadsheet.

The DO in mass units was already in the ROS files.

The loop data were then added to the spreadsheet which was then sorted on event/date/time.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on date, time, sample method and pressure.

(Note For future reference: On another occasion this step failed several times and no cause could be found. Closing and reopening IOS SHELL and doing a reboot did not help. Closing all EXCEL files and opening and closing some convert spreadsheet routines in other projects but not actually running them eventually worked – the program conversion program seemed to remember an input file that had since been fixed.)

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2018-01-surface.loop.  The start and end location information was at the end of the header comments, so it was moved to just below he Administration section.

Final checks turned up an error in the position of one loop file. Once that was fixed the track plot looked reasonable. There is an odd pattern between P22 and P20, but it matches the TSG track and was at a time when heavy weather had interrupted the cruise plan. Only loops were taken during that period.

Plots of variables against time or longitude looked ok.

Particulars – from log book and Chief Scientist’s sampling notes

PAR OFF: 25, 26, 27, 33, 37, 43, 53, 62, 64, 72, 73, 74, 79

Out of Order firing: 17, 19, 20, 36, 37, 69, 76, 84, 86, 95.

Deployment Method –Rosette brought to surface, pumps turned on, rosette taken to 10m and back up. Start archiving and wait 30s, then cast run.

CTD: #1222 for casts 1-33; #0443 for casts 36-100. 45 casts with sampling, 2 without out, 1 test cast

TSG: No flow meter, remote temperature did not work. Time interval 30s for file 1, 5s for file 2-6.

# Test cast; bottle s fired but no CHE file needed.

14. Niskin 5 – top cap was open – samples taken but shouldn’t have – data won’t be good.

17. Sampling interrupted by fire/boat drill for 25min.

33. Malfunction of CTD at 2190 of downcast, NO CHE file.

36. Switch to CTD 0443.

36-37. Wrong fluorometer gain entered in config file. Should be 10X.

41. Winch problems – many stops on way down. Stopped at 145db and retrieved – Cast 42 started for this site.

59.  Station name needs editing to PA-011.

59-60. Sample #295 used twice.

60. Six Niskins closed at 5. Only Niskins 1 and 2 needed – remove rest from CHE file.

76. Niskin 23 not used.

# 82. Station name format needs editing – should be P09. Fixed

# 95. Station name wrong – should be 41. Fixed

# CRUISE SUMMARY     CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1222 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information - 1222 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 18Feb2017 | Factory |  |  |
| Conductivity | 2280 | 23Feb2017 | Factory |  |  |
| Secondary Temp. | 2663 | 18Feb2017 | Factory |  |  |
| Secondary Cond. | 2754 | 22Feb2017 | Factory |  |  |
| Transmissometer | 1185DR | 7Aug2017 | Factory |  |  |
| SBE 43 DO sensor | 3038 | 24Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3642 |  |  |  |  |
| Pressure Sensor | 1222 | 03Mar2015 | Factory |  |  |
| Altimeter | 70613 | 21Mar2016 | Factory |  |  |

| Calibration Information - 0443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 02Sep2017 | Factory |  |  |
| Conductivity | 2128 | 14Sep2017 | Factory |  |  |
| Secondary Temp. | 4752 | 02Sep2017 | Factory |  |  |
| Secondary Cond. | 2399 | 14Sep2017 | Factory |  |  |
| Transmissometer | 1185DR | 7Aug2017 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 2Sep2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3641 |  |  |  |  |
| Pressure Sensor | 0443 | 31Aug2017 | Factory |  |  |
| Altimeter | 70613 | 21Mar2016 | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3274       Cruise ID#:	2018-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3274 | 4Mar17 | Factory |  |  |
| Conductivity | 3274 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | SCF3654 | May2017 |  |  |  |
| Temperature SBE38 | 271 | 3Dec15 |  |  |  |