## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.6 | Run at 2018/06/04 13:41:56 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_BOT | Sal1 |  | Sal1 - Sal_BOT
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_BOT | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2018-01-0013.mrgcor1 | 2018-01-0013.samcor1 | 274 | 4.846 | 4.4846 | 0.15319 | -0.3614 | 32.201 | 32.2046 | 0.0027067 | 0.0036049
2018-01-0014.mrgcor1 | 2018-01-0014.samcor1 | 241 | 5.201 | 4.9641 | 0.17005 | -0.23693 | 30.7934 | 30.7926 | 0.0020651 | -0.00081825
2018-01-0018.mrgcor1 | 2018-01-0018.samcor1 | 352 | 5.373 | 5.1284 | 0.11333 | -0.24459 | 31.1229 | 31.1289 | 0.0074139 | 0.00597
2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 194 | 1251.02 | 1250.54 | 0.26469 | -0.4873 | 34.4618 | 34.462 | 0.00023636 | 0.00024796
2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 301 | 1250.34 | 1250.24 | 0.46379 | -0.099731 | 34.4624 | 34.462 | 0.0002399 | -0.00038147
2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 234 | 999.866 | 999.653 | 0.41828 | -0.21313 | 34.3832 | 34.3816 | 0.00024153 | -0.0015564
2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 241 | 800.969 | 800.732 | 0.19112 | -0.23718 | 34.2919 | 34.2925 | 0.00027953 | 0.00063324
2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 241 | 4.206 | 3.9697 | 0.37041 | -0.23629 | 32.2098 | 32.2079 | 0.00034356 | -0.0018997
2018-01-0025.mrgcor1 | 2018-01-0025.samcor1 | 277 | 4.705 | 4.6513 | 0.13297 | -0.053736 | 32.5413 | 32.5394 | 0.00055822 | -0.0018654
2018-01-0026.mrgcor1 | 2018-01-0026.samcor1 | 287 | 4.615 | 4.2516 | 0.14496 | -0.36343 | 32.5285 | 32.5289 | 0.0011508 | 0.00043106
2018-01-0027.mrgcor1 | 2018-01-0027.samcor1 | 276 | 4.252 | 4.371 | 0.28765 | 0.11898 | 32.5343 | 32.5328 | 0.00021426 | -0.0015183
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 241 | 1999.43 | 1999.19 | 0.41608 | -0.23535 | 34.5875 | 34.5874 | 0.00023356 | -7.6294e-05
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 240 | 1501.56 | 1501.33 | 0.39348 | -0.23352 | 34.5136 | 34.5171 | 0.00036361 | 0.0035438
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 218 | 1250.8 | 1250.65 | 0.51958 | -0.15039 | 34.4585 | 34.4592 | 0.00028033 | 0.00066757
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 241 | 1002.08 | 1001.84 | 0.22426 | -0.23682 | 34.3963 | 34.397 | 0.00023694 | 0.00066757
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 134 | 799.451 | 799.341 | 0.50902 | -0.11023 | 34.3055 | 34.3051 | 0.00022788 | -0.0003891
2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 337 | 6.314 | 6.245 | 0.28382 | -0.068986 | 32.5327 | 32.532 | 0.0014918 | -0.00073624
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 388 | 3001.37 | 3001.28 | 0.47141 | -0.088623 | 34.6531 | 34.6477 | 0.00022113 | -0.0054092
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 390 | 3001.45 | 3001.32 | 0.48454 | -0.12793 | 34.6503 | 34.6477 | 0.00022127 | -0.0026093
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 241 | 2499.92 | 2499.92 | 0.30284 | 0 | 34.6281 | 34.6283 | 0.00022269 | 0.00015259
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 85 | 2001.92 | 2001.58 | 0.57886 | -0.33899 | 34.5858 | 34.5826 | 0.0004425 | -0.0031738
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 241 | 1501.19 | 1501.19 | 0.53867 | 0 | 34.5078 | 34.5059 | 0.00056106 | -0.0018883
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 206 | 1250.67 | 1250.46 | 0.54417 | -0.20752 | 34.4371 | 34.4366 | 0.00047089 | -0.00048828
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 241 | 999.767 | 999.767 | 0.07434 | -0.00018311 | 34.3685 | 34.369 | 0.00024976 | 0.00054169
2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 241 | 4.729 | 4.7288 | 0.18333 | -0.00017452 | 32.5069 | 32.5032 | 0.00021785 | -0.003727
2018-01-0042.mrgcor1 | 2018-01-0042.samcor1 | 287 | 4.871 | 5.0277 | 0.16047 | 0.15673 | 32.4351 | 32.4316 | 0.00021789 | -0.0034561
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 3359.17 | 3359.17 | 0.42789 | -0.00024414 | 34.6628 | 34.6586 | 0.00024389 | -0.0042343
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 3001.02 | 3001.02 | 0.48193 | 0.00024414 | 34.651 | 34.6492 | 0.00026022 | -0.0018043
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 118 | 2497.21 | 2497.05 | 0.54089 | -0.16528 | 34.632 | 34.6317 | 0.000312 | -0.00027084
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 212 | 2251.62 | 2251.84 | 0.3837 | 0.21802 | 34.612 | 34.6115 | 0.0002694 | -0.00049591
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 2001.12 | 2001.12 | 0.33486 | 0.00036621 | 34.5909 | 34.5893 | 0.00026535 | -0.0015602
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 1750.92 | 1750.92 | 0.47466 | 0 | 34.5528 | 34.5509 | 0.00024462 | -0.0019455
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 225 | 1499.18 | 1499.25 | 0.62876 | 0.073486 | 34.5046 | 34.5034 | 0.00031761 | -0.0012016
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 168 | 1251.1 | 1251.19 | 0.68685 | 0.083862 | 34.4438 | 34.4422 | 0.00029329 | -0.001606
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 210 | 1000.83 | 1000.65 | 0.47677 | -0.18152 | 34.3674 | 34.3667 | 0.00029411 | -0.00072861
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 235 | 799.946 | 799.972 | 0.63329 | 0.026184 | 34.282 | 34.2819 | 0.00056436 | -0.00012207
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 601.066 | 601.066 | 0.38111 | 6.1035e-05 | 34.1572 | 34.1552 | 0.00023583 | -0.0019798
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 399.245 | 399.245 | 0.26375 | 0.00021362 | 33.991 | 33.9907 | 0.00028492 | -0.00032043
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 184 | 298.796 | 298.996 | 0.42454 | 0.19995 | 33.9079 | 33.9081 | 0.00062623 | 0.00019074
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 249.854 | 249.854 | 0.19536 | 0.00016785 | 33.8814 | 33.8809 | 0.0003803 | -0.00051117
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 240 | 197.41 | 197.406 | 0.64855 | -0.0042114 | 33.8697 | 33.8678 | 0.00027032 | -0.0018692
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 176.653 | 176.653 | 0.4859 | -0.00019836 | 33.848 | 33.8459 | 0.0002575 | -0.0021324
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 149 | 150.828 | 150.809 | 0.56778 | -0.018722 | 33.8257 | 33.8065 | 0.00061781 | -0.019238
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 125.374 | 125.374 | 0.18371 | -4.5776e-05 | 33.5851 | 33.5822 | 0.00038871 | -0.0029221
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 99.345 | 99.3446 | 0.55205 | -0.00039673 | 33.2732 | 33.2813 | 0.003982 | 0.0080833
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 74.085 | 74.0854 | 0.2866 | 0.00039673 | 32.4596 | 32.4526 | 0.00052981 | -0.0069618
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 239 | 48.794 | 48.8029 | 0.59042 | 0.0088882 | 32.4528 | 32.4502 | 0.0014774 | -0.0025597
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 241 | 24.373 | 24.3725 | 0.41546 | -0.00049782 | 32.4537 | 32.4507 | 0.00024931 | -0.0030022
2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 308 | 5.01 | 4.8688 | 0.26004 | -0.14125 | 32.453 | 32.4505 | 0.00023054 | -0.0024567
2018-01-0044.mrgcor1 | 2018-01-0044.samcor1 | 272 | 5.197 | 5.1379 | 0.43183 | -0.05911 | 32.4283 | 32.4257 | 0.00024738 | -0.0026283
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 76 | 3500.4 | 3500.45 | 0.59335 | 0.045898 | 34.6695 | 34.6695 | 0.00029215 | -4.5776e-05
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 95 | 2500.68 | 2500.84 | 0.5548 | 0.15527 | 34.6199 | 34.6223 | 0.00027091 | 0.0024109
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 213 | 1501.12 | 1500.97 | 0.48153 | -0.14307 | 34.5009 | 34.5 | 0.00031518 | -0.00089645
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 110 | 1250.6 | 1250.43 | 0.59164 | -0.16565 | 34.4429 | 34.4439 | 0.00027072 | 0.0010071
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 241 | 1000.16 | 1000.16 | 0.46508 | -0.00048828 | 34.3675 | 34.368 | 0.00025738 | 0.00046158
2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 241 | 4.708 | 4.7081 | 0.49625 | 0.00010347 | 32.505 | 32.5024 | 0.00035994 | -0.0026016
2018-01-0059.mrgcor1 | 2018-01-0059.samcor1 | 241 | 10.292 | 10.2916 | 0.20251 | -0.00041008 | 32.4936 | 32.4923 | 0.00021236 | -0.0012817
2018-01-0059.mrgcor1 | 2018-01-0059.samcor1 | 241 | 7.408 | 7.4083 | 0.12741 | 0.00034046 | 32.4955 | 32.4922 | 0.00079649 | -0.0033226
2018-01-0059.mrgcor1 | 2018-01-0059.samcor1 | 241 | 4.69 | 4.6901 | 0.068358 | 6.628e-05 | 32.4927 | 32.4922 | 0.00030366 | -0.00048828
2018-01-0060.mrgcor1 | 2018-01-0060.samcor1 | 511 | 5.325 | 5.1559 | 0.33354 | -0.16905 | 32.4744 | 32.4733 | 0.00078684 | -0.001091
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 592 | 2006.34 | 2006.23 | 0.51521 | -0.10413 | 34.5794 | 34.581 | 0.00027174 | 0.0015945
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 627 | 2005.8 | 2005.85 | 0.55561 | 0.05188 | 34.5803 | 34.581 | 0.00026776 | 0.00068665
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 477 | 2004.54 | 2004.65 | 0.507 | 0.11682 | 34.5815 | 34.581 | 0.00025334 | -0.00052261
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 373 | 2004.05 | 2004.33 | 0.46464 | 0.28772 | 34.5814 | 34.581 | 0.00024556 | -0.00043869
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 502 | 2004.65 | 2004.72 | 0.53113 | 0.072388 | 34.5817 | 34.581 | 0.0002545 | -0.00072479
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 610 | 2006.17 | 2006.13 | 0.52549 | -0.036133 | 34.5795 | 34.581 | 0.00026803 | 0.0014954
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 207 | 2007.52 | 2006.98 | 0.32514 | -0.53967 | 34.5813 | 34.581 | 0.00027863 | -0.00028992
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 531 | 2006.61 | 2006.37 | 0.47902 | -0.23914 | 34.5812 | 34.581 | 0.00027389 | -0.00020981
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 418 | 2004.23 | 2004.47 | 0.47089 | 0.24231 | 34.5793 | 34.581 | 0.00025054 | 0.0016708
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 102 | 2002.98 | 2003.48 | 0.41911 | 0.50354 | 34.5765 | 34.581 | 0.00021403 | 0.0044823
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 381 | 2004.1 | 2004.38 | 0.45437 | 0.27026 | 34.5783 | 34.581 | 0.00024922 | 0.0026627
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 593 | 2006.26 | 2006.2 | 0.51109 | -0.063721 | 34.5737 | 34.581 | 0.00027036 | 0.0072937
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 370 | 2007.09 | 2006.68 | 0.4236 | -0.41687 | 34.5796 | 34.581 | 0.0002753 | 0.0014
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 622 | 2005.85 | 2005.9 | 0.54605 | 0.052856 | 34.5864 | 34.581 | 0.00026699 | -0.0054092
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 409 | 2004.23 | 2004.46 | 0.46229 | 0.22864 | 34.5776 | 34.581 | 0.00025135 | 0.0033722
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 405 | 2004.22 | 2004.45 | 0.45815 | 0.23413 | 34.5795 | 34.581 | 0.00025067 | 0.0014725
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 637 | 2005.83 | 2005.87 | 0.56339 | 0.0448 | 34.5776 | 34.581 | 0.00026687 | 0.0033913
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 474 | 2006.84 | 2006.52 | 0.4753 | -0.3158 | 34.5764 | 34.581 | 0.00027274 | 0.0045967
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 623 | 2005.87 | 2005.9 | 0.54696 | 0.029541 | 34.5775 | 34.581 | 0.00026749 | 0.0034904
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 551 | 2004.8 | 2004.83 | 0.5828 | 0.027954 | 34.5796 | 34.581 | 0.00025609 | 0.0013733
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 592 | 2005.14 | 2005.18 | 0.60715 | 0.042358 | 34.5804 | 34.581 | 0.00025947 | 0.00057983
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 623 | 2005.87 | 2005.9 | 0.54696 | 0.030518 | 34.5799 | 34.581 | 0.00026749 | 0.001091
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 635 | 2005.67 | 2005.73 | 0.57309 | 0.062988 | 34.5797 | 34.581 | 0.00026477 | 0.0012856
2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 241 | 5.317 | 5.3172 | 0.21505 | 0.00015783 | 32.4719 | 32.471 | 0.00023661 | -0.00085449
2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 207 | 3500.75 | 3500.56 | 0.57018 | -0.19263 | 34.6692 | 34.6698 | 0.00026802 | 0.00056457
2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 241 | 3001.6 | 3001.6 | 0.36652 | -0.00048828 | 34.6506 | 34.6524 | 0.00023354 | 0.0018311
2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 182 | 2500.68 | 2500.72 | 0.5681 | 0.046387 | 34.6255 | 34.6264 | 0.00023328 | 0.00090408
2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 231 | 2001.39 | 2001.44 | 0.59005 | 0.051514 | 34.5796 | 34.5799 | 0.00027076 | 0.00033188
2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 289 | 5.106 | 5.1313 | 0.2037 | 0.025263 | 32.4857 | 32.4838 | 0.00026579 | -0.0018692
2018-01-0063.mrgcor1 | 2018-01-0063.samcor1 | 462 | 4.641 | 4.5837 | 0.31405 | -0.057264 | 32.4906 | 32.4882 | 0.00025238 | -0.0023956
2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 226 | 3499.55 | 3499.49 | 0.55906 | -0.066894 | 34.67 | 34.6701 | 0.00024917 | 9.5367e-05
2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 219 | 2999.72 | 2999.82 | 0.55184 | 0.104 | 34.6528 | 34.6538 | 0.00025001 | 0.0010414
2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 209 | 2499.1 | 2498.87 | 0.42121 | -0.23511 | 34.6423 | 34.6279 | 0.0002415 | -0.014431
2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 193 | 1999.49 | 1999.6 | 0.56406 | 0.10925 | 34.5814 | 34.582 | 0.00028839 | 0.00062942
2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 359 | 4.829 | 4.8215 | 0.16839 | -0.0075154 | 32.4965 | 32.4954 | 0.00022546 | -0.0010567
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 281 | 3500.57 | 3500.41 | 0.61028 | -0.16309 | 34.6696 | 34.6684 | 0.00028334 | -0.001152
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 296 | 3500.53 | 3500.33 | 0.6123 | -0.19482 | 34.6668 | 34.6684 | 0.00028327 | 0.0016441
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 232 | 3001.5 | 3001.54 | 0.63989 | 0.039551 | 34.652 | 34.6522 | 0.00025949 | 0.00016785
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 96 | 2500.49 | 2500.45 | 0.57235 | -0.038086 | 34.6267 | 34.6264 | 0.00026286 | -0.0002594
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 201 | 1999.92 | 1999.9 | 0.59436 | -0.019165 | 34.5805 | 34.5813 | 0.00024535 | 0.00080109
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 230 | 1500.1 | 1500.05 | 0.63551 | -0.049072 | 34.4912 | 34.4919 | 0.00028229 | 0.00068665
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 241 | 1249.05 | 1249.05 | 0.22279 | 0.00012207 | 34.4362 | 34.4354 | 0.00027479 | -0.00078964
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 241 | 999.07 | 999.07 | 0.3908 | 0.00024414 | 34.3566 | 34.3548 | 0.00025841 | -0.0018463
2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 241 | 4.803 | 4.8028 | 0.15704 | -0.00023222 | 32.4433 | 32.4402 | 0.00024845 | -0.0030708
2018-01-0073.mrgcor1 | 2018-01-0073.samcor1 | 314 | 5.903 | 5.6912 | 0.22175 | -0.21177 | 32.4712 | 32.4687 | 0.00058606 | -0.0024757
2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 155 | 3499.64 | 3499.64 | 0.5803 | -0.0026856 | 34.6673 | 34.6674 | 0.00026825 | 6.485e-05
2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 206 | 2999.4 | 2999.59 | 0.57655 | 0.18921 | 34.653 | 34.653 | 0.00026722 | 3.0518e-05
2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 239 | 2500.44 | 2500.44 | 0.57803 | 0.0085449 | 34.6251 | 34.6297 | 0.00026886 | 0.0045509
2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 198 | 1999.06 | 1999.31 | 0.61823 | 0.24756 | 34.5861 | 34.5854 | 0.0002797 | -0.00071716
2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 306 | 4.259 | 4.2865 | 0.19909 | 0.027477 | 32.4588 | 32.4604 | 0.00030169 | 0.0015564
2018-01-0075.mrgcor1 | 2018-01-0075.samcor1 | 297 | 4.681 | 4.7738 | 0.10704 | 0.092778 | 32.465 | 32.4623 | 0.00024546 | -0.0027199
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 349 | 3500.18 | 3500.38 | 0.31366 | 0.20581 | 34.6635 | 34.6648 | 0.00023441 | 0.0012512
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 349 | 3500.48 | 3500.38 | 0.31366 | -0.092285 | 34.666 | 34.6648 | 0.00023441 | -0.0012474
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 3000.22 | 3000.22 | 0.40874 | 0.00048828 | 34.6534 | 34.6538 | 0.00024982 | 0.0003624
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 2500.65 | 2500.65 | 0.30522 | -0.00024414 | 34.6237 | 34.6282 | 0.00022822 | 0.0044517
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 2000.4 | 2000.4 | 0.19559 | -0.00024414 | 34.5787 | 34.582 | 0.00022901 | 0.0033073
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 1500.27 | 1500.27 | 0.32017 | -0.00024414 | 34.502 | 34.5034 | 0.00031585 | 0.0013847
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 1248.8 | 1248.8 | 0.21794 | 0 | 34.4348 | 34.4371 | 0.00023718 | 0.0023384
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 1000.36 | 1000.37 | 0.098601 | 0.00030518 | 34.3592 | 34.3604 | 0.00094628 | 0.0012283
2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 241 | 4.319 | 4.3193 | 0.13926 | 0.00029468 | 32.4901 | 32.49 | 0.00022601 | -0.00014114
2018-01-0080.mrgcor1 | 2018-01-0080.samcor1 | 325 | 5.279 | 4.9074 | 0.35175 | -0.37155 | 32.3686 | 32.3672 | 0.00021483 | -0.0014038
2018-01-0081.mrgcor1 | 2018-01-0081.samcor1 | 399 | 4.698 | 4.7851 | 0.16176 | 0.08713 | 32.1439 | 32.1433 | 0.00048253 | -0.00063324
2018-01-0082.mrgcor1 | 2018-01-0082.samcor1 | 312 | 5.334 | 5.3624 | 0.11348 | 0.028452 | 32.1282 | 32.125 | 0.00034554 | -0.0032196
2018-01-0084.mrgcor1 | 2018-01-0084.samcor1 | 241 | 5.104 | 5.1037 | 0.033902 | -0.00024509 | 31.7208 | 31.7113 | 0.00027006 | -0.00951
2018-01-0086.mrgcor1 | 2018-01-0086.samcor1 | 241 | 4.636 | 4.6364 | 0.027108 | 0.00037766 | 30.6364 | 30.6282 | 0.0002957 | -0.0082073
2018-01-0088.mrgcor1 | 2018-01-0088.samcor1 | 241 | 49.196 | 49.1956 | 0.047924 | -0.0003891 | 30.0779 | 30.0898 | 0.0073299 | 0.011927
2018-01-0088.mrgcor1 | 2018-01-0088.samcor1 | 241 | 5.132 | 5.1323 | 0.030356 | 0.00035286 | 29.6723 | 29.6596 | 0.00618 | -0.012714
2018-01-0092.mrgcor1 | 2018-01-0092.samcor1 | 241 | 5.094 | 5.0944 | 0.029887 | 0.00040674 | 26.7609 | 25.9512 | 0.05348 | -0.80971
2018-01-0098.mrgcor1 | 2018-01-0098.samcor1 | 241 | 4.997 | 4.9965 | 0.028758 | -0.00048971 | 27.8101 | 27.6088 | 0.013793 | -0.20129
2018-01-0100.mrgcor1 | 2018-01-0100.samcor1 | 241 | 4.533 | 4.5327 | 0.031747 | -0.00026131 | 27.9756 | 27.8693 | 0.028466 | -0.10627
Create Fit Sheet
Plot
Plot

## File List
Number of Files= |  | 33
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2018-01-0013.mrgcor1 | 2018-01-0013.samcor1 | 11 | 11 | 2018-01-0013.mrgcor1 & 2018-01-0013.samcor1
2 | 2018-01-0014.mrgcor1 | 2018-01-0014.samcor1 | 13 | 13 | 2018-01-0014.mrgcor1 & 2018-01-0014.samcor1
3 | 2018-01-0018.mrgcor1 | 2018-01-0018.samcor1 | 15 | 15 | 2018-01-0018.mrgcor1 & 2018-01-0018.samcor1
4 | 2018-01-0020.mrgcor1 | 2018-01-0020.samcor1 | 17 | 21 | 2018-01-0020.mrgcor1 & 2018-01-0020.samcor1
5 | 2018-01-0025.mrgcor1 | 2018-01-0025.samcor1 | 23 | 23 | 2018-01-0025.mrgcor1 & 2018-01-0025.samcor1
6 | 2018-01-0026.mrgcor1 | 2018-01-0026.samcor1 | 25 | 25 | 2018-01-0026.mrgcor1 & 2018-01-0026.samcor1
7 | 2018-01-0027.mrgcor1 | 2018-01-0027.samcor1 | 27 | 27 | 2018-01-0027.mrgcor1 & 2018-01-0027.samcor1
8 | 2018-01-0032.mrgcor1 | 2018-01-0032.samcor1 | 29 | 34 | 2018-01-0032.mrgcor1 & 2018-01-0032.samcor1
9 | 2018-01-0037.mrgcor1 | 2018-01-0037.samcor1 | 36 | 43 | 2018-01-0037.mrgcor1 & 2018-01-0037.samcor1
10 | 2018-01-0042.mrgcor1 | 2018-01-0042.samcor1 | 45 | 45 | 2018-01-0042.mrgcor1 & 2018-01-0042.samcor1
11 | 2018-01-0043.mrgcor1 | 2018-01-0043.samcor1 | 47 | 69 | 2018-01-0043.mrgcor1 & 2018-01-0043.samcor1
12 | 2018-01-0044.mrgcor1 | 2018-01-0044.samcor1 | 71 | 71 | 2018-01-0044.mrgcor1 & 2018-01-0044.samcor1
13 | 2018-01-0053.mrgcor1 | 2018-01-0053.samcor1 | 73 | 78 | 2018-01-0053.mrgcor1 & 2018-01-0053.samcor1
14 | 2018-01-0059.mrgcor1 | 2018-01-0059.samcor1 | 80 | 82 | 2018-01-0059.mrgcor1 & 2018-01-0059.samcor1
15 | 2018-01-0060.mrgcor1 | 2018-01-0060.samcor1 | 84 | 84 | 2018-01-0060.mrgcor1 & 2018-01-0060.samcor1
16 | 2018-01-0061.mrgcor1 | 2018-01-0061.samcor1 | 86 | 109 | 2018-01-0061.mrgcor1 & 2018-01-0061.samcor1
17 | 2018-01-0062.mrgcor1 | 2018-01-0062.samcor1 | 111 | 115 | 2018-01-0062.mrgcor1 & 2018-01-0062.samcor1
18 | 2018-01-0063.mrgcor1 | 2018-01-0063.samcor1 | 117 | 117 | 2018-01-0063.mrgcor1 & 2018-01-0063.samcor1
19 | 2018-01-0064.mrgcor1 | 2018-01-0064.samcor1 | 119 | 123 | 2018-01-0064.mrgcor1 & 2018-01-0064.samcor1
20 | 2018-01-0072.mrgcor1 | 2018-01-0072.samcor1 | 125 | 133 | 2018-01-0072.mrgcor1 & 2018-01-0072.samcor1
21 | 2018-01-0073.mrgcor1 | 2018-01-0073.samcor1 | 135 | 135 | 2018-01-0073.mrgcor1 & 2018-01-0073.samcor1
22 | 2018-01-0074.mrgcor1 | 2018-01-0074.samcor1 | 137 | 141 | 2018-01-0074.mrgcor1 & 2018-01-0074.samcor1
23 | 2018-01-0075.mrgcor1 | 2018-01-0075.samcor1 | 143 | 143 | 2018-01-0075.mrgcor1 & 2018-01-0075.samcor1
24 | 2018-01-0079.mrgcor1 | 2018-01-0079.samcor1 | 145 | 153 | 2018-01-0079.mrgcor1 & 2018-01-0079.samcor1
25 | 2018-01-0080.mrgcor1 | 2018-01-0080.samcor1 | 155 | 155 | 2018-01-0080.mrgcor1 & 2018-01-0080.samcor1
26 | 2018-01-0081.mrgcor1 | 2018-01-0081.samcor1 | 157 | 157 | 2018-01-0081.mrgcor1 & 2018-01-0081.samcor1
27 | 2018-01-0082.mrgcor1 | 2018-01-0082.samcor1 | 159 | 159 | 2018-01-0082.mrgcor1 & 2018-01-0082.samcor1
28 | 2018-01-0084.mrgcor1 | 2018-01-0084.samcor1 | 161 | 161 | 2018-01-0084.mrgcor1 & 2018-01-0084.samcor1
29 | 2018-01-0086.mrgcor1 | 2018-01-0086.samcor1 | 163 | 163 | 2018-01-0086.mrgcor1 & 2018-01-0086.samcor1
30 | 2018-01-0088.mrgcor1 | 2018-01-0088.samcor1 | 165 | 166 | 2018-01-0088.mrgcor1 & 2018-01-0088.samcor1
31 | 2018-01-0092.mrgcor1 | 2018-01-0092.samcor1 | 168 | 168 | 2018-01-0092.mrgcor1 & 2018-01-0092.samcor1
32 | 2018-01-0098.mrgcor1 | 2018-01-0098.samcor1 | 170 | 170 | 2018-01-0098.mrgcor1 & 2018-01-0098.samcor1
33 | 2018-01-0100.mrgcor1 | 2018-01-0100.samcor1 | 172 | 172 | 2018-01-0100.mrgcor1 & 2018-01-0100.samcor1

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal1 - Sal_BOT) | Pressure | 8
2 | (Sal1 - Sal_BOT) | Pressure | 25

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00043744 |  | X | CALC Y
m= | 4.1438e-08 | 8.0866e-07 |  | Std Dev of Resid |  | 0.00082389 |  | 799.341 | -3.8684e-05
b= | -7.1807e-05 | 0.0009892 |  | Resid Sum of Squares |  | 4.0727e-06 |  | 1999.19 | 1.1035e-05
# of Values Extracted = |  |  | 17
# of Values Used in Fit = |  |  | 8
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 4.4846 | 32.2046 | 0.0027067 | 0.0036049 | -7.1621e-05 | -0.0036765 |  | 1250.54 | 0.00024796
2 | NO | 4.9641 | 30.7926 | 0.0020651 | -0.00081825 | -7.1601e-05 | 0.00074665 |  | 1250.24 | -0.00038147
3 | NO | 5.1284 | 31.1289 | 0.0074139 | 0.00597 | -7.1594e-05 | -0.0060416 |  | 999.653 | -0.0015564
4 | YES | 1250.54 | 34.462 | 0.00023636 | 0.00024796 | -1.9988e-05 | -0.00026794 |  | 800.732 | 0.00063324
4 | YES | 1250.24 | 34.462 | 0.0002399 | -0.00038147 | -2e-05 | 0.00036147 |  | 1999.19 | -7.6294e-05
4 | YES | 999.653 | 34.3816 | 0.00024153 | -0.0015564 | -3.0384e-05 | 0.001526 |  | 1250.65 | 0.00066757
4 | YES | 800.732 | 34.2925 | 0.00027953 | 0.00063324 | -3.8627e-05 | -0.00067187 |  | 1001.84 | 0.00066757
4 | NO | 3.9697 | 32.2079 | 0.00034356 | -0.0018997 | -7.1642e-05 | 0.0018281 |  | 799.341 | -0.0003891 | -2.3365e-05 | 0.00076294
5 | NO | 4.6513 | 32.5394 | 0.00055822 | -0.0018654 | -7.1614e-05 | 0.0017938
6 | NO | 4.2516 | 32.5289 | 0.0011508 | 0.00043106 | -7.1631e-05 | -0.00050269
7 | NO | 4.371 | 32.5328 | 0.00021426 | -0.0015183 | -7.1626e-05 | 0.0014466
8 | YES | 1999.19 | 34.5874 | 0.00023356 | -7.6294e-05 | 1.1035e-05 | 8.7328e-05
8 | NO | 1501.33 | 34.5171 | 0.00036361 | 0.0035438 | -9.5955e-06 | -0.0035534
8 | YES | 1250.65 | 34.4592 | 0.00028033 | 0.00066757 | -1.9983e-05 | -0.00068756
8 | YES | 1001.84 | 34.397 | 0.00023694 | 0.00066757 | -3.0293e-05 | -0.00069787
8 | YES | 799.341 | 34.3051 | 0.00022788 | -0.0003891 | -3.8684e-05 | 0.00035041
8 | NO | 6.245 | 32.532 | 0.0014918 | -0.00073624 | -7.1548e-05 | 0.00066469
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00098839 |  | X | CALC Y
m= | -9.3314e-08 | 3.5458e-07 |  | Std Dev of Resid |  | 0.0022186 |  | 601.066 | 0.00054747
b= | 0.00060355 | 0.00081224 |  | Resid Sum of Squares |  | 0.00034455 |  | 3500.56 | 0.0002769
# of Values Extracted = |  |  | 113
# of Values Used in Fit = |  |  | 72
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
9 | YES | 3001.28 | 34.6477 | 0.00022113 | -0.0054092 | 0.00032349 | 0.0057327 |  | 3001.28 | -0.0054092
9 | YES | 3001.32 | 34.6477 | 0.00022127 | -0.0026093 | 0.00032349 | 0.0029327 |  | 3001.32 | -0.0026093
9 | YES | 2499.92 | 34.6283 | 0.00022269 | 0.00015259 | 0.00037028 | 0.00021769 |  | 2499.92 | 0.00015259
9 | YES | 2001.58 | 34.5826 | 0.0004425 | -0.0031738 | 0.00041678 | 0.0035906 |  | 2001.58 | -0.0031738
9 | YES | 1501.19 | 34.5059 | 0.00056106 | -0.0018883 | 0.00046347 | 0.0023518 |  | 1501.19 | -0.0018883
9 | YES | 1250.46 | 34.4366 | 0.00047089 | -0.00048828 | 0.00048687 | 0.00097515 |  | 1250.46 | -0.00048828
9 | YES | 999.767 | 34.369 | 0.00024976 | 0.00054169 | 0.00051026 | -3.1425e-05 |  | 999.767 | 0.00054169
9 | NO | 4.7288 | 32.5032 | 0.00021785 | -0.003727 | 0.00060311 | 0.0043301 |  | 3359.17 | -0.0042343
10 | NO | 5.0277 | 32.4316 | 0.00021789 | -0.0034561 | 0.00060308 | 0.0040592 |  | 3001.02 | -0.0018043
11 | YES | 3359.17 | 34.6586 | 0.00024389 | -0.0042343 | 0.0002901 | 0.0045244 |  | 2497.05 | -0.00027084
11 | YES | 3001.02 | 34.6492 | 0.00026022 | -0.0018043 | 0.00032352 | 0.0021279 |  | 2251.84 | -0.00049591
11 | YES | 2497.05 | 34.6317 | 0.000312 | -0.00027084 | 0.00037054 | 0.00064139 |  | 2001.12 | -0.0015602
11 | YES | 2251.84 | 34.6115 | 0.0002694 | -0.00049591 | 0.00039343 | 0.00088934 |  | 1750.92 | -0.0019455
11 | YES | 2001.12 | 34.5893 | 0.00026535 | -0.0015602 | 0.00041682 | 0.001977 |  | 1499.25 | -0.0012016
11 | YES | 1750.92 | 34.5509 | 0.00024462 | -0.0019455 | 0.00044017 | 0.0023857 |  | 1251.19 | -0.001606
11 | YES | 1499.25 | 34.5034 | 0.00031761 | -0.0012016 | 0.00046365 | 0.0016653 |  | 1000.65 | -0.00072861
11 | YES | 1251.19 | 34.4422 | 0.00029329 | -0.001606 | 0.0004868 | 0.0020928 |  | 799.972 | -0.00012207
11 | YES | 1000.65 | 34.3667 | 0.00029411 | -0.00072861 | 0.00051018 | 0.0012388 |  | 601.066 | -0.0019798
11 | YES | 799.972 | 34.2819 | 0.00056436 | -0.00012207 | 0.00052891 | 0.00065098 |  | 3500.45 | -4.5776e-05
11 | YES | 601.066 | 34.1552 | 0.00023583 | -0.0019798 | 0.00054747 | 0.0025273 |  | 2500.84 | 0.0024109
11 | NO | 399.245 | 33.9907 | 0.00028492 | -0.00032043 | 0.0005663 | 0.00088673 |  | 1500.97 | -0.00089645
11 | NO | 298.996 | 33.9081 | 0.00062623 | 0.00019074 | 0.00057565 | 0.00038492 |  | 1250.43 | 0.0010071
11 | NO | 249.854 | 33.8809 | 0.0003803 | -0.00051117 | 0.00058024 | 0.0010914 |  | 1000.16 | 0.00046158
11 | NO | 197.406 | 33.8678 | 0.00027032 | -0.0018692 | 0.00058513 | 0.0024543 |  | 2006.23 | 0.0015945
11 | NO | 176.653 | 33.8459 | 0.0002575 | -0.0021324 | 0.00058707 | 0.0027195 |  | 2005.85 | 0.00068665
11 | NO | 150.809 | 33.8065 | 0.00061781 | -0.019238 | 0.00058948 | 0.019827 |  | 2004.65 | -0.00052261
11 | NO | 125.374 | 33.5822 | 0.00038871 | -0.0029221 | 0.00059185 | 0.0035139 |  | 2004.33 | -0.00043869
11 | NO | 99.3446 | 33.2813 | 0.003982 | 0.0080833 | 0.00059428 | -0.0074891 |  | 2004.72 | -0.00072479
11 | NO | 74.0854 | 32.4526 | 0.00052981 | -0.0069618 | 0.00059664 | 0.0075585 |  | 2006.13 | 0.0014954
11 | NO | 48.8029 | 32.4502 | 0.0014774 | -0.0025597 | 0.000599 | 0.0031587 |  | 2006.98 | -0.00028992
11 | NO | 24.3725 | 32.4507 | 0.00024931 | -0.0030022 | 0.00060128 | 0.0036034 |  | 2006.37 | -0.00020981
11 | NO | 4.8688 | 32.4505 | 0.00023054 | -0.0024567 | 0.0006031 | 0.0030598 |  | 2004.47 | 0.0016708
12 | NO | 5.1379 | 32.4257 | 0.00024738 | -0.0026283 | 0.00060307 | 0.0032314 |  | 2003.48 | 0.0044823
13 | YES | 3500.45 | 34.6695 | 0.00029215 | -4.5776e-05 | 0.00027691 | 0.00032269 |  | 2004.38 | 0.0026627
13 | YES | 2500.84 | 34.6223 | 0.00027091 | 0.0024109 | 0.00037019 | -0.0020407 |  | 2006.2 | 0.0072937
13 | YES | 1500.97 | 34.5 | 0.00031518 | -0.00089645 | 0.00046349 | 0.0013599 |  | 2006.68 | 0.0014
13 | YES | 1250.43 | 34.4439 | 0.00027072 | 0.0010071 | 0.00048687 | -0.00052021 |  | 2005.9 | -0.0054092
13 | YES | 1000.16 | 34.368 | 0.00025738 | 0.00046158 | 0.00051023 | 4.8647e-05 |  | 2004.46 | 0.0033722
13 | NO | 4.7081 | 32.5024 | 0.00035994 | -0.0026016 | 0.00060311 | 0.0032047 |  | 2004.45 | 0.0014725
14 | NO | 10.2916 | 32.4923 | 0.00021236 | -0.0012817 | 0.00060259 | 0.0018843 |  | 2005.87 | 0.0033913
14 | NO | 7.4083 | 32.4922 | 0.00079649 | -0.0033226 | 0.00060286 | 0.0039255 |  | 2006.52 | 0.0045967
14 | NO | 4.6901 | 32.4922 | 0.00030366 | -0.00048828 | 0.00060312 | 0.0010914 |  | 2005.9 | 0.0034904
15 | NO | 5.1559 | 32.4733 | 0.00078684 | -0.001091 | 0.00060307 | 0.0016941 |  | 2004.83 | 0.0013733
16 | YES | 2006.23 | 34.581 | 0.00027174 | 0.0015945 | 0.00041635 | -0.0011782 |  | 2005.18 | 0.00057983
16 | YES | 2005.85 | 34.581 | 0.00026776 | 0.00068665 | 0.00041638 | -0.00027027 |  | 2005.9 | 0.001091
16 | YES | 2004.65 | 34.581 | 0.00025334 | -0.00052261 | 0.00041649 | 0.00093911 |  | 2005.73 | 0.0012856
16 | YES | 2004.33 | 34.581 | 0.00024556 | -0.00043869 | 0.00041652 | 0.00085521 |  | 3500.56 | 0.00056457
16 | YES | 2004.72 | 34.581 | 0.0002545 | -0.00072479 | 0.00041649 | 0.0011413 |  | 3001.6 | 0.0018311
16 | YES | 2006.13 | 34.581 | 0.00026803 | 0.0014954 | 0.00041635 | -0.001079 |  | 2500.72 | 0.00090408
16 | YES | 2006.98 | 34.581 | 0.00027863 | -0.00028992 | 0.00041628 | 0.00070619 |  | 2001.44 | 0.00033188
16 | YES | 2006.37 | 34.581 | 0.00027389 | -0.00020981 | 0.00041633 | 0.00062614 |  | 3499.49 | 9.5367e-05
16 | YES | 2004.47 | 34.581 | 0.00025054 | 0.0016708 | 0.00041651 | -0.0012543 |  | 2999.82 | 0.0010414
16 | YES | 2003.48 | 34.581 | 0.00021403 | 0.0044823 | 0.0004166 | -0.0040657 |  | 1999.6 | 0.00062942
16 | YES | 2004.38 | 34.581 | 0.00024922 | 0.0026627 | 0.00041652 | -0.0022461 |  | 3500.41 | -0.001152
16 | YES | 2006.2 | 34.581 | 0.00027036 | 0.0072937 | 0.00041635 | -0.0068774 |  | 3500.33 | 0.0016441
16 | YES | 2006.68 | 34.581 | 0.0002753 | 0.0014 | 0.0004163 | -0.00098369 |  | 3001.54 | 0.00016785
16 | YES | 2005.9 | 34.581 | 0.00026699 | -0.0054092 | 0.00041638 | 0.0058256 |  | 2500.45 | -0.0002594
16 | YES | 2004.46 | 34.581 | 0.00025135 | 0.0033722 | 0.00041651 | -0.0029557 |  | 1999.9 | 0.00080109
16 | YES | 2004.45 | 34.581 | 0.00025067 | 0.0014725 | 0.00041651 | -0.001056 |  | 1500.05 | 0.00068665
16 | YES | 2005.87 | 34.581 | 0.00026687 | 0.0033913 | 0.00041638 | -0.0029749 |  | 1249.05 | -0.00078964
16 | YES | 2006.52 | 34.581 | 0.00027274 | 0.0045967 | 0.00041632 | -0.0041804 |  | 999.07 | -0.0018463
16 | YES | 2005.9 | 34.581 | 0.00026749 | 0.0034904 | 0.00041638 | -0.0030741 |  | 3499.64 | 6.485e-05
16 | YES | 2004.83 | 34.581 | 0.00025609 | 0.0013733 | 0.00041648 | -0.00095681 |  | 2999.59 | 3.0518e-05
16 | YES | 2005.18 | 34.581 | 0.00025947 | 0.00057983 | 0.00041644 | -0.00016339 |  | 2500.44 | 0.0045509
16 | YES | 2005.9 | 34.581 | 0.00026749 | 0.001091 | 0.00041638 | -0.00067462 |  | 1999.31 | -0.00071716
16 | YES | 2005.73 | 34.581 | 0.00026477 | 0.0012856 | 0.00041639 | -0.00086916 |  | 3500.38 | 0.0012512
16 | NO | 5.3172 | 32.471 | 0.00023661 | -0.00085449 | 0.00060306 | 0.0014575 |  | 3500.38 | -0.0012474
17 | YES | 3500.56 | 34.6698 | 0.00026802 | 0.00056457 | 0.0002769 | -0.00028767 |  | 3000.22 | 0.0003624
17 | YES | 3001.6 | 34.6524 | 0.00023354 | 0.0018311 | 0.00032346 | -0.0015076 |  | 2500.65 | 0.0044517
17 | YES | 2500.72 | 34.6264 | 0.00023328 | 0.00090408 | 0.0003702 | -0.00053388 |  | 2000.4 | 0.0033073
17 | YES | 2001.44 | 34.5799 | 0.00027076 | 0.00033188 | 0.00041679 | 8.4913e-05 |  | 1500.27 | 0.0013847
17 | NO | 5.1313 | 32.4838 | 0.00026579 | -0.0018692 | 0.00060308 | 0.0024723 |  | 1248.8 | 0.0023384 | 0.00040118 | 0.002204
18 | NO | 4.5837 | 32.4882 | 0.00025238 | -0.0023956 | 0.00060313 | 0.0029988
19 | YES | 3499.49 | 34.6701 | 0.00024917 | 9.5367e-05 | 0.000277 | 0.00018164
19 | YES | 2999.82 | 34.6538 | 0.00025001 | 0.0010414 | 0.00032363 | -0.00071778
19 | NO | 2498.87 | 34.6279 | 0.0002415 | -0.014431 | 0.00037038 | 0.014801
19 | YES | 1999.6 | 34.582 | 0.00028839 | 0.00062942 | 0.00041696 | -0.00021246
19 | NO | 4.8215 | 32.4954 | 0.00022546 | -0.0010567 | 0.0006031 | 0.0016598
20 | YES | 3500.41 | 34.6684 | 0.00028334 | -0.001152 | 0.00027692 | 0.001429
20 | YES | 3500.33 | 34.6684 | 0.00028327 | 0.0016441 | 0.00027692 | -0.0013672
20 | YES | 3001.54 | 34.6522 | 0.00025949 | 0.00016785 | 0.00032347 | 0.00015562
20 | YES | 2500.45 | 34.6264 | 0.00026286 | -0.0002594 | 0.00037023 | 0.00062963
20 | YES | 1999.9 | 34.5813 | 0.00024535 | 0.00080109 | 0.00041694 | -0.00038415
20 | YES | 1500.05 | 34.4919 | 0.00028229 | 0.00068665 | 0.00046358 | -0.00022307
20 | YES | 1249.05 | 34.4354 | 0.00027479 | -0.00078964 | 0.000487 | 0.0012766
20 | YES | 999.07 | 34.3548 | 0.00025841 | -0.0018463 | 0.00051033 | 0.0023566
20 | NO | 4.8028 | 32.4402 | 0.00024845 | -0.0030708 | 0.00060311 | 0.0036739
21 | NO | 5.6912 | 32.4687 | 0.00058606 | -0.0024757 | 0.00060302 | 0.0030788
22 | YES | 3499.64 | 34.6674 | 0.00026825 | 6.485e-05 | 0.00027699 | 0.00021214
22 | YES | 2999.59 | 34.653 | 0.00026722 | 3.0518e-05 | 0.00032365 | 0.00029313
22 | YES | 2500.44 | 34.6297 | 0.00026886 | 0.0045509 | 0.00037023 | -0.0041807
22 | YES | 1999.31 | 34.5854 | 0.0002797 | -0.00071716 | 0.00041699 | 0.0011342
22 | NO | 4.2865 | 32.4604 | 0.00030169 | 0.0015564 | 0.00060315 | -0.00095325
23 | NO | 4.7738 | 32.4623 | 0.00024546 | -0.0027199 | 0.00060311 | 0.003323
24 | YES | 3500.38 | 34.6648 | 0.00023441 | 0.0012512 | 0.00027692 | -0.0009743
24 | YES | 3500.38 | 34.6648 | 0.00023441 | -0.0012474 | 0.00027692 | 0.0015243
24 | YES | 3000.22 | 34.6538 | 0.00024982 | 0.0003624 | 0.00032359 | -3.8804e-05
24 | YES | 2500.65 | 34.6282 | 0.00022822 | 0.0044517 | 0.00037021 | -0.0040815
24 | YES | 2000.4 | 34.582 | 0.00022901 | 0.0033073 | 0.00041689 | -0.0028905
24 | YES | 1500.27 | 34.5034 | 0.00031585 | 0.0013847 | 0.00046356 | -0.00092118
24 | YES | 1248.8 | 34.4371 | 0.00023718 | 0.0023384 | 0.00048702 | -0.0018514
24 | NO | 1000.37 | 34.3604 | 0.00094628 | 0.0012283 | 0.00051021 | -0.00071812
24 | NO | 4.3193 | 32.49 | 0.00022601 | -0.00014114 | 0.00060315 | 0.0007443
25 | NO | 4.9074 | 32.3672 | 0.00021483 | -0.0014038 | 0.0006031 | 0.0020069
26 | NO | 4.7851 | 32.1433 | 0.00048253 | -0.00063324 | 0.00060311 | 0.0012363
27 | NO | 5.3624 | 32.125 | 0.00034554 | -0.0032196 | 0.00060305 | 0.0038227
28 | NO | 5.1037 | 31.7113 | 0.00027006 | -0.00951 | 0.00060308 | 0.010113
29 | NO | 4.6364 | 30.6282 | 0.0002957 | -0.0082073 | 0.00060312 | 0.0088104
30 | NO | 49.1956 | 30.0898 | 0.0073299 | 0.011927 | 0.00059896 | -0.011328
30 | NO | 5.1323 | 29.6596 | 0.00618 | -0.012714 | 0.00060308 | 0.013317
31 | NO | 5.0944 | 25.9512 | 0.05348 | -0.80971 | 0.00060308 | 0.81032
32 | NO | 4.9965 | 27.6088 | 0.013793 | -0.20129 | 0.00060309 | 0.2019
33 | NO | 4.5327 | 27.8693 | 0.028466 | -0.10627 | 0.00060313 | 0.10687
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE