## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 26 March 2023 | Added HPLC S.H. |

## PROCESSING NOTES

Cruise: 2018-05

Agency: OSD

Location: Strait of Georgia / Vancouver Harbour

Project: Strait of Georgia / Vancouver Harbour Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: 2 February 2018 – 8 February 2018

Processed by: Germaine Gatien

Date of Processing: 12 February 2018 – 4 May 2018

Number of original HEX files:  	81	   	Number of CTD files: 80*

Number of bottle files: 	30			Number of bottle casts processed: 29 *

Number of original TSG files: missing		Number of processed TSG files:	 0

* 1bottle file created accidentally and 2 were combined in single file.

* 1 CTD file with only surface data not processed.

# INSTRUMENT SUMMARY

Two CTDs were used during this cruise:

- A Seabird 911+ CTD (#1222) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#3038), a SeaPoint Fluorometer (#3642), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (0691) and an altimeter (#62355). A Surface PAR (#20518) was mounted on the boat deck.

- A Seabird 25 CTD (0404) with an ECO Fluorometer (#4185 and a SBE 43 DO sensor (#1176) attached was used without a rosette.

A thermosalinograph (Seacat 21 S/N 3274) was mounted with a Wet Labs WETstar fluorometer (S/N SCF3654) and a remote temperature sensor.

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

Note that as of January 2017 a change has been made to the threshold levels for quality flags 3and 4 for Extracted Chlorophyll. Flag=3 will now be 10%CV (formerly 15%) and Flag=4 will now be 30% (formerly 50%).

The Daily Science Log, rosette log sheets and analysis logs were in good order with detailed comments on problems. There was no mention of the Thermosalinograph. The cruise report makes it clear that it was present but that there were problems in getting it working.

The deck pressure was recorded on 2 occasions. This was helpful in assessing the pressure sensor accuracy.

There were 3 casts using an SBE25 without a rosette. They were given non-standard names and contained no information to indicate station name or event number. The only information available was the internally recorded time which was used to determine the event number; the computer time in the SBE25 files appears to be in PDT rather than PST. To assist in converting such files it is helpful to have headers entered with station name, water depth, latitude and longitude. No configuration file was provided; one was constructed based on the sensors listed in the log although no pH sensor was listed.

The pumps were turned off too quickly after the final bottle was fired during cast #7; wait at least 5 seconds so that CTD data collected around firing time come from a time when the pumps are on.

For cast #7 conversion was rerun with a smaller window.

The BL file for cast #1 was empty. A file was constructed based on plots of pressure versus scan numbers and choosing scan numbers from about 20s after the CTD stopped. This enabled conversion and the descent rates in the converted data are close to zero so the choices look appropriate.

The primary temperature/salinity pair was selected for archiving for most casts. However, for cast #82 there were apparent problems in the flow in the primary system near the surface so the secondary sensors were chosen for that cast.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore or when there is a large range of CHL values. The comparison for this cruise showed most fluorescence values close to extracted chlorophyll.

The fluorescence from the last 2 casts in

Recalibration of the salinity and dissolved oxygen channels was based on the results of 2018-01 which immediately followed this cruise and had sampling in deep waters and in conditions likely to result in better flushing of Niskin bottles. The comparisons between bottles and CTD data for 2018-05 do show evidence of poor flushing. The recalibration of channels Oxygen:Dissolved:SBE and Salinity:T0:C0 applied the same corrections as for 2018-01 while the correction of Salinity:T1:S1 was

The CTD dissolved oxygen correction based on comparison to bottles is lower than found during cruise 2018-01 which immediately followed it. The later cruise had extensive sampling in waters where Niskin bottles are likely to have flushed quite well, so the DO correction used for 2018-01 was applied to these data. The salinity comparisons also show evidence of incomplete flushing of Niskin bottles, so the results of 2018-01 were applied to both salinity channels.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 50db

±0.2 mL/L from 50 to 150db

±0.06 mL/L from 150db to 250db

±0.02 mL/L below 250db

As of May 3, 2018 no data from the TSG were available. There were problems setting up the system during the cruise, so there may be no useful data.

# PROCESSING SUMMARY

PART I – CTD 911+

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and dissolved oxygen sensors was checked. This was the first use of those sensors since they were last serviced with the exception of the DO sensor was used on a camera with no calibration sampling during cruise 2017-36. The sensors were all used during 2018-01 with calibration sampling in deep offshore waters; the cruise occurred after 2018-01 but the bottle comparison was run first to provide information for this cruise. .

##### Hysteresis Tests

There was no deep sampling so no hysteresis study was done.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-05-ctd.xmlcon. Depth was included.

The BL file for cast #1 was empty. A file was constructed based on plots of pressure versus scan numbers and choosing scan numbers from about 20s after the CTD stopped. This enabled conversion and the descent rates in the converted data are close to zero so the choices look appropriate.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. No significant outliers were found.

A preliminary header check was run and it was found that the pumps were off for part of the surface bottle of file #7. The pumps were shut off immediately after firing – a wait of 5seconds is recommended. The cast was reconverted using a smaller window.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. A few adjustments had to be made to the sample numbers:

- Cast #75 – Bottle 11 did not fire.

- Cast #78 – accidental firing of one bottle – no further processing to be done.

- Cast #98 has a single bottle that should be added to cast #97. The files were renamed 97.bota and 97.botb and combined using JOIN to create a new file 2018-05-0097.bot. The file was opened in Ultraedit to change the bottle number for the final bottle from 1 to 7.

- Cast #104 & 106 - Sample numbers were out of order and had to be rearranged in the ADDSAMP file. Niskin #1 for cast 104 was not sampled so was dropped from ADDSAMP.

- Cast #110 – salinity sample labelled as 298 and the value suits that level but log sheet says it was to come from 297. Assume label is correct.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

For cast #104 Niskin #1 was removed from the SAMAVG file.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-05-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-05chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-05chl.csv. Loop data were saved in a separate file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-05oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2018-05oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2018-05SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 14 to 28 days after collection.

The files were simplified and saved as 2018-05sal.csv. File 2018-05sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-05_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-05-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a few discrepancies in CHL samples coming from a different depth than on the log but the labels match the spreadsheet entries and the values are reasonable. There were also problems in the bottle file for #106, both in the ADDSAMP file and in the nutrient file. The nutrient data for cast #106 were split in the analyst’s spreadsheet with data from 104 between them. This led to errors in the creation of the NUT file. These errors were fixed and the merges were repeated.

There were a number of discrepancies between the log sheets and the samples actually taken:

- There were no DO samples from cast 57 to match those indicated on the rosette log sheet; no raw data were found so it is presumed no samples were taken.

- There were many more samples for event 96 than are shown on the rosette log. The DO data from the bottles and CTD match reasonably well so the data do appear to be from that cast.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were 7 outliers, mostly from the top 50m.

- Cast 31 – The sample came from a high-gradient region so that the difference between bottle and CTD might be due to only a slight vertical offset between them. No change is needed.

- Cast #53 – Sample #122 could not be from the rosette entry for that number since the bottle did not close. The salinity and dissolved oxygen samples look like they came from Niskin 1 (sample #121). All samples were renamed as 121 with 3 flags. Compare was rerun after that correction and this sample was much closer to the CTD values, but still slightly out of line. The bottle salinity had a value that is lower than the CTD. This was a bottom bottle so this is likely due to incomplete flushing. The CTD salinity was also very noisy.

- Cast #57 – There is a lot of noise in the CTD salinity during the stop but most of the difference is likely due to incomplete flushing of the CTD. The bottle salinity can be found in the CTD profile about 5m below the firing level. No flag is suggested.

- Cast #97 - The CTD salinity is noisy and the values of both bottle and CTD can be found during the stop. No flag is suggested.

- Cast #104 - The CTD salinity is noisy and the values of both bottle and CTD can be found during the stop. No flag is suggested.

- Cast #110 & 111 - The bottles were fired at the bottom of the casts so poor flushing would produce bottle values lower than the CTD as is found here. The water would have to have come from about 6m higher in both cases. No flag is suggested.

The only flag changes recommended are to sample 121. The cases of incomplete flushing mean that the bottles do not compare well with the CTD salinity, but the values likely accurately reflect the contents of the bottles.

When the outliers above are excluded there remain 3 minor outliers in the top 100m that are likely due to incomplete flushing with bottle values higher than the CTD. Another case has the bottle value higher than the CTD but it was taken near the bottom where incomplete flushing leads to lower bottle values. Other bottles were excluded because the CTD data were noisy during the 10s window.

When those outliers are also excluded the primary salinity was found to be higher than the bottles by 0.0020psu and the secondary by 0.0030psu with standard deviations of 0.0032 and 0.0033 respectively.

The pressure and time dependence differ little between the two.

Note that the salinity sample from cast 106 was misplaced in the file used for this comparison, but that value was not included in the comparison, so the results are unaffected.

For full details for the COMPARE run see file 2018-05-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When this sensor was used during 2018-01, which followed this cruise, the near-surface waters were well mixed whereas for this cast there is often a large vertical gradient near the surface, especially in Indian Arm where the gradients at the surface were particularly high. Another difference is that flushing of Niskin bottles was likely good for most bottles during 2018-01, but not for this cruise during which descent rates were quite steady.

The results for 2018-01 were:

CTD DO Corrected = CTD DO * 1.0327 + 0.0627 (R2 = 0.97)

The trendline for 2018-05 when all data are included is:

CTD DO Corrected = CTD DO * 1.0345 + 0.0432   (R2 = 0.08)

That looks reasonably close to the 2018-01 but includes a lot of scatter and some extreme outliers. As outliers are gradually removed based on residuals the slope is reduced and the offset increased.

Generally we expect the CTD DO to read lower than bottles because calibration drift is almost always towards lower DO. If there is poor Niskin flushing then the bottles are likely to contain deeper water with lower DO. So it is not a surprise to find that the most significant outliers in the direction of bottles having lower DO than the CTD come mostly from Indian Arm and Vancouver Harbour where the descent rate of the CTD was very steady. Outliers in the opposite direction also come from Indian Arm but at the surface where the vertical gradients in DO were extremely high so that the vertical separation between the CTD and bottle became significant.

Bottles fired at the bottom will tend to have DO reading too high if flushing is poor. When bottom bottle differences were displayed they showed that most did fall below or within the general fit; none showed the CTD to be higher than bottles. When the fit of bottom bottles was forced to have the offset found during 2018-05 the only bottles that don’t come close to the fit are from Indian Arm, Vancouver Harbour and Saanich Inlet. The vertical DO gradient was quite high near the bottom for all those bottles so poor flushing will create larger errors than elsewhere.

The data from the central part of the Strait of Georgia were examined in the hopes that flushing was better there. However, the descent rates still look quite steady. There are fewer obvious cases of poor flushing but the fit does suggest that there are small effects.

CTD DO Corrected = CTD DO * 1.0148 + 0.1214   (R2 = 0.18)

For this program we usually have casts from near the mouth of Juan de Fuca Strait where the sea state ensures that Niskin bottles get flushed better than in the Strait of Georgia, but this cruise did not include that area.

The fit from 2018-01 is likely the best choice for recalibration of the CTD data. To check if this makes sense, a plot was produced with the 2018-01 trendline superimposed on the 2018-05 central Strait of Georgia data. It shows that for CTD DO values between 2 and 4mL/L the differences fall fairly close to that trendline. For most of the DO values from 5 to 7mL/L the differences are well above the trendline. This makes sense because the lower values generally come from depths where the vertical gradients are low, limiting errors due to incomplete flushing. The higher DO values come from depths where larger vertical gradients are found, so poor flushing creates a larger effect. This is not a simple thing to analyze, since the vertical gradients can be so high that the distance between the CTD and Niskin results in an error larger than and of opposite sign to the one due to poor flushing.

All outliers could be explained by high gradients and/or poor flushing. No flag changes were recommended.

It was impossible to assess which of the replicates was better in the 2 cases that were flagged as replicate outliers.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The ratio FL/CHL versus Extracted CHL looks typical of these sensors. It starts high for very low CHL and drops rapidly to about 1as CHL values rise. The ratio continues to drop slowly but is close to 1 for 0.3< CHL<1.3ug/L. There are no bottles with CHL>1.3ug/L.

For full details of the comparison see file 2018-05-fl-chl-comp1.xlsx.

The analysts were notified of problems that will affect their QF files.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-05-ctd.

The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

- The Dissolved Oxygen, pH,  transmissivity and fluorescence profiles look normal

- PAR and SPAR values were mostly very low but some casts have a larger range and look reasonable.

- The altimetry looks reasonable near the bottom though there are frequent spikes so values will need to be checked.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. None of the casts were deep enough to determine if there is any pressure dependence. The differences are small and similar to the shallowest comparisons made during 2018-01 which are shown highlighted below. The salinity differences are consistent with the differences found in the bottle comparison from 2018-01 when the secondary salinity was higher than the primary by an average of 0.0014psu.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-05-0046 | 380 | -0.0004 | +0.0001 | +0.0014 | High, steady |
| 2018-05-0067 | 330 | -0.0002 | +0.0001 | +0.0016 | “ |
| 2018-05-0089 | 400 | -0.0002 | +0.0001 | +0.0014 | “ |
| 2018-01-0026 | 1000 | -0.0005 | +0.0001 | +0.0012 | High, Mod |
| 2018-01-0032 | 1000 | -0.0002 | +0.0001 | +0.0011 | High, Noisy |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0072 | 1000 | -0.0005 | -0.0002 | -0.0023 | High, XNoisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. There was an error in station name for cast #38 and a number of station names had formats that didn’t match other casts so they were changed for consistency.

The header check was run and no problems were found.

The cruise track was plotted and added to the end of this report. (Note that the SBE25 casts were included.)

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±5m the readings were checked. Based on those checks the water depths were changed to match the log book entries for casts 12, 37, 47 and 53.

A surface check was run and shows an average surface pressure for the cruise was 1.5db which is reasonable for the Vector in calm waters.

During cast #1 the minimum pressure recorded was -0.227db. This occurred when the pumps were off at the end of the cast. The transmissivity drops to 0 when the pressure reaches -0.02db which suggests that the pressure is accurate. Deck measurements made during the cruise provide further confirmation of the pressure accuracy. During cast 47 the deck pressure was -0.03db and during cast 75 it varied from +0.01 to -0.04db.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space.

SHIFT was run twice on all SBE911 casts using -0.6 records for the primary conductivity and -0.5 records the secondary. Salinity was recalculated for both channels.

pH

Tests were run on 3 casts to determine the best setting to align pH with temperature. A setting of +50 records looked best. SHIFT was run on all SBE911casts using +50 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were two warnings about casts being less than 10db:

- Cast #98 was run just to fire another bottle to go with the previous cast. It will not be processed further.

- Cast #110 was in very shallow water so the warning can be ignored as no data were lost.

##### Other Comparisons

Previous experience with these sensors –

This was the first use for the pressure, temperature and conductivity sensors since they were last serviced. The dissolved oxygen sensor was used during 2017-36 but was on a CTD with a camera on a net and no calibration sampling.

The sensors were used during 2018-01 which was run immediately after this cruise. The primary salinity was found to be high by about 0.0019 and the secondary by 0.0038. The dissolved oxygen was recalibrated using slope 1.0327 and offset 0.0627

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only temperature data that fell outside the climatology were from the bottom of casts #7 and #10 (in the southern Strait of Georgia, near the Gulf Islands) which were slightly low at the very bottom of the cast. Salinity was also below the historic minimum at the bottom of the same two casts and also for cast #73 (in the north-west part of the Strait of Georgia.) There are few observations in this area in February. These slight excursions are not suggestive of calibration problems.

Repeat Casts – There were a few repeat casts:

- Casts 98 was a very shallow repeat of 97, so of no use for comparison

- Casts 38 and 90 were too far apart in time to test repeatability of instruments but does give some idea of the size of temporal variations. There was a lot of small-scale vertical variability in cast #38 and much smoother profiles for #90.

- Casts 81 and 82 were run using different CTDs so are useful in evaluating the calibration of the SBE25. The data from both casts were bin-averaged before editing and recalibration and compared in EXCEL. See 2018-05-sbe25-sbe911-comp.xls.

  - The Dissolved Oxygen data were lower than those from the 911 CTD. When the SBE911 DO values after recalibration were plotted against SBE25 values, the following fit was found:

SBE25 DO Corrected = SBE25 DO * 1.043 + 0.0812

The fit is poor only in the high gradient zone between 3 and 5mL/L.

The correction varies from 5.75% at DO 6mL/L to 8.6% at DO 2mL/L.

If the fit is forced through the origin it is:

SBE25 DO Corrected = SBE25 DO * 1.066

.

  - The SBE25 temperature was noisier than the primary but there was no consistent trend.

  - The ECO fluorescence from the SBE25 was lower than the SBE fluorescence from the SBE911, by about 0.05ug/L near the bottom and by 0.1ug/L at mid-depths, but in the top 25m the ECO reads higher by about 0.1ug/L. The value of the ECO fluorescence is ~0.04ug/L near the bottom of the cast while that of the SBE sensor is ~0.09ug/L. The SBE911 is closest to the extracted CHL samples, though it was only close for the surface bottle. This may partly due to poor flushing of Niskin bottles.

  - pH values were consistently lower by about 0.03 at depth and 0.055 above 100db.

  - Salinity was lower at the surface, higher at 10db and low by about 0.05 in the upper 100m and by 0.01below 100db. On a T-S surface the SBE25 salinity is lower than the CTD by about 0.015 between 130db and 200db and by 0.01psu below 200db.

For the fluorescence and pH it is not obvious that one sensor is better than the other. For most channels there is a lot of noise and we have only the one comparison available. The temperature looks good and the salinity differences may be at least partly due to real changes. The fact that the salinity looks lowest relative to the SBE911 in areas of high gradients suggests that differences are partly due to poorer conductivity response or poor alignment between temperature and conductivity. Only the DO comparison is strong enough to consider recalibration of the SBE25 values. The differences were consistent and the SBE911 DO calibration is pretty well understood from studies during 2018-01.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish the primary and secondary channels. The primary salinity was chosen for editing and eventual archiving because it was a little closer to the bottles and the secondary may have a few more spikes and unstable features.

However, cast #83 had poor primary salinity, likely due to a problem with flow to the conductivity sensor; the secondary data also had problems but looked better overall so T1/C1 were selected for editing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and bottom of cast that appear to be corrupted by shed wakes or ship effects. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some unstable features but those are expected in this region of active mixing. No further editing was applied.

##### Recalibration

The primary salinity was found to be higher than the bottle salinity by an average of 0.0020psu and the secondary was high by 0.0030. The standard deviations in the comparison are 0.0032 for both channels.

The comparison during 2018-01 found the primary salinity to be high by 0.0019psu and secondary by 0.0038. The standard deviation was 0.0008. We normally see a pretty good agreement between the differences between the primary and secondary comparisons and the differences between the 2 salinity channels during downcast as show in section10. The 2018-01 difference is 0.0019 which is a little higher than expected while 2018-05 is 0.0010 which is lower. A closer look at the particular bottles where the primary and secondary salinity differ most shows they have noisy CTD salinity with rapidly varying differences so that the exact moment of firing makes a difference. The bottles where the differences are closest to the downcast results show little noise. When only those bottles are included in the comparison the primary salinity is found to be high by 0.0025psu and the secondary by 0.0041psu.  This raises the differences for both channels but much more to the secondary, so it is likely that noise in the secondary channel during bottle stops may account for the poor match to the 2018-01 results. The primary comparisons agree reasonably well.

2018-01 had comparisons have lower standard deviations, bottles from waters with lower vertical salinity gradients and probably had better flushing of Niskin bottles, so results from that cruise will be applied to this one. The primary calibration looks quite secure and that will be used for all but one cast. The secondary may be slightly off, but will lead to better results than would be achieved by not recalibrating.

Dissolved oxygen recalibration was discussed in section 5.

File 2018-05-SBE911-recal.ccf was prepared to subtract 0.0019psu from channel Salinity:T0:C0 and 0.0038psu from the Salinity:T1:C1, to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0327 + 0.0627

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The primary salinity was found to be low by 0.0004psu and the secondary salinity was found to be low by 0.0008psu.

See file 2018-05-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen. When 3 extreme outliers are excluded the average of differences in the DO fit was +0.024mL/L and the standard deviation was 0.087mL/L. The differences near the surface are generally higher than bottles as is expected due to incomplete flushing of Niskin bottles. Most of the differences are small for DO<4mL/L.  The fit looks much better than before recalibration.

See file 2018-05-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.043mL/L (standard deviation of 0.049mL/L). This is expected due to the incomplete flushing of Niskin bottles. Near the surface the differences are all positive, while below 250m the DO values are generally lower than bottles; both of those results are consistent with incomplete flushing.

No further recalibration is justified. See 2018-05-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #83 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #83 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to remove the header entry TIME UNITS, fix the vessel name and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 50db

±0.2 mL/L from 50 to 150db

±0.06 mL/L from 150db to 250db

±0.02 mL/L below 250db

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on the processing see document: 2018-05_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found. This will be repeated after the SBE25 data are completed so all can be reported together.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~77% to 100% with the lowest values in Vancouver Harbour. Most values were between 92% and 100%. These values are low but most casts were very well mixed which does keep DO saturation down.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts except #83 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #83 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-05-bottles-final.xls. The entries were compared with the rosette log sheets and one problem was found in cast #106. There was an error in the merge process, which was corrected and all affected steps were repeated.

Standards check and a header check were run on all files. A few problems were found and resolved.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

PART II – SBE25

##### Seasave

This step was completed at sea. The names were non-standard.

##### Preliminary Steps

The cruise summary sheet was completed.

No configuration file was provided. A file was found with serial numbers that matched those recorded in the headers. These proved to be incorrect, producing bad conductivity values. The temperature and conductivity sensors recorded in the file headers were not the ones listed in the log.

A new configuration file was found with the appropriate parameters and it was saved as 2018-05-sbe25-ctd.xmlcon.

##### Conversion of Raw Data

Conversion would not work because the sensor numbers in the RAW file headers and configuration file did not match. The RAW files were copied to a different folder and the sensor serial numbers in the headers were edited.

All files were converted using configuration file 2018-05-sbe-ctd.xmlcon. They were renamed in standard format; times were used to match to log entries.

Plots were examined and all expected channels were present. The fluorescence values for the last two casts are very high and fairly constant. The only extracted chlorophyll sampling available anywhere close to this site comes from several days earlier and from sites that were not particularly well-mixed, but they show no evidence of high CHL. Since this sensor is not pumped the high values would not be due to clogging of the duct and something adhering to the sensor would not likely happen at 2 casts. It is most likely that the cap was left on.

The deployment method was to soak the CTD at about 10m for between 1.6 and 2.5 minutes.

The pressures where conductivity drops suddenly during upcasts ranged from -0.2 and -0.3db; ±1db is the accuracy quoted by SeaBird for the pressure sensor so this looks reasonable.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

There are frequent small steps in the pressure and some small reversals. Tests were run to see if just filtering pressure was sufficient to improve the salinity and it was not.

Tests were run to see whether the best choice is a Low-Pass Filter or a Windows Filter size 3, 5 or 7, choosing in each case to filter temperature and conductivity as well as pressure. The low-pass filter may over-smooth the pressure, but the results in T-S space look much improved. The Windows cosine filter size 7 looks even better while size 5 looks worse than the low-pass filter.

WFILTER was run with cosine file width 7.

##### ALIGNCTD

During 2016-40, 2017-41, 2017-52 and 2017-53 when this DO sensor was also used the DO channel was advanced by 2.5s. That setting looks reasonable for these data as well.

ALIGNCTD was run to apply a 2.5s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

Casts were examined after this step and the changes are small but generally upcast and downcast T-S traces become closer.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Station names, locations and water were added to the headers of the IOS Header files.

ADD TIME CHANNEL was run to add 7 hours to the casts.

CLEAN was run to add event numbers.

##### Checking Headers

The cross-reference check was run. The times are now correct.

CLIP was used to remove initial records from the soak period to ensure that data are only kept from after the pumps came on and the sensors were able to equilibrate. For cast #81 1200 scans were removed and for the other two casts 780 scans were removed.  Plots were examined and show the number of records removed were appropriate

Track plots and Header Check were run together with the SBE911 data and no problems were found.

Track plots were produced using event numbers and station names. They were added to the end of this report.

The surface check shows an average of 0.1db on files before CLIP was applied and 0.8db afterwards. The associated salinity values are extremely low before CLIP and much higher afterwards.

##### SHIFT

Fluorescence

The fluorometer on the SBE25 was not pumped so alignment is not usually run; plots confirm that it is unnecessary.

Conductivity

It has been found during other cruises using this equipment that no setting improved most casts. The casts from this cruise are not suitable to test this.

SHIFT was not applied to conductivity.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; the alignment looks very good for 2 of the 3 casts. The 3rd is difficult to judge but likely any adjustment would improve one area and make another worse.

SHIFT was not run on DO.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017 cruises using this equipment. A setting of +10 records looked appropriate for these data. SHIFT was run to advance the pH signal by 10 records.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and/or bottom of all 3 casts and to clean salinity for 1 cast. Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined and no problems were found.

##### Other calibration checks

Sensor History – The sensors were used on 6 previous cruises since they were last serviced. There was no calibration sampling for 5 of those cruise. During 2017-78 there was some surface sampling that suggested the SBE25 salinity was high but the comparison is not considered reliable. There was one inter-ship comparison in June 2017 that suggested that the salinity accuracy was good and that dissolved oxygen was probably reading low.

Comparison of repeat casts – For a comparison of casts 81 (SBE25) and 82 (SBE911), see section15 and document 2018-05-sbe25-sbe911-comp.xls. There may have been real changes between the two casts especially given that the comparison was based on pressure. Conclusions from that comparison were:

- Temperature data were very close with a median difference of 0.0004C°.

- pH and Fluorescence data were both slightly lower from the SB25 but there is no way to say which CTD was more accurate; the differences varied with pressure in a non-linear way that probably reflect the varying ability to resolve large gradients.

- Dissolved oxygen values from the SBE25 are clearly low. The comparison leads to a fit but there are insufficient data to trust the results except to say that the SBE25 values appear to be low by at least 5%. A larger correction is indicated by these data, but given too many uncertainties in the comparison a conservative correction of 5% looks reasonable.

- Salinity values from the SBE25 were lower by a median value of 0.009psu but the standard deviation was 0.06psu. The SBE25 sometimes read significantly higher or lower than the SBE911 which is likely due to poorer resolution in high vertical gradients.

Historic Ranges – All temperature and salinity data were within the climatology.

Post-cruise calibrations – None were available.

##### CALIBRATE

The only SBE25 variable for which we have sufficient evidence to justify recalibration is dissolved oxygen. It is likely low by at least 5%, so file 2018-05-sbe911-recal.ccf was prepared to multiply all DO values by 1.05.

##### Filter

The fluorescence channel was filtered using a median filter, size 5. This removed large spikes but a few slightly negative values remained at depth in one cast. The dark value of this sensor is so close to zero that instrumental noise sometimes appears as slightly negative values.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

After bin-averaging most of the negative fluorescence values had disappeared, but a few very small ones remained.

CLEAN was run to replace a few negative values in Fluorescence with pad values.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels. Channel Fluorescence:URU:Wetlabs was removed from casts #116 and 118.

Dissolved Oxygen was derived in mass units.

Saturation rates were calculated and plots show near-surface saturation varied from 87 to 90% which is in the same range as other casts from this cruise.

REORDER was used to get the 2 dissolved oxygen channels together.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal

and unedited except that some records were removed in editing temperature

and salinity.

Dissolved oxygen was recalibrated based on a comparison with a cast run at the same

site using a SBE911. DO values from the SBE25 were clearly low, as expected

from this type of sensor. The correction applied should be considered a

rough estimate.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Channel Fluorescence:URU:Wetlabs was removed from files 2018-05-0116 and 2018-05-0118

as values were very high throughout the water column; the cap was likely left

on the sensor.

For details on the processing see document: 2018-05_Processing_Report.doc.

The standards check was run and no problems were found.

A cross-reference listing was produced for the whole cruise.

A header check was run including files from both CTD s. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PART III - TSG

##### Thermosalinograph Data

No data files were available at the time of original processing of this cruise. It is known that there were problems setting up the instrument, and possibly no data were found to be useful. When the computer on which any data might be stored is accessible, data may be found.

Particulars

Test cast – Fluorometer on SBE911 not working – replaced.

Surface PAR mounted on boat deck rather than aft rail.

5. Bottle 6 spigot loose connection.

7. Very strong tidal flow, long lead aft on rosette line, could not get rosette deeper than 202m.

Bottle 5 tripped but did not seal - no data.

27. Top 0-rings replaced on bottles 15 and 16.

40. Bottle 1 was empty.

47. Deck Pressure reading -0.03db.

53. Went from 100 back down to 127 for bottle 7. Bottle 1 not fired. Two bottles fired at 125m but Niskin 6 not sampled.

62. Foggy.

65. Fog lifts.

75. Deck pressure 0.01 to -0.04db.

Bottle 11 did not fire – took HPLC from Bottle 9 which was the next closest extracted CHL sample.

78. 1 bottle fired accidentally – no bottle file needed.

98. Surface cast only - just to fire bottle 7 to go with event 97. No CTD or CHE cast prepared.

104/106. Sample numbers out of order.

116. SBE25 used off the side.

116-118. Looks like cap left on fluorometer.

CRUISE SUMMARY - CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1222 | Yes | Yes |
| 2 | SEABIRD | 25 | 404 | No | Yes |

| Calibration Information - 1222 |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sensor |  |  |  | Pre-Cruise |  |  | Post Cruise |  |  |  |  |
| Name |  | S/N |  | Date | Location |  | Date |  | Location |  |  |
| Temperature |  | 2106 |  | 18Feb2017 | Factory |  |  |  |  |  |  |
| Conductivity |  | 2280 |  | 23Feb2017 | Factory |  |  |  |  |  |  |
| Secondary Temp. |  | 2663 |  | 18Feb2017 | Factory |  |  |  |  |  |  |
| Secondary Cond. |  | 2754 |  | 22Feb2017 | Factory |  |  |  |  |  |  |
| Transmissometer |  | 1185DR |  | 7Aug2017 | Factory |  |  |  |  |  |  |
| SBE 43 DO sensor |  | 3038 |  | 24Mar2017 | Factory |  |  |  |  |  |  |
| PAR |  | 4565 |  | 16Mar2011 | Factory |  |  |  |  |  |  |
| SeaPoint Fluor. |  | 3640 |  |  |  |  |  |  |  |  |  |
| pH sensor |  | 691 |  | 15Jan2015 | Factory |  |  |  |  |  |  |
| Surface PAR |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |  |
| Pressure Sensor |  | 1222 |  | 03Mar2015 | Factory |  |  |  |  |  |  |
| Altimeter |  | 62355 |  | 21Mar2016 | Factory |  |  |  |  |  |  |
| Calibration Information - 0404 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  |  |  |  | Post Cruise |  |  |  |
| Name | S/N |  | Date |  |  | Location |  | Date |  | Location |  |
| Temperature | 2449 |  | 16Dec2015 |  |  | Factory |  |  |  |  |  |
| Conductivity | 1764 |  | 23Feb2017 |  |  | Factory |  |  |  |  |  |
| pH | 0852 |  | 8Mar2017 |  |  | Factory |  |  |  |  |  |
| ECO Fluorometer | 4185 |  | 7Mar2017 |  |  | Factory |  |  |  |  |  |
| SBE43 Oxygen | 1176 |  | 2Mar2017 |  |  | Factory |  |  |  |  |  |
| Pressure | 482 |  | 28Feb2017 |  |  | Factory |  |  |  |  |  |

NOTE: At the time of original processing no data from the TSG were available. This section is included in case any data are found. The TSG was installed but not working well, so there may not be any useful data.

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3274       Cruise ID#:	2018-05

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3274 | 4Mar17 | Factory |  |  |
| Conductivity | 3274 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | SCF3654 | May2017 |  |  |  |
| Temperature SBE38 | 271 | 3Dec15 |  |  |  |