## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2018-013

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, B.C.

Party Chief: Wyeth M.					Location: Strait of Georgia

Project: Hard Bottom Longline Hook Survey (South)	Platform: Neocaligus

Date: 7 August 2018 – 31 August 2018

Processed by: Germaine Gatien

Date of Processing: 17 June 2019 – 18 June 2019

Number of original HEX files:	45		Number of CTD files: 45

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 7580) was mounted with oxygen sensor #0249.

# SUMMARY OF QUALITY AND CONCERNS

There was a digital version of the logbook available and a spreadsheet with header information which included positions in the format required for easy addition to the headers. This was much appreciated as it saves time in processing. The file names were non-standard, but included information that made it easy to change them to standard format. The event numbers run from 1001 to 1045.

The CTD was soaked at about 1db for about 1 or 2 minutes after the pump came on.

There was no calibration sampling for either salinity or dissolved oxygen. The calibrations were from February 2017, so some calibration drift is likely.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were non-standard.

The raw files were copied to another folder where they were renamed.

##### Preliminary Steps

A digital version of the log book and a spreadsheet with positions and comments for casts were provided.

The event numbers go from 1001 to 1045.

The cruise summary sheet was completed.

A single configuration file used for all casts.

There were no parameters on file at IOS for the temperature, conductivity, dissolved oxygen and pressure calibrations, so the values could not be checked.

There was no history available for this CTD since the previous recalibration.

The configuration file was saved as file 2018-013-ctd.xmlcon.

##### Conversion of Raw Data

All casts were converted using file 2018-013-ctd.xmlcon.

The CTD was soaked around 1m for about 1 to 2 minutes after the pump came on.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the conductivity to force it to have the same response as the temperature. Pressure was smooth and did not require filtering.

##### ALIGNCTD

Alignctd was run to advance the DO voltage by +1.5 seconds which brought the DO voltage traces into better alignment with the temperature traces.  Temperature was advanced by +0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Plots of temperature, salinity and dissolved oxygen versus pressure were examined to see that data look reasonable and they do.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth. Spreadsheet “HBLL2018_CTD_summary_forIOS.xlsx” provided that information. One worksheet had the data with position formats suitable for IOS SHELL, so this step was easy. The file was saved as 2018-013_hdr_merge.csv.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH including the start time for each cast.

A cross-reference listing was produced and looks fine.

A preliminary header check was also run and no errors were found. The minimum pressure found in the data was -0.06db; that came from the beginning of a cast when conductivity was ~0. At the end of the same cast the minimum pressure was ~-0.01db with ~0 conductivity. The pressure calibration appears to be good given that the initial accuracy quoted for this type of pressure gauge is ±1db.

The surface check is not very useful for this type of sampling as there may be records that are in water but with the pump not on. The average surface pressure was +0.1db, but the associated salinity values were near zero. The one cast with “in water” salinity values had a surface value of about +0.2db, so the pressure looks ok.

Track plots look good and were added to the end of this report.

##### SHIFT

Conductivity

Small shifts to conductivity were tested to see if they improved the salinity, as judged by T-S plots. None of the settings appeared to improve the data, so no shift was applied.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. There are few distinct features to allow a good test, but there are no signs that further adjustment is needed.

##### DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

Plots were made to see what editing was needed and it is clear that some bad near-surface data remain.

CLIP was run to remove data with pressure <1.2db which come from the soak period or the very beginning of the cast, but even with that step some poor data remain near the surface.

All casts were examined in CTDEDIT and some editing was applied to all but 1 cast. This involved removing a few surface records or records corrupted by shed wakes near the bottom and/or light cleaning of salinity where unstable features appear to be due to minor misalignments between temperature and salinity.

All CLIP files were copied to EDT and then the CTDEDIT output files were copied to EDT to produce a complete set of files.

##### Inter-comparisons

Sensor History – The sensors have been recalibrated since they were last used.

Comparison of repeat casts –There were no repeat casts and nearby casts are too shallow to expect good repetition.

Historic Ranges – There is local climatology for most of the casts. Profile plots with the local climatology superimposed show all salinity data fell within the normal range except for low salinity values in the top 5m at 2 casts in Malaspina Strait. Temperatures were above the upper limit for parts of most casts in the Gulf Islands. It is not surprising since temperatures have been found to be higher in many areas in the past few years and the climatology derives from 1988 to 2008. Moreover, there were likely not many data available from the Gulf Islands for the climatology and the 3-standard-deviation climatology is considered too severe for near-shore casts. So these excursions are not considered indicative of a calibration problem. Temperatures were all within the climatology for all other casts.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling and no useful history to enable estimates of calibration drift. No recalibration was applied.

##### Bin Average and REMOVE

The files were bin-averaged using 1db bins.

T-S plots were examined. Some very small unstable features remain but may well be real.

CHANGE UNITS was run to derive dissolved oxygen in mass units.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

REORDER was run to get the 2 dissolved oxygen channels together.

ADD TIME CHANNEL was used to get the time in UTC by adding 7 hours.

##### Dissolve Oxygen Saturation

Plots were made to check the DO saturation rate at the surface. Values ranged between 75% and 130% with the most common values being between 100% and 120%.  This range is not surprising given the mix of vertically well-mixed casts in areas of strong tidal mixing where low saturation values are expected and other areas with higher vertical stratification so that higher DO saturation is likely. We can conclude little about the calibration of the DO sensor except to say that the range looks reasonable.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

The Oxygen:Dissolved:SBE channel is nominal and unedited except that some

records were removed in editing temperature and salinity.

There was no salinity or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

For details on the processing see the report: 2018-013-proc.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2018-013 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 7 August 2018 End: 31 August 2018 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Wyeth M. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 7580 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE19+/7580	Cruise ID#:		2018-013

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 7580 | 19Feb2017 | Factory |  |  |
| Conductivity | 7580 | 19Feb2017 | Factory |  |  |
| SBE43 Oxygen | 0249 | 16Feb2017 | Factory |  |  |
| Pressure | 7580 | 14Feb2017 | Factory |  |  |