## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 12 Oct.2018 | After initial processing it was discovered that a different TSG had been in use than was indicated in the con file and the data were not archived as the results were bad. The data were reprocessed and added to the archive. PART III of this report was rewritten. |

## PROCESSING NOTES

Cruise: 2018-029

Agency: OSD

Location: Strait of Georgia /Juan de Fuca Strait

Project: Strait of Georgia Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: 6 April 2018 – 13 April 2018

Processed by: Germaine Gatien

Date of Processing: 5 July 2018 – 19 July 2018

Number of original HEX files:  	84 	   Number of CTD files: 83 (81 SBE911, 2 SBE25)

Number of bottle files: 	24		   Number of bottle casts processed: 24 (1 with no chemistry)

Number of original TSG files: 3		   Number of processed TSG files: 3

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the primary pump, a SeaPoint Fluorometer (#3641) on the secondary pump, a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (0691) and an altimeter (#62355). A Surface PAR (#20518) was mounted on the boat deck.

A thermosalinograph (SeaCat 21 S/N 3363) was mounted with a remote temperature sensor.

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+ v5.1e.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log contained no information about the SBE25 that was used for 2 casts; no configuration file was provided. Otherwise the log and rosette sheets were in good order with detailed comments on problems.

Sample labels were put into the Daily Science Log Book in cases where samples were planned but accidentally not taken. In one case there was no note on the rosette sheet about this. The analysts need that information so they can make an entry in the analysis spreadsheet with a pad value, flag 9 and a comment to indicate what happened. It is not unknown for someone to be searching for a particular sample long after a cruise. A note in the log book is useful as well.

For more than half of the CTD casts the pumps were turned off less than 5seconds after the last bottle was fired. Waiting at least 5s ensures that pumped CTD data can be captured during a 10s-window centred on the firing times.

There were 2 casts using an SBE25 without a rosette. It was very helpful that the headers had station name, water depth, latitude and longitude entered so that the information was automatically entered into IOS headers when the data were converted to IOS Header format. No configuration file was provided; one from a previous cruise was used.

The ECO Fluorescence and SBE dissolved oxygen channels were removed from the SBE25 files.

- The fluorometer data were not usable as was also found when it was used during 2018-05. The sensor performed properly during one cruise between 2018-05 and 2018-029 so it is presumed that the problem arose on both cruises because the cap was left on the sensor.

- The SBE25 dissolved oxygen concentration and surface saturation are much lower than those from nearby casts using the SBE911 and lower than any other casts during the cruise, including some in well-mixed waters where lower values are expected. The same sensor was used in February and March of 2018 with no sign of a significant problem. Temperature and salinity values are in good agreement with nearby casts. There may have been a deployment problem such as a problem with the plumbing. Preliminary data from a June 2018 cruise suggest that the sensor was working properly then.

Cast #108 was a shallow cast to collect samples. The cast went to only 6db so no CTD file was prepared.

No sample data is available at this time, so the CHE file contains only CTD data collected during the bottle stops. Cast #130 has a full profile for this site, station BS5.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

Recalibration of channel Salinity:T1:C1 was based on a comparison with bottles and a preliminary examination of data from the cruise that followed this one, 2018-039, plus previous observations of the effect of incomplete flushing on Niskin bottles. The recalibration of channel Oxygen:Dissolved:SBE was based on the results of 2018-01 which had offshore sampling with a large range of DO values and in conditions likely to result in better flushing of Niskin bottles.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 50db

±0.2 mL/L from 50 to 100db

±0.15 mL/L from 100db to 300db

±0.03 mL/L below 300db

The TSG data are available upon request but at this time the temperature and salinity values are considered suspicious. This will be reviewed when the instrument is next calibrated at the factory or if information becomes available from another cruise.

# PROCESSING SUMMARY

PART I – SBE911

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors was checked. They were all used during just 1 cruise, 2018-01, since they were last calibrated.

- The calibration control file was checked and a small change was made to factor E in the dissolved oxygen sensor parameters, based on hysteresis tests during 2018-01.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-029-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. A few outliers were found in casts 9, 93 and 97, mostly in the secondary salinity. CTDEDIT was used to clean the salinity data. The output files were copied to *.BOT.

A preliminary header check was run and a few problems were found:

- There was no station name entered for cast #1, so that was added.

- For more than half of the stations the pumps were turned off less than 5seconds after the last bottle was fired. Please wait at least 5s so pumped CTD data can be captured during a 10s-window centred on the firing times. The unpumped values were removed from each BOT file using a text editor. CLEAN was run to update the channel limits.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-029-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-029chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-029chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-029oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2018-029oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2018-029SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 8 to 11 days after collection. The files were simplified and saved as 2018-029sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-029_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-029-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Two problems were found:

- Event #59 which was called #58 on the rosette sheet. The CHL data had been identified correctly but not the other samples.

- Event #73 was missing DO sample #166 – this sample was on the DO analysis log but was entered with the wrong sample #.

- Event #108 has no sample numbers and currently contains no samples, but a CHE file will be produced because Dr. Sutherland did take samples and needs the CTD data to go with them.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There are 4 outliers that were removed from the comparison. After that exclusion the fit for the primary salinity was fairly flat with the CTD salinity low by an average of 0.0011psu and standard deviation of 0.003psu. The secondary has more pressure dependence with salinity low by an average of 0.0064psu and a standard deviation of 0.005psu.

When bottles above 125db are excluded the trendline for the primary remains the same while the secondary becomes much flatter with pressure. We expect better results at depth because higher vertical gradients near the surface make the effects of incomplete flushing more significant. The primary is low by the same average of 0.0011psu with a standard deviation of 0.003psu and the secondary is low by 0.0055psu with standard deviation of 0.005psu.

The analyst was informed about 4 large outliers:

- Cast 1 –The CTD data show little variability during the 10s window but just as the CTD stopped the data do look noisy. So the bottle analysis is likely fine. No flag suggested.

- Cast #22 – The CTD salinity is variable throughout the stop so the difference of 0.1psu from bottle likely reflects real change through the stop. No flag suggested.

- Cast #66 – There was quite a lot of noise in the CTD salinity. The bottle value looks closer to the 300db data. Choosing Niskin #3 is unusual, so maybe there was a mis-sample. However, that is not clear. The sample may be from the wrong bottle but the evidence is not clear. Analyst added flag 4 and comment.

- Cast #93 – This time the 300db sample looks more like the 250db data. There is no evidence of a misfire in other variables. The CTD data was not noisy during the stop. If we assume it is really from 250m the CTD is lower than the bottle by 0.0015psu whereas if it is assumed to be from 300db the CTD is high by 0.0659psu. This looks like a mis-sample. Analyst added flag 4 and comment.

There were no casts near the mouth of Juan de Fuca Strait where the flushing of bottles is generally best. The most westerly casts in the Strait were:

- Event #18 at station 72. The standard deviation was low in the CTD data during the stop. The primary salinity was high by 0.00005psu and the secondary was low by 0.0038psu.

- Event #13 at station 69 also has quiet CTD salinity; the primary was low by 0.0009psu and the secondary was low by 0.0045psu.

Those 2 casts show the CTD being slightly closer to bottles than the average of the whole cast. This may be reflective of better flushing of bottles. The Strait of Georgia bottles are likely to contain higher salinity water than in situ values. The differences between the bottles shows that secondary to be lower than the primary by between 0.002 and 0.004psu with the largest differences seen where the samples are a little deeper, so that may be due to some pressure dependence. There are not enough data to conclude that and the standard deviations are large.

Using the 5 deepest bottles (excluding outliers) the primary is low by 0.0011 and the secondary by 0.0057psu, with the secondary lower than the primary by ~0.0046. None of those bottles comes from Juan de Fuca Strait.

It is a little puzzling that the primary salinity seems flatter in shallow water than the secondary, but a closer look shows that one bottle at 100db has a large influence on the fits; the standard deviation in the CTD data is very high for that bottle. Excluding that bottle would make the secondary fit a little flatter and the primary less so. So the fit below 125m is a better choice.

For full details for the COMPARE run see file 2018-029-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When this sensor was used during 2018-01 the near-surface waters were well mixed whereas for this cruise there are often large vertical gradients near the surface. Another difference is that flushing of Niskin bottles was likely good for most bottles during 2018-01, but not for this cruise during which descent rates were quite steady.

The results for 2018-01 (Line P) were:

CTD DO Corrected = CTD DO * 1.0283 + 0.0491 (R2 = 0.95)

The trendline for 2018-029 when all casts are included and outliers removed based on residuals is:

CTD DO Corrected = CTD DO * 1.0117 + 0.0812   (R2 = 0.28)

When only the casts in Juan de Fuca are included and 2 outliers removed the fit is:

CTD DO Corrected = CTD DO * 1.0285 + 0.0030   (R2 = 0.69)

When 7 casts with descent rates indicative of conditions likely to improve Niskin flushing were included and outliers excluded based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0263 + 0.0198   (R2 = 0.66)

In areas where flushing is not efficient the Niskin contents may come from lower in the water column where DO values are lower. Since the DO sensor does read low, in general, as a result of calibration drift, the flushing error offsets the calibration error. In the comparison with bottles this might lead to the DO sensor looking low by less than it really is, or even make it appear to be reading high. So the results from the different fits above are consistent with previous results in more protected waters. The Juan de Fuca results are reasonably close to those from Line P which were based on offshore observations where bottle flushing is expected to be quite good.

The fit from 2018-01 is likely the best choice for recalibration of the CTD data. To check if this makes sense, a plot was produced with the 2018-01 trendline superimposed on the 2018-029. It shows that for CTD DO values between 2 and 3.5mL/L the differences fall fairly close to that trendline. For most of the DO values from 5 to 7mL/L the differences are above the trendline. This makes sense because the lower values generally come from depths where the vertical gradients are low, limiting errors due to incomplete flushing. The higher DO values come from depths where larger vertical gradients are found, so poor flushing has a larger effect. This is a very rough analysis -the vertical gradients can be so high that the distance between the CTD and Niskin results in an error larger than and of opposite sign to the one due to poor flushing and there were many DO reversals to complicate interpretation.

All outliers could be explained by high gradients and/or poor flushing. No flag changes were recommended. The values flagged by the analyst looked only slightly out of line; no changes to flags are recommended.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises. However, the fluorescence values are lower relative to Extracted CHL than usually seen. Similar results were found last April in the same area and using a different fluorometer. When the CHL and Fluorescence are plotted against file pair number, it is clear that the ratio is higher in Juan de Fuca Strait and low in the Strait of Georgia. In the casts closest to the mouth of Juan de Fuca the ratio is close to 1.

For full details of the comparison see file 2018-029-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-029-ctd.

The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

- There was a lot of noise in the temperature, especially the primary, but it is mostly in the upcast.

- The Dissolved Oxygen, pH,  transmissivity, fluorescence, PAR and SPAR profiles look normal

- The altimetry looks reasonable near the bottom though there are frequent spikes so values will need to be checked.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. That setting was used during 2018-01.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 4 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. None of the casts were deep enough to determine if there is any pressure dependence. The differences in temperature and conductivity are small but the combined effect makes for a significant difference in salinity. The salinity differences are consistent with the differences found in the bottle comparison from Juan de Fuca (~0.004psu) and in the deeper waters of the Strait of Georgia (~0.0045psu).

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0059 | 300 | ~0 | -0.0004 | -0.004 | High, V.Steady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The SBE25 casts were processed up to the point of header checks so they could all be run together.

The header check was run and the only problem found was that there was no station name for event #1. That was added to the IOS and CLN files.

A cross-reference list was checked against the log book. The only problem found was in Event #1. It looks like the computer was turned on more than 10 minutes before acquisition started. The positions and times from the log book were used to replace the headers in the file.

The cruise track was plotted and added to the end of this report including the SBE25 casts.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. Based on those checks the water depths were changed to match the log book entries for casts 12, 48 and 128. The altimetry header was removed from cast #108 since it was only a surface cast and never got close to the bottom. The header entry for event 136 was changed because the data near the bottom was very noisy; a plot was examined and extrapolation to the bottom shows that 4m is a reasonable value. For the bottle files only cast #108 needed to be edited.

A surface check was run and shows an average surface pressure for the cruise was 1.7db which is reasonable for the Vector.

There was a problem with the wire coming off the sheave during cast #5, so there was a slow recovery. The CTD was held between +0.2db and -0.3db for about 45s with acquisition on and pumps running. About 1s after the values became negative the conductivity fell to near-zero. When pressure was ~ -0.2db the transmissivity was ~0%/m but then increased. So the pressure appears to be within -0.2db. Given that spray may affect readings even when the CTD is slightly out of the water, we cannot deduce that it is reading slightly low. The pressure is likely accurate to within +/- 0.2db.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space.

SHIFT was run twice on all SBE911 casts using -0.5 records for the primary conductivity and -0.4 records the secondary. Salinity was recalculated for both channels.

pH

Tests were run on 3 casts to determine the best setting to align pH with temperature. A setting of +50 records looked best. SHIFT was run on all SBE911casts using +50 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was that pressure was <10db for cast #108. That was a surface-only cast.

For cast #140 the pumps were not on for the first drop, so the CTD was lowered again with pumps on. It was not clear that a new file had not been started, but that was the case. So the data were removed from the shfpH file for which the pumps were off, DELETE was rerun. The results were much better.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 2018-01. The primary salinity was found to be high by ~0.0017 and the secondary low by ~0.0007. The dissolved oxygen was recalibrated using slope 1.0283 and offset 0.0491

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Cast #27 had salinity above the climatology maximum between 8 and 30db. It also has unstable features in T-S space. Cast #128 has spikes at the bottom of the cast; the CTD touched bottom during this cast, and the near-bottom T-S data are unstable. All other T and S data fell within the climatology.

Repeat Casts – The only repeat cast was too shallow to be useful.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The choice of which T-S pair to choose for editing and eventually archiving is not obvious. The primary salinity appears to be closer to bottle values, though the effect of incomplete flushing reduces confidence in that result. The primary salinity is likely slightly high and the secondary, low; the secondary conductivity may have drifted more than the primary. However, the primary conductivity has many more spikes than the secondary data so they will not require as much editing. The secondary was also chosen when these sensors were used during 2018-01.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and bottom of casts that appear to be corrupted by shed wakes or ship effects. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some unstable features but those are expected in this region of active mixing. No further editing was applied.

##### Recalibration

The secondary salinity was found to be low by between 0.004 and 0.007psu depending on what casts were used in the comparison. This is a marked change from cruise 2018-01 in February when the channel was found to be low by 0.0007psu. Differences in shallower water are often larger, which is likely due to higher vertical gradients increasing errors due to incomplete flushing of Niskin bottles. However, the differences between the two sensors had also increased since 2018-01 from about 0.001 to 0.004psu. Complicating the issue was that there appeared to be some pressure dependence in the primary salinity during 2018-01 so that it was reading higher at 300m than at 1000m whereas the secondary was quite flat with pressure. So it is impossible to judge whether one sensor has drifted significantly or both sensors have drifted in opposite directions.

From the two casts in the middle of Juan de Fuca Strait (the most westerly but not near the mouth) the secondary salinity is low by about 0.004psu. During a cruise in 2017-63 it was found that the bottle comparison from western Juan de Fuca showed a sensor pair producing salinity higher by about 0.002psu than from other parts of the cruise. So it seems likely that flushing errors affect the comparisons by at least 0.002psu. That would suggest the salinity correction should be between +0.002 and +0.005psu.

Using the 5 deepest bottles (excluding outliers) the primary is low by 0.0011 and the secondary by 0.0057psu, with the secondary lower than the primary by ~0.0046. None of those bottles comes from Juan de Fuca Strait. The differences between bottles is consistent with the findings in section 9

A final check was to do look at preliminary data from cruise 2018-039 which used the same equipment in May 2018. Checks were made of 19 bottles from 4 deep offshore casts where flushing is likely to be good. The samples were from 500db or below. The secondary salinity was found to be low by median values of -0.0039, -0.0018, -0.0039 and -0.0034psu for the 4 casts. Using the whole set the median difference was -0.0036psu. The CTD data during bottle stops were averaged but outliers had not been removed first, nor were standard deviations examined. Also there is some pressure dependence and not all samples came from the same depths, so this is a rough comparison.  A quick check of the primary salinity shows it to be high by 0.0008psu which implies the secondary salinity is lower than the primary by 0.0044psu. The median difference between channels shows the secondary lower by 0.0038psu. So the difference between channels is roughly consistent with the findings in section 9. The secondary salinity appears to be reading low by ~0.0036psu in May 2018.

See 2018-39-test-sal-comp.csv.

Dissolved oxygen recalibration was discussed in section 4. The results of cruise 2018-01 were considered the most reliable information available.

File 2018-029-SBE911-recal.ccf was prepared to add 0.0035psu to channel Salinity:T1:C1, to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0283 + 0.0491

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The secondary salinity was found to be low by 0.002psu, as expected. See file 2018-029-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen. When data from the top5m and Saanich Inlet are removed plus outliers based on residuals the average of differences in the DO fit was +0.033mL/L and the standard deviation was 0.021mL/L. While the CTD appears to be reading high, this is expected due to incomplete flushing of Niskin bottles. The differences near the surface are generally higher than in deeper waters due to higher near-surface gradients combined with incomplete flushing of Niskin bottles. The fit looks much better than before recalibration. See file 2018-029-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.068mL/L (standard deviation of 0.043mL/L). This is expected due to the incomplete flushing of Niskin bottles. The differences show less depth dependence than is usually seen which may be because of the complex vertical gradients. For the bottles fried at the bottom of casts, the DO was high by 0.003mL/L with a standard deviation of 0.141mL/L. At the bottom of casts any effect of poor flushing is expected to make the CTD look lower than bottles, though low gradients at the bottom usually lead to such errors being small. As usual, Saanich Inlet comparison is out of line with all others, having a very steep DO gradient. No further recalibration is justified. See 2018-029-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

The secondary salinity recalibration was partly based on a comparison with bottles,

taking into account conditions likely to result in incomplete flushing of Niskin

bottles. Also considered was the preliminary comparison with bottles from an

offshore cruise that followed and used the same equipment.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 50db

±0.2 mL/L from 50 to 100db

±0.15 mL/L from 100db to 300db

±0.03 mL/L below 300db

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on the processing see document: 2018-029_Processing_Report.doc.

At this point cast #108 was dropped from the files to be archived as it contains only data to 6m. The bottle file is of interest to Dr. Sutherland, but there is already a full profile at this site (#130) so the CTD file is of little use.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found. This will be repeated after the SBE25 data are completed so all can be reported together.

Profile and T-S plots were examined and no problems were found. There are some unstable features in T-S space where waters are well-mixed, but those are very small and possibly real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~75% to 125% with the exception of Saanich Inlet where it was ~160%. The lowest values were in Juan de Fuca Strait, Haro Strait and other well-mixed sites, while values were highest in more open areas of the Strait of Georgia. The values look reasonable for the area and season.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-029-bottles-final.xls. The entries were compared with the rosette log sheets. A few problems were found:

- 9 flags should have been added to the nutrient data from samples 126 and 128; they had already been added to the dissolved oxygen channel. They were added to the CHE file and the comment was amended.

- The chlorophyll data for event #66 were missing because they had been entered as event 60. Processing steps were repeated for that cast and the CHL spreadsheet was amended.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

PART II – SBE25

##### Seasave

This step was completed at sea. There were only 2 files; the names were changed to the new standard format with cruise number changed from 29 to 029.

##### Preliminary Steps

The cruise summary sheet was completed.

No configuration file was provided. A file was found with serial numbers that matched those recorded in the headers and it was saved as 2018-029-sbe25-ctd.xmlcon.

##### Conversion of Raw Data

All files were converted using configuration file 2018-029-sbe25-ctd.xmlcon.

Plots were examined and all expected channels were present. But the fluorescence signal looks bad. This fluorometer also produced unbelievable values during 2 casts of 2018-05 in February, but good values in March during 2018-96. It was presumed that the cap was left on during the 2 casts in the February cruise and that is probably what happened for these 2 casts as well.

The deployment method was to lower the CTD to about 10m and return to the surface, a process taking about 1.6 minutes.

The pressure at which conductivity drops suddenly during upcasts was about -0.8db; ±1db is the accuracy quoted by SeaBird for the pressure sensor so this looks reasonable. With only 2 casts to judge by, recalibration is not justified.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

There are frequent small steps in the pressure and some small reversals. Tests were run during 2018-05 to see if just filtering pressure was sufficient to improve the salinity and it was not. A variety of filters were tested and the best choice was found to be a Windows cosine filter size 7.

WFILTER was run on pressure, depth, temperature and conductivity with cosine file width 7.

##### ALIGNCTD

During 2016-40, 2017-41, 2017-52, 2017-53 and 2018-05 when this DO sensor was also used the DO channel was advanced by 2.5s. That setting looks reasonable for these data as well.

ALIGNCTD was run to apply a 2.5s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on both casts using the default setting: (α, 1/β) = (0.04, 8).

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Station names, locations and water were included in the CNV file headers and were automatically added to the IOS Header files upon conversion.

Times were in PDT, so ADD TIME CHANNEL was run to add 7 hours to the casts.

CLEAN was run to add event numbers.

##### Checking Headers

The cross-reference check was run. The times are now correct. No other problems were found.

CLIP was used to remove 800 initial records from the soak period to ensure that data are only kept from after the pumps came on and the sensors were able to equilibrate.  Plots were examined and show the number of records removed were appropriate

Track plot and Header Check were run together with the SBE911 data and no problems were found.

The surface check shows an average of -0.4db on files before CLIP was applied and 0.7db afterwards. The associated salinity values are extremely low before CLIP and much higher afterwards.

##### SHIFT

Fluorescence

The fluorometer on the SBE25 was not pumped so alignment was not needed.

Conductivity

It has been found during other cruises using this equipment that no setting improved most casts. The casts from this cruise are not suitable to test this. SHIFT was not applied to conductivity.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; the alignment looks good. SHIFT was not run on DO.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017 cruises using this equipment. A setting of +10 records looked appropriate for these data. SHIFT was run to advance the pH signal by 10 records.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and bottom of one cast and a few records corrupted by shed wakes from the other; salinity was cleaned lightly in both casts. Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined and no problems were found.

##### Other calibration checks

Sensor History – The sensors were used on 7 previous cruises since they were last serviced. There was no calibration sampling for 6 of those cruise. During 2017-78 there was some surface sampling that suggested the SBE25 salinity was high but the comparison is not considered reliable. There was one inter-ship comparison in June 2017 that suggested that the salinity accuracy was good and that dissolved oxygen was probably reading low.

Comparison of repeat casts – None were available.

Historic Ranges – All temperature and salinity data were within the climatology.

Post-cruise calibrations – None were available.

Inter-comparison of the 2 CTDs – There was no inter-comparison during this cruise, but during 2018-05 there were 2 casts at one station using the same SBE25 CTD used for this cruise. While it was compared with a different SBE911+ the results are included as they give provide some assessment of the SBE25 data.

  - The SBE25 temperature was noisier than the primary but there was no consistent trend in the differences.

  - pH values were consistently lower, by about 0.03 at depth and 0.055 above 100db. There is no calibration information to determine which sensor is more accurate.

  - Salinity was lower at the surface, higher at 10db and low by about 0.05 in the upper 100m and by 0.01below 100db. On a T-S surface the SBE25 salinity is lower than the CTD by about 0.015 between 130db and 200db and by 0.01psu below 200db. The salinity differences may be at least partly due to real changes. The fact that the salinity looks lowest relative to the SBE911 in areas of high gradients suggests that differences are partly due to poorer conductivity response or poor alignment between temperature and conductivity.

  - Only the DO comparison is strong enough to consider recalibration of the SBE25 values. The differences were consistent and the SBE911 DO calibration is pretty well understood from studies during 2018-01.

##### CALIBRATE

The only SBE25 variable for which we have sufficient evidence to justify recalibration is dissolved oxygen. It is likely low by at least 5%, so file 2018-029-sbe911-recal.ccf was prepared to multiply all DO values by 1.05.

##### Filter

This step was skipped since the fluorescence data are bad.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Fluorescence:URU:Wetlabs , Oxygen:Voltage, Oxygen:Dissolved:SBE Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units.

Saturation rates were calculated and plots show near-surface saturation at about 46% and 49% whereas those from nearby casts using the SBE911 varied from 82 to 87%. So the dissolved oxygen looks very low. This is an area where vertical mixing could draw oxygen down but even in Haro Strait values are well above these. The same sensor was used in February and March of 2018 with no sign of a significant problem. Temperature and salinity values are in good agreement with nearby casts. There may have been a deployment problem such as a problem with the plumbing connection. Examination of a cast from June 2018 suggests that the values were reasonable then, but that is a very rough judgment. Given no way to establish that the values are reliable and a very strong suggestion that they are not, a return was made to the REMOVE step and channel Oxygen:Dissolved:SBE was removed as well.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

Conductivity, Transmissivity and pH:SBE data are nominal and unedited except that

some records were removed in editing temperature and salinity.

Channel Fluorescence:URU:Wetlabs was removed from files 2018-029-0028 and 2018-029-0029

as values were very high throughout the water column; the cap was likely left

on the sensor.

Channel Oxygen:Dissolved:SBE was removed from files 2018-029-0028 and 2018-029-0029

because values were very low and out of line with values observed at nearby casts

where data were collected using an SBE911.

For details on the processing see document: 2018-029_Processing_Report.doc.

CLEAN was run on cast #108 to remove the Sample_Number channel as it contains only pad values.

Header Edit was rerun on that file and a note of explanation was added.

The standards check was run and no problems were found.

A cross-reference listing was produced for the whole cruise. A problem was noted in the Baynes Sound casts which had been given station names that match the log book but are not the standard ones used for this program, nor do they match the ones used for the GRAB events. For example Stn 5 should be BS5. These were fixed in the CTD files for events 128 to 154 and for the CHE file for event #108.

A header check was run including files from both CTDs. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PART III – TSG – THIS PART WAS REPLACED IN OCTOBER 2018

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

The 3 configuration files were identical but they are for a different TSG than was actually in use during this cruise. File SBE21_2488-4mar17.xmlcon was obtained with the proper calibration parameters.

b.) Conversion of Files

In order to convert the files the serial numbers had to be changed in the HDR and HEX files

The 3 hex files were converted using file SBE21_2488-4mar17.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The first file is about 40 minutes long but the traces look fine. The second file contains a section of very noisy salinity between casts 28 and 29. The weather was rough at that time so this may be due to bubbles. The third file looks ok. There was no TSG data for the Baynes Sound section of the cruise.

The track plot looks fine.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 2db, and then again using 3db. These were exported to 2 spreadsheets which were added to file 2018-029-ctd-tsg-comp.xls. There were 40 casts with data at 2db which overlapped with TSG files. Using 3db produced 49 matches.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to as 2018-029-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no position differences >0.0014º and the median differences were both <0.0001º. So the TSG clock worked well.

d.) Comparison of T and S from TSG and CTD data

The differences between the TSG and CTD were large especially in the central Strait of Georgia where the surface gradients were generally higher. The plot below is based on the differences at 3db; the differences at 2db were similar. Using all casts the temperature is high by a median of 0.135C°and the salinity is low by a median of 0.307psu, with standard deviations of 0.11C° and 0.90psu.

When outliers are removed based on either TSG temperature or salinity having high standard deviations over the 2minutes used to derive the median, the temperature is high by a median of 0. 133C° and salinity is low by a median of 0.300psu. The standard deviations are lower at 0.024C° for temperature and much lower for salinity at 0.031psu.

(See 2018-029-ctd-tsg-comp.xls.)

Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the 5th use since then. All the cruises were on the Tully with a longer loop than the Vector.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-06: The temperature in the lab was higher than CTD temperatures by from 0.15Cº to 0.20Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-23: The temperature in the lab was higher than CTD temperatures by from 0.16Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.004Cº. The TSG Salinity read lower than the CTD salinity by ~0.29psu and that appeared to be due to bubbles. There were no loop samples. No recalibration was applied to the TSG data.

2017-08: The temperature in the lab was higher than CTD temperatures by a median of about 0.018Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-09: The temperature in the lab was higher than CTD temperatures by a median of about 0.028Cº. The salinity traces was full of spikes; salinity was low by a median value of .0064psu or by 0.018psu if only the 10 casts with the lowest standard deviations in the TSG salinity were included.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by about 0.13Cº using both 2db and 3db CTD data.  In previous cruises using this equipment the lab temperature was high by about 0.18Cº, but they were all cruises on the Tully with a longer loop so that more heating is expected.

3. The TSG Salinity is lower than that from the 2db and 3db CTD data by about 0.30psu. Bubbles in the loop can sometimes cause salinity errors though there is no evidence of large bubbles during this cruise except during one noisy patch.

4. There are few good calibration checks for this TSG while on the Vector. Looking at 2 other TSGs that have been used on the Vector shows that the temperature in the lab has been higher than CTDs by between 0.1 and 0.2Cº with 0.13Cº being the value noted most often, so the value found for this cruise looks reasonable. For salinity the results have been highly variable with values being low by between 0.04psu and 0.93psu with the largest differences associated with the noisiest comparisons.

5. The temperature seems to be reliable and subtracting 0.13Cº would give a reasonable substitute for intake temperature.

6. The salinity values are too low. This could be due to calibration drift, bubbles in the water or the loop drawing water from very close to the surface, above the minimum depth in our CTD data. The 3rd explanation seems unlikely as temperatures would probably also be out of line if this were the case. The variability for those cruises using TSG #3363 shows no temporal dependence, so these is no suggestion of calibration drift. That makes bubbles seem a more likely explanation. If that is the case then it is surprising that the effect is quite consistent except for a few noisy patches. So the bubbles are likely due ship effects or the nature of the flow in the loop rather than oceanic conditions which would be more variable.

7. A proxy for intake temperature will be created by subtracting 0.13 Cº to the lab temperature.

8. Salinity will be recalibrated based on the comparison with CTD data by subtracting 0.30psu, but a note will be added to the header that this is a rough estimate.

9. Temperature will be reported with only 3 significant figures and salinity with only 2 significant figures as confidence in the data does not justify the usual 4 figures. This may draw users attention to the warning in the headers.

f.) Editing

No editing was required.

g.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2018-029-tsg-recal1.ccf to subtract 0.13Cº to Temperature:Primary and to add 0.30psu from Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference, Flag and Position:New channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH, to add the depth of sampling to the header, to remove entries for station and event number, to change channel names to standard names and formats (except for temperature and salinity as noted in the comments below) and to add the following comments:

TSG Data Processing

-------------------

There was no loop sampling, intake temperature sensor or flow rate meter.

A comparison of temperature and salinity was made with co-incident CTD casts.

The temperature in the lab was found to be higher than the CTD casts by a median

value of 0.13C degrees. A proxy intake temperature was created by adding 0.13 C

degrees from the lab temperature.

The TSG salinity was lower than that from the CTD salinity with a median difference

of about 0.30psu compared to CTD data from 2m and 3m. This difference may be due

to small bubbles in the loop.

Salinity:T0:C0: Salinity was recalibrated based on the comparison with CTD data.

Temperature:Primary - The temperature data from the lab were recalibrated based

on comparison with CTD data to create a proxy for intake temperature.

Temperature:Lab - These are uncorrected temperature data from the TSG.

See document 2018-029-Processing_Report.doc for more details.

-----------------------------------------------------------------------------------------------------------------------

WARNING: THESE DATA SHOULD BE USED WITH CAUTION AS THE CALIBRATION OF THE

TEMPERATURE AND CONDUCTIVITY SENSORS IS BASED ON A COMPARISON WITH CTD

CASTS AND THERE WAS A LOT OF SCATTER IN THAT COMPARISON.

DATA ARE REPORTED WITH FEWER SIGNIFICANT FIGURES THAN USUAL TO REFLECT

UNCERTAINTIES IN CALIBRATION.

THE PERFORMANCE OF THE TSG MAY DIFFER WHEN THE SHIP IS IN MOTION.

-----------------------------------------------------------------------------------------------------------------------

The files were saved as 2018-029-000*.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars

Surface PAR mounted on boat deck rather than aft rail

5. Wire off sheave at 51m – recover very slowly, then redid the cast

24-27. Chains on

28. SBE25 – fluorometer cap left on

30. pH cap off for 4+ hours

64. Interesting oxygen, temp signal @50m

70. pH cap left off

128. Touched bottom (editing applied in processing to remove effects)

130 – 138. Log note “CTD only” but entered as ROS so presume CTD was mounted in the rosette

CRUISE SUMMARY - CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  |  | CTD Calibration Sheet Competed? |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0443 |  | Yes |  |  | Yes |  |  |  |  |  |
| 2 | SEABIRD | 25 |  |  |  | Yes |  |  | No |  |  |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  | Post Cruise |  |  |  |  |
| Name |  |  | S/N |  | Date |  | Location |  |  | Date |  | Location |  |  |
| Temperature |  |  | 2710 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Conductivity |  |  | 2128 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Temp. |  |  | 4752 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Cond. |  |  | 2399 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 7Aug2017 |  | Factory |  |  |  |  |  |  |  |
| SBE 43 DO sensor |  |  | 0997 |  | 2Sep2017 |  | Factory |  |  |  |  |  |  |  |
| PAR |  |  | 4565 |  | 16Mar2011 |  | Factory |  |  |  |  |  |  |  |
| SeaPoint Fluor. |  |  | 3641 |  |  |  |  |  |  |  |  |  |  |  |
| pH sensor |  |  | 692 |  | 2Mar2018 |  | Factory |  |  |  |  |  |  |  |
| Surface PAR |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 31Aug2017 |  | Factory |  |  |  |  |  |  |  |
| Altimeter |  |  | 62355 |  | 21Mar2016 |  | Factory |  |  |  |  |  |  |  |
| Calibration Information - 0404 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  |  | Post Cruise |  |  |  |
| Name |  | S/N |  |  | Date |  |  | Location |  |  | Date |  | Location |  |
| Temperature |  | 2449 |  |  | 16Dec2015 |  |  | Factory |  |  |  |  |  |  |
| Conductivity |  | 1764 |  |  | 23Feb2017 |  |  | Factory |  |  |  |  |  |  |
| pH |  | 0852 |  |  | 8Mar2017 |  |  | Factory |  |  |  |  |  |  |
| ECO Fluorometer |  | 4185 |  |  | 7Mar2017 |  |  | Factory |  |  |  |  |  |  |
| SBE43 Oxygen |  | 1176 |  |  | 2Mar2017 |  |  | Factory |  |  |  |  |  |  |
| Pressure |  | 482 |  |  | 28Feb2017 |  |  | Factory |  |  |  |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2018-030

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |