## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | ### Updated channel names & formats in TOB files. G.G. |
| ### 26 March 2023 | ### Added HPLC. S.H. |
| 12June2020 | Sample #s 186-188 changed to 9186 to 9188 in file 2018-030-0059.CHE |

## PROCESSING NOTES

Cruise: 2018-030

Agency: OSD

Location: Strait of Georgia /Juan de Fuca Strait

Project: Strait of Georgia Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: 11 June 2018 – 16 June 2018

Processed by: Germaine Gatien

Date of Processing: 13 August 2018 – 15 October 2018

Number of original HEX files:  	88(1 split) 	Number of CTD files:  87

Number of bottle files: 	28 (1split)	   	Number of bottle casts processed: 27

Number of original TSG files: 	1	   	Number of processed TSG files: 1

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#953DR), a SBE 43 DO sensor (#3038), a SeaPoint Fluorometer (#3685), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (0692) and an altimeter. A Surface PAR (#20518) was mounted on the boat deck.

A thermosalinograph (SeaCat 21 S/N 3363) was used.

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+ v5.1e.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log contained no information about the equipment used and personnel; the usual page for that information was missing. TSG data were initially missing but 1 file was found on a shipboard computer. It only covered the first 24 hours of the cruise. A note in the log indicates that a file was started about 1.5 hours after the end of the 1st file but no other data have been found.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

Primary temperature and salinity were chosen for all casts except cast #98 during which the secondary data were bad below 320m on the downcast. The problem did not affect the bottle file.

Event #60 was run just to fire a 3rd bottle that was missed in the previous event. The two files were combined and named 2018-030-0059.CHE. Event #59 contained the full downcast.

The primary and secondary salinity channels were found to be low by 0.003psu and 0.002psu respectively, using data from below 50m and excluding 2 outliers. When data from 4 casts near the mouth of Juan de Fuca Strait are used to minimize errors due to incomplete flushing of Niskin bottles,  the salinity channels are low by an average of 0.002psu and 0.001psu. Those values are considered more reliable. They may overcorrect if there remain some errors due to flushing, but that error is likely to be small, so the corrected salinity would still be within ±0.002psu.

Similarly, the SBE DO data were recalibrated based on samples collected from the 2 westernmost casts in Juan de Fuca Strait to minimize the effects of incomplete flushing of Niskin bottles.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 50db

±0.15 mL/L from 75 to 175db

±0.06 mL/L below 175db

There was only 1 TSG file found covering only 24 hours of the cruise, and only a few CTD casts were available for comparison. Recalibration was based on the results of cruise 2018-029. The data are not considered as reliable as usual so users are warned to use the data with caution and values are reported with fewer significant figures than usual.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the dissolved oxygen sensor was obtained. The conductivity and dissolved oxygen sensors had not been used for any other cruises on record since they were last calibrated.

- The calibration control file was the same throughout the cruise. No errors were found in the calibration parameters. The file from cast #2 was saved as 2018-030-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-030-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. A few outliers were found in casts #34 and 128 in the primary salinity. CTDEDIT was used to clean the salinity data. The output files were copied to *.BOT.

There was also a need to join files 59 and 60. First the BOT files were joined and a text editor was used to fix the bottle numbers. The ADDSAMP file needed adjusting as well.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-030-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-030chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-030chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-030oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2018-030oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2018-030SAL.xlsx; there were no duplicates. The analyses were carried out in a temperature-controlled lab 3 to 9 days after collection. The files were simplified and saved as 2018-030sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-030_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-030-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

For cast #126 there was a repeated sample number that had to be changed from 291 to 9291, so the order of the SAL, NUTS and OXY files had to be reordered before the merges would work.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Two problems were found:

- Event #4 had all bottles fired but only Niskin 1 to 14 had sampling. The lines for Niskin 15 to 24 were removed from the SAMAVG file and the merge process was rerun.

- Event #59 was missing the samples from Niskin #3. The merge process was rerun.

- Event #86 had samples not indicated on the rosette sheet. The analysis spreadsheets were examined and confirm there were many more samples taken than indicated. It looks like only duplicates were entered for Oxy and Nut, and no CHL was entered. No change is needed.

- Event #126 was missing because the last step had been missed, so it was rerun for that cast.

The export step was rerun and no further problems were found.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The scatter in the fit was very high. The 7 bottles in the top 30m all had high salinity values compared to the CTD salinity which is to be expected when flushing is incomplete in the presence of high gradients. Next bottles that were fired at the bottom of casts were examined. Of these, 2 are outliers suggesting poor flushing, while 2 look ok. The latter 2 came from casts with very low gradients near the bottom.

The primary salinity was found to be low by an average of 0.003psu but the standard deviation is 0.007psu. The secondary was low by an average of 0.002psu with a standard deviation of 0.006psu.

When only the 4 casts nearest the mouth of Juan de Fuca Strait are included the primary is low by 0.002psu and the secondary by 0.001 with standard deviations of 0.0008psu and 0.0007psu, respectively. This is consistent with better flushing of Niskin bottles due to more vertical motion during casts.

Outliers are all explained by poor flushing and/or noisy CTD salinity.

The only pattern noted in plots versus time is due to the shallowest casts being at the end of the cruise.

For full details for the COMPARE run see file 2018-030-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When this sensor was used during 2018-01 the near-surface waters were well mixed whereas for this cruise there are often large vertical gradients near the surface. Another difference is that flushing of Niskin bottles was likely good for most bottles during 2018-01, but not for this cruise during which descent rates were quite steady. However, there were not many bottles using this sensor during 2018-01.

The results for 2018-01 (Line P) in February were:

CTD DO Corrected = CTD DO * 1.0327 + 0.0627 (R2 = 0.97)

The trendline for 2018-030 when all casts are included and outliers are removed based on high standard deviations in the CTD DO and on residuals is:

CTD DO Corrected = CTD DO * 1.0312 + 0.0626   (R2 = 0.64)

When only the 4 westernmost casts in Juan de Fuca are included and 4 outliers removed the fit is:

CTD DO Corrected = CTD DO * 1.0436 + 0.0295   (R2 = 0.89)

The 2 westernmost casts in Juan de Fuca have the noisiest descent; when only those are included and 5 near-surface points removed the fit is:

CTD DO Corrected = CTD DO * 1.0445 + 0.0245   (R2 = 0.92)

Since the descent rate is quite steady for most of this cruise we expect flushing to be poor, but in Juan de Fuca, especially near the mouth, the descent rate was noisy so bottles are likely to be flushed more efficiently. In areas where flushing is not efficient the Niskin contents may come from lower in the water column where DO values are lower. Since the DO sensor does read low, in general, as a result of calibration drift, the flushing error offsets the calibration error. The larger correction suggested by the Juan de Fuca results is likely more reliable.

The significant outliers were examined. Most are explained by high standard deviations in the CTD DO. A few more came from near the surface where gradients were high and sometimes with reversed sign, and the deepest sample from Saanich Inlet where the gradient was fairly high. No flag changes were recommended.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises. However, the fluorescence values are lower relative to Extracted CHL than usually seen, but similar results were found in April 2018 in the same area and using a different fluorometer. When the CHL and Fluorescence are plotted against file pair number, it is clear that the ratio is higher in Juan de Fuca Strait and low in the Strait of Georgia. In the casts closest to the mouth of Juan de Fuca the ratio is close to 1. This reflects low CHL in Juan de Fuca and high CHL in the Strait of Georgia.

For full details of the comparison see file 2018-030-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-030-ctd.

The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. As usual the T, C and Sal channel pairs track well on downcasts but differ significantly during upcasts. There were a few spikes in the secondary conductivity in one cast. The Dissolved Oxygen, pH, altimetry, transmissivity, fluorescence, PAR and SPAR profiles look normal

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. That setting was used during other recent uses of this sensor.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 4 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. None of the casts were deep enough to determine if there is any pressure dependence. The salinity differences are consistent with the differences found in the bottle comparison. The temperatures are farther apart than in the 2 previous cruises and conductivity and salinity channels are closer. The data in shaded cells are from previous cruises using the same sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0059 | 300 | ~0 | -0.0004 | -0.004 | High, V.Steady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |
| 2018-030-0050 | 300 | -0.0008 | +0.00003 | +0.0009 | High, Steady |
| 2018-030-0056 | 300 | -0.0010 | -0.00002 | +0.0007 | F.High, V.Noisy |
| 2018-030-0086 | 300 | -0.0013 | -0.00003 | +0.0008 | F.High, Noisy |

The differences are small though there are changes in both temperature and conductivity that combine to give lower salinity differences.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run and no problems were found.

A cross-reference list was checked against the log book. The only problem found was in the format of 2 station names. Those were fixed. There were a few discrepancies in dates or times, but in all cases the files look correct and the log records were clearly wrong.

The cruise track was plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. Based on those checks the water depths were changed in the CLN files to match the log book entries for casts #5, 7, 23, 32, 33, 35, 36, 47, 53, 58, 77, 79, 86, 87, 92, 145, 146 and 150. Two of those changes were applied to the SAMAVG files as well. Plots were examined for other casts but the altimetry header looked appropriate. There was likely some shoaling during casts and small differences in water depth can account for the check values.

The MERGE process was repeated for the 2 SAMAVG files affected.

A surface check was run and shows an average surface pressure for the cruise was 2.6db which is a little higher than usual for the Vector. The lowest pressure measured was 0.7db at the end of cast #95. The transmissivity fell to 0 at about the pressure but then increased to about 15%/m. Salinity was about 23psu and fluorescence about 1.5ug/L, but the pumps were off at the time. These values are reasonable for near the surface. There is insufficient evidence to say whether the values are a little low, but the sensor had bene serviced fairly recently, so it is unlikely there had been significant calibration drift.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

pH

Tests were run on 5 casts to determine the best setting to align pH with temperature. A setting of +50 records looked best, as it did for the same sensor during 2018-029. SHIFT was run on all SBE911casts using +50 records.

Conductivity

Tests were run on 5 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. The results varied from cast to cast with some cases of any setting making some near-surface features worse, but those were associated with reversals in temperature that look suspicious.

SHIFT was run twice on all SBE911 casts using -0.6 records for the primary conductivity and -0.5 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The oxygen sensor was used during 2017-36, 2018-05 and 2018-01, but only the last of those cruises provided a good comparison. The pressure and conductivity sensors have not been used since they were last calibrated at the factory.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There were a few minor excursions with bottom waters at cast #10 and bottom water at cast #58 being warmer and fresher, while cast #82 was a little warmer near the bottom. Excursions at the top and bottom are small and not unusual and do not suggest a calibration problem.

Repeat Casts – There were no repeat casts and nearby casts are too shallow to provide a reasonable comparison.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary salinity appears to be closer to bottle values, but there are a few spikes in the secondary T-S plots that are not seen in the primary. The primary channels were also chosen for editing except for cast #98. During that cast the primary T and S were bad from 320m to the bottom of the downcast plus below 240m of the upcast. This bad data does not affect the CHE file because all sampling was shallow.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. Most of the data removed are from near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing.

After editing T-S plots were examined for all casts. There are some unstable features but those are expected in this region of active mixing. No further editing was applied.

##### Recalibration

The most reliable dissolved oxygen data appear to be those from the 2 casts nearest the mouth of Juan de Fuca where flushing of bottles appears to be best.  The results using the 4 casts nearest the mouth produce a similar result but the fit is a little tighter and flushing likely is better using 2 casts only.

The primary and secondary salinity channels were found to be low by 0.003psu and 0.002psu respectively, using data from below 50m and excluding 2 outliers. When data from 4 casts near the mouth of Juan de Fuca Strait are used the 2 channels are low by an average of 0.002psu and 0.001psu. Those values are considered more reliable. There are not enough salinity samples to make a reasonable comparison based on just the 2 casts nearest the mouth. Using +0.002 and +0.001psu may overcorrect if there remain some errors due to flushing, but that error is likely to be small, and the corrected salinity is likely to be within ±0.002psu.

File 2018-030-SBE911-recal.ccf was prepared to add 0.002psu to channel Salinity:T0:C0 and 0.001psu to channel Salinity:T1:C1and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0445 + 0.0245   (R2 = 0.92)

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The primary salinity was found to be low by an average of 0.0002psu and the secondary salinity low by 0.00002psu. See file 2018-030-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen. Using the same data as for the fit the average difference shows the CTD to be higher than the bottles by 0.003mL/L where it was lower by 0.150mL/L before recalibration.  The fit looks excellent between 50m and 200m; CTD values are slightly high near the surface and slightly low below 200m. The differences near the surface are generally higher than in deeper waters due to higher near-surface gradients combined in the presence of incomplete flushing of Niskin bottles. The fit is much better than before recalibration. See file 2018-030-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.04mL/L (standard deviation of 0.06L/L). This is expected due to the incomplete flushing of Niskin bottles. For the bottles fired at the bottom of casts, the DO was low by 0.01mL/L with a standard deviation of 0.06mL/L. At the bottom of casts any effect of poor flushing is expected to make the CTD look lower than bottles, though low gradients at the bottom usually lead to such errors being small. No further recalibration is justified. See 2018-030-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #98 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #98 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 50db

±0.15 mL/L from 75 to 175db

±0.06 mL/L below 175db

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on the processing see document: 2018-030_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S space where waters are well-mixed, but those are very small and possibly real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~70% to 145%. The highest values were in Baynes Sound and the lowest in Haro Strait and eastern Juan de Fuca Strait and other well-mixed sites. Values in the more open areas of the Strait of Georgia were between 90% and 110%. The values look reasonable for this area where variability is usually very high.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts except #98 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #98 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-030-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

The configuration file used at sea was for the wrong TSG.

An initial processing and comparison with CTD data was similar to results during 2018-029 when lab temperatures were much lower than expected and salinity much higher than CTD salinity. It was discovered that the TSG on the Vector was #2488, not #3363.

A new configuration file was prepared.

b.) Conversion of Files

The wrong serial number was in the HDR and HEX files so that had to be changed to enable conversion.

There was only 1hex file. It was converted using the new file 2018-030-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. The traces look fine with no spiking in salinity.

The track plot looks fine but covers only the first day of the cruise.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 2.5db. These were exported to a spreadsheet which was saved as 2018-030-ctd-tsg-comp.xls. There were only 8 casts with data that overlapped with TSG files.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to 2018-030-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no position differences >0.0008º and the median differences were both 0.0001º. So the TSG clock worked well.

d.) Comparison of T and S from TSG and CTD data

The TSG temperature was found to be higher than the CTD temperature by a median value of 0.28C° (standard deviation 0.18C°) and salinity was low by a median of 0.06psu (standard deviation 0.07psu).

There are no useful data from 2m during downcasts. There are some shallower data from upcasts but they tend to be corrupted by shed wakes. In a few cases the data from upcasts looks more reliable and there is some suggestion from cast #5 and from later casts that did not overlap with the TSG data that the temperature might be higher at 2m and salinity lower so that differences might be lower for temperature and higher for salinity.

See 2018-030-ctd-tsg-comp.xls.

Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and have been used for 6 other cruises.

The temperature and conductivity sensors were recalibrated in March 2017 and this was the fourth use since then. All of these cruises were on the Tully except for 2018-029.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-06: The temperature in the lab was higher than CTD temperatures by from 0.15Cº to 0.20Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-23: The temperature in the lab was higher than CTD temperatures by from 0.16Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.004Cº. The TSG Salinity read lower than the CTD salinity by ~0.29psu and that appeared to be due to bubbles. There were no loop samples. No recalibration was applied to the TSG data.

2017-08: The temperature in the lab was higher than CTD temperatures by a median of about 0.018Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

2017-09: The temperature in the lab was higher than CTD temperatures by a median of about 0.028Cº. The salinity traces was full of spikes; salinity was low by a median value of .0064psu or by 0.018psu if only the 10 casts with the lowest standard deviations in the TSG salinity were included.

2017-029 (Vector): The temperature in the lab was higher than CTD temperatures by a median of about 0.0135Cº. The TSG salinity was lower than the CTD by a median value of 0.307psu or by 0.300psu if 12 casts with the highest standard deviations in the TSG temperature or salinity were excluded.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by about 0.27Cº. We would expect it to be reading higher by about 0.1 to 0.2 Cº, but the scatter in the comparison and the few data available make this comparison weak. We have no downcast data from 2m for the casts included in the comparison. Upcast data are mostly corrupted by shed wakes from below, but there are was one cast hinting that temperatures are higher at 2m than at 2.5m by enough to bring differences more into line with the results of cruise 2018-029. From some casts that do not intersect with TSG records, there is further evidence that temperatures at 2m are higher.

3. The TSG Salinity is lower than that from the 2.5db CTD data by about 0.05psu. During 2018-029 it was much lower than that. A possible explanation for that is that bubbles in the water were a bigger issue for the earlier cruise but there was no obvious sign of that in time-series plots. Large scatter in the few data available for 2018-030 make this comparison of limited value.

4. For the 2018-029 TSG data there were more data from near 2m and a comparison between the results at 2m and 3m showed little difference implying low average vertical gradients. Profiles showed a mix of results near 2m with some fairly high vertical gradients but with frequent reversals. So with 40 casts available and a mix of gradients these results are likely more reliable than the 2018-030 comparison.

5. The 2018-030 results are out of line with the history of this instrument. That may be due to having to use 2.5m-data and having so few data in the comparison. It is best to use the results of cruise 2018-029 with a warning that this is a rough estimate.

6. A proxy for intake temperature will be created by subtracting 0.13 Cº to the lab temperature based on the results of 2018-029 and the salinity will be recalibrated by subtracting 0.30psu based on the results of 2018-029.

f.) Editing

No editing was required.

g.) Recalibration

Add Channels was used to create Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2018-030-tsg-recal1.ccf to subtract 0.13Cº to Temperature:Primary and to add 0.30psu from Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan Number, Temperature:Difference, Flag and Position:New channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

Sampling System

---------------

The seawater is drawn from a 5cm standpipe which is flush with the hull

and located 2.0 metres below the water line. There is 1.6m of 5cm pipe

from the hull to the Moyno 2L6 SSQ pump and the water is then pumped

approximately 26 metres through a 3.8cm stainless steel pipe up to the

main lab. The system pressure is regulated by a 3.8cm Singer 106RPS valve.

In the lab, the seawater flows directly through the Seabird model 21

Sea-cat thermosalinograph and the exhaust goes directly overboard.

TSG Data Processing

-------------------

There was no loop sampling, intake temperature sensor or flow rate meter.

A comparison of temperature and salinity was made with co-incident CTD casts,

of which there were only 8. There was a lot of scatter in the comparison. When

a few outliers were excluded the temperature in the lab was found to be higher

than the CTD casts by a median value of 0.27C degrees and the TSG salinity was

lower than CTD salinity by a median value of 0.054psu. We expect the lab

temperature to be higher than CTD data but not by >0.2 C degrees.

There were no downcast CTD data from 2m in the comparison and there is some

evidence that the vertical gradients were significant. During the previous cruise,

2018-029, with many more CTD casts in the comparison and more data available

from close to 2m, no significant difference was found between the comparisons

at 2m and 3m. So the results of 2018-029 were used to recalibrate these data.

Salinity:T0:C0 - TSG salinity data recalibrated by adding 0.30psu.

Temperature:Primary - Proxy for intake temperature created by subtracting 0.13 C

degrees from the temperature recorded by the TSG in the lab.

Temperature:Lab - These are uncorrected temperature data from the TSG.

See document 2018-030-Processing_Report.doc for more details.

----------------------------------------------------------------------------------

WARNING: THESE DATA SHOULD BE USED WITH CAUTION AS THE CALIBRATION OF THE

TEMPERATURE AND CONDUCTIVITY SENSORS IS BASED ON A COMPARISON WITH CTD CASTS FROM ANOTHER CRUISE.

DATA ARE REPORTED WITH FEWER SIGNIFICANT FIGURES THAN USUAL TO REFLECT UNCERTAINTY IN CALIBRATION.

THE PERFORMANCE OF THE TSG MAY DIFFER WHEN THE SHIP IS IN MOTION.

----------------------------------------------------------------------------------

The file was saved as 2018-030-0001.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars

4. Bottles 15-22 fired for testing only – not needed in rosette file.

16. TSG display blank – no signal. Reset and worked for about 30 minutes. Restarted ~01:00 13 June.

35. Depth change during downcast (335 to 319).

36. Depth change during downcast (147-160).

43. No hdr, bl, xmlcon files – hex file converted properly to CNV. No bottles fired, so no problem.

60. Just run to trip bottle 3 for event 59. JOIN.

75. CTD cable twisted.

98. Feature in fluorescence signal at 240m on upcast. Primary channels bad near bottom of cast.

128. Block moved so brought CTD back to 10 off bottom. Although altimeter read 5m off bottom, rosette looks clean on recovery.

CRUISE SUMMARY - CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 953DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  | 3038 |  | 24Mar2017 |  | Factory |  |  |  |  |
| PAR |  |  | 4565 |  | 16Mar2011 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3685 |  |  |  |  |  |  |  |  |
| pH sensor |  |  | 692 |  | 2Mar2018 |  | Factory |  |  |  |  |
| Surface PAR |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | ? |  | ? |  | Factory |  |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2018-030

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |