## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 March 2023 | ### Added HPLC. S.H. |
| 22 March 2021 | Made corrections to NUTS in bottle files 2018-032-0008 and 2018-032-0053. S.H. |

## PROCESSING NOTES

Cruise: 2018-032

Agency: OSD

Location: Strait of Georgia /Juan de Fuca Strait

Project: Strait of Georgia Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: 21 November 2018 – 26 November 2018

Processed by: Germaine Gatien

Date of Processing: 11 March 2019 – 3 April 2019

Number of original HEX files:  	68		Number of CTD files:  68

Number of bottle files: 	26	   		Number of bottle casts processed: 26

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#953DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (3685), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (692) and an altimeter (43281). A Biosperical/Licor Surface PAR (#20518) was mounted on the boat deck.

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was JN937-DHT411-7KWHR.

The deck unit was a Seabird model 11+.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) Kit #2.

An IOS rosette with 24 10L bottles was used.

There was no thermosalinograph in use during this cruise.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and Sampling logs were in good order with details about problems encountered.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available. The fluorescence values look lower compared to extracted chlorophyll than usual, but it is   hard to judge this since most chlorophyll values were very low. When the fluorometer plumbing was cleaned in spring 2019, a small obstruction at the intake was found.

There was a problem with the salinometer used to analyze samples for this cruise. Tests were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”.  They established a linear correction to be applied to all salinity samples analyzed on salinometer #68572 between January 2019 and March 2019. Based on those corrected bottle values and the history of the sensors, the CTD salinity values are believed to be within +/-0.004psu.

The SBE DO data were recalibrated based on samples collected from the 4 westernmost casts in Juan de Fuca Strait to minimize the effects of incomplete flushing of Niskin bottles.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.35 mL/L from 0 to 50db

±0.15 mL/L from 50 to 100db

±0.10 mL/L from 100 to 200db

±0.03 mL/L below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained and were in excellent order.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the dissolved oxygen , temperature, conductivity and pressure sensors was obtained.

- The configuration files used at sea had no errors. It was saved as 2018-032-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-032-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. A few outliers were found in casts 25, 28 and 35 so those were opened in CTDEDIT. Light editing was applied as follows:

- Cast 25 – Secondary salinity at 40m

- Cast 28 – Secondary salinity at 40m

- Cast 35 – Primary salinity at 225m

The output files were copied to BOT files so there is a complete set of BOT files, edited or not.

A preliminary header check was run and fluorescence in Saanich Inlet was near the maximum so checks will be made later to see if it went off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

A few problems were encountered and fixed in the addsamp file:

- Cast #75 is shown to have 12 bottles fired on the log sheet, but Niskin #12 is not found in the CTD ROS file; that bottle was not fired and there was no sampling from that bottle so it was dropped from the ADDSAMP file.

- The dissolved oxygen samples from cast #82 were miss-named as #81.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

Bottles 10-24 were removed from the SAMAVG file for event #1 as there was no sampling from those bottles.

Niskin #17 was removed from file 2018-032-0082.samavg. No sample number was assigned to that bottle – all samples were identified as being from Niskin #16 at the same depth.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-032-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-032chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-032chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-032oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2018-032oxy.csv. That file was converted into individual *.OXY files.

Some bottles were flagged due to problems that will affect other samples as well, so after the bottle samples are merged the relevant flags will be added to the other samples.

SALINITY

Salinity analysis was obtained in file 2018-032SAL.xlsx; there were no duplicates. The analyses were carried out in a temperature-controlled lab 66 to 71 days after collection. The files were simplified and saved as 2018-032sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-032_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-032-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a few chlorophyll samples that were from a different bottle than shown on the sampling log, but the depths are the intended ones, and the planned bottle hadn’t been fired.

At this point the flags that the oxygen analyst indicated should apply to all samples were changed in the other sample files and then the MERGE process was rerun.

##### Compare

Salinity

Compare was run with pressure as reference channel.

Problems with the salinometer were detected when COMPARE was run during 2018-031 processing. The differences were very large and roughly inversely proportional to salinity. A quick check of this cruise showed similar problems even though the CTD sensors were different. Cruise 2019-003 was also examined and the comparison was similar. Salinity samples from 2018-031 and 2018-032 were analyzed in late January 2019 using Autosal #68572 and those from 2018-003 were run in late February 2019.

It appeared that the Autosal was producing poor results in early 2019. It had recently been serviced at the factory. Tests were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”. They established a linear correction to be applied to all salinity analyzed on salinometer #68572 between January and March 2019. The Autosal was returned to the factory in mid-March 2019. The correction was found to be:

Salinity_Corrected = 1.01135 * Salinity measured -0.3976

File 2018-032-recal-sal-bot was prepared and used to apply that correction to the MRGCLN2 files with output MRGCORBOT and COMPARE was rerun. There is a lot of scatter in the result with an unusually large standard deviation in the CTD data. But when bottles are excluded that have standard deviation in the CTD salinity >0.0008psu and depth of sample <150db, the average shows the primary salinity to be low by 0.007psu and the secondary low by 0.006psu. The difference between the comparison for primary and secondary sensors is consistent with the bottle differences noted in section 9. These values are lower than those found during the last use of these sensors. That difference may be due to calibration drift, evaporation/adsorption of samples during the 2-month wait for analysis, incomplete flushing of bottles and errors in determining the Autosal correction.

The most satisfying result is that the plot of differences versus time shows no obvious trend.

Since these data were part of the evidence to support the correction, it is not wise to recalibrate based on these results. The results do suggest that the bottle correction was appropriate.

For full details for the COMPARE run see file 2018-032-sal-comp1.xls and salinometer-calibration-study.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

This was the first use of this sensor since it was last calibrated at the factory.

The initial fit using all casts but excluding outliers based on residuals was:

CTD DO Corrected = CTD DO * 1.0309 + 0.0219 (R2 = 0.81)

To minimize the effect of poor flushing of Niskin bottles, the 4 bottle casts nearest the mouth of Juan de Fuca Strait were chosen since they have noisier descent rates. The fit was:

CTD DO Corrected = CTD DO * 1.0324 + 0.0096 (R2 = 0.89)

The significant outliers were examined. Most are explained by high standard deviations in the CTD DO. Most were near the surface and associated with fairly high standard deviation in the CTD DO. One came from the bottom of Saanich Inlet. That cast is generally not used in comparisons due to very high vertical DO gradients and the sensor’s slow response in hypoxic conditions. No further quality flags are justified.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises. However, the fluorescence values are lower relative to Extracted CHL than usually seen, but not as low as seen during 2018-031 when a sensor was presumed to have malfunctioned. In this case there are too few chlorophyll values >1ug/L to make a judgment and the only one >5ug/L was in Saanich Inlet where there was extremely high variability in fluorescence at 2db.

When the fluorometer plumbing was cleaned in spring 2019, a small obstruction at the intake was found.

It is possible that this may account for the low values found.

For full details of the comparison see file 2018-032-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-032-ctd.

The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. As usual the T, C and Sal channel pairs track well on downcasts and upcasts are much noisier. There appear to be more spikes in the primary temperature than in the secondary. The Dissolved Oxygen, pH, transmissivity, fluorescence, pH, PAR and SPAR profiles look normal. The PAR and SPAR values are close at the surface. The altimetry was fine near the bottom except for casts #10 and #11 during which the log notes that the chains were on and interfered with the altimetry.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +3s certainly improves the alignment and overall looks like a good choice for both DO sensors. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. None of the casts were deep enough to determine if there is any pressure dependence. The data in shaded cells are from the only previous cruise that used these sensors since they were last calibrated.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-030-0050 | 300 | -0.0008 | +0.00003 | +0.0009 | High, Steady |
| 2018-030-0056 | 300 | -0.0010 | -0.00002 | +0.0007 | F.High, V.Noisy |
| 2018-030-0086 | 300 | -0.0013 | -0.00003 | +0.0008 | F.High, Noisy |
| 2018-032-0053 | 370 | +0.0005 | +0.0001 | +0.0007 | High, Steady |
| 2018-032-0063 | 400 | +0.0004 | +0.0001 | +0.0009 | High, FSteady |
| 2018-032-0079 | 340 | +0.0002 | +0.0001 | +0.0010 | Mod, Noisy |

The salinity differences are consistent with the differences found in the bottle comparison and similar to the previous use. The temperatures are closer than in the previous cruise and conductivity values further apart.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. Fluorescence has high values during cast #1, but a plot does not suggest that it went off-scale.

A cross-reference list was checked against the log book. No problems were found.

The cruise track was plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. Based on those checks the altimetry header was removed from 2 files due to interference from chains and 1 during which the CTD had not reached within 15m of the bottom. For 7 casts the water depths were changed in the CLN files to match the log book entries if those were better or to an estimated value. 4 bottle files were also affected so the changes were applied to SAMAVG, MRG, MRGCLN2 and MRGCOR1 files.

A surface check was run and shows an average surface pressure for the cruise was 2.5db which is reasonable for the Vector. The lowest pressure measured was 0.9db at the end of cast #17; the associated salinity was ~31psu so this is clearly in water. There was a single deck pressure reading of -1.1db so it is possible that pressures are a little low, but given that readings tend to vary, we need more than one to justify recalibration. The sensor was recalibrated in March 2018 so it is unlikely there had been significant calibration drift.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

pH

Tests were run on 3 casts to determine the best setting to align pH with temperature. A setting of +40 records looked best.

Conductivity

Tests were run on 4 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. The results varied from cast to cast with some cases of any setting making some near-surface features worse, but those were associated with reversals in temperature that look suspicious.

SHIFT was run twice on all SBE911 casts using -0.6 records for the primary conductivity and -0.4 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The oxygen sensor has not been used since it was last calibrated at the factory. The pressure and conductivity sensors were used during 2018-030 in June 2018. The primary salinity was found to be lower than bottles by 0.002psu and the secondary lower by 0.001psu.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions in temperature were at stations 101 and 103 at the mouth of Juan de Fuca where temperatures were a little higher than the range maximum. Salinity was also a little low at station 103. This looks like a case of deeper mixing than usual. There were some salinity values that were slightly above the climatology maximum at mid-depths and near the bottom, east of Texada Island and at station 56. The excursions are not systematic, the climatology is old and not really suitable for near-shore observations, so these do not suggest a calibration problem.

Repeat Casts – There were no repeat casts and nearby casts are too shallow to provide a reasonable comparison.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary salinity appears to be closer to bottle values, but there are fewer spikes in the primary T-S plots that are not seen in the primary. The primary channels were also chosen for editing as was also the case during 2018-030.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. Most of the data removed are from near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing except for event #31.

After editing T-S plots were examined for all casts. There are some unstable features, most of which look likely to be real as this is an area of active mixing. Event #33 at station 51 looks odd in the top 8m but there is no obvious instrumental cause for this. No further editing was applied.

##### Recalibration

The most reliable dissolved oxygen data appear to be those from the 4 casts nearest the mouth of Juan de Fuca where flushing of bottles appears to be best.  This will be used for recalibration:

The primary and secondary salinity channels were found to be low by 0.007psu and 0.006psu respectively, using data from below 150m and excluding cases with high standard deviation in the CTD salinity. There were problems with the Autosal used to analyze the salinity bottles so this result is not trusted. During the previous use of these sensors the primary CTD salinity was found to be low by 0.002 and the secondary by 0.001psu. So the primary salinity will be corrected by adding 0.002psu.

File 2018-032-recal1.ccf was prepared to add 0.002psu to channel Salinity:T0:C0 and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0324 + 0.0096 (R2 = 0.89)

This correction was first applied to the SAM and MRGCORBOT files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The primary salinity was found to be low by an average of 0.005psu. This may be partly due to calibration drift and evaporation/adsorption during storage, but is likely primarily due to errors in the bottle values. See file 2018-032-sal-comp3.xls for details.

COMPARE was rerun for dissolved oxygen. Using the same data as for the fit the average difference shows the CTD to be higher than the bottles by 0.005mL/L.  The standard deviation is 0.014mL/L and the fit looks good over the whole range. When only the casts from western Juan de Fuca Strait are used selecting the same points as in the original comparison, the CTD salinity is low by an average of 0.0002mL/L with a standard deviation of 0.015mL/L. See file 2018-032-DO-comp2.xls for details.

At this stage flag 3 was added to all Salinity:Bottle values in the MRGCOR1 files. There will be a comment in the headers to explain the problem with the bottle analyses.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When a few outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.02mL/L (standard deviation of 0.09L/L). This is expected due to the incomplete flushing of Niskin bottles. No further recalibration is justified.

A plot of differences versus pressure was used to make an estimate of errors at different depths.

See 2018-032-dox-comp3.xlsx for details.

##### Correction to pH channel

An error was discovered at this point. The step to align the primary conductivity channel was run on the wrong files so the alignment of the pH channel was lost. The SHFpH files put through DELETE. Merge was run to add the aligned pH channel to the COR1 files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

There was a problem with the salinometer used to analyze samples for this cruise. Tests

were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”.

They established a linear correction to be applied to all salinity analyzed on salinometer

#68572 between January and March 2019. Based on those corrected bottle values and the history of the sensors, the CTD salinity values are believed to be within +/-0.004psu.

The fluorescence values look lower compared to extracted chlorophyll than usual, but it is

hard to judge this since most chlorophyll values were very low. When the fluorometer plumbing

was cleaned in spring 2019, a small obstruction at the intake was found.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.35 mL/L from 0 to 50db

±0.15 mL/L from 50 to 100db

±0.10 mL/L from 100 to 200db

±0.03 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see document: 2018-032_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S space where waters are well-mixed, but those are very small and possibly real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~55% to 95%. Those values are unusually low.  Most casts were very well mixed in the top 25m. The highest values were from casts with a somewhat larger vertical gradient in the top 10m. A check was made again that the CTD DO was close to bottles at the surface and it was.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-032-bottles-final.xls. The entries were compared with the rosette log sheets. For cast #72 it was discovered that for one bottle in the file with no sample number (due to a misfire) had not been removed. That line was removed from the MRGREO file, the number of records amended and HEADEDIT was rerun.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files. No problems were found.

##### Thermosalinograph Data

There was no thermosalinograph available for this cruise.

Particulars

1. Test cast – bottles 10-24 fired for testing only – removed from CHE file.

9. Rough landing

10. Chains on – interfered with altimeter.

11. Chains on – interfered with altimeter.

12. Chains off – altimeter ok.

20. deck pressure ~1.10db.

20. Cleaned and reseated the primary conductivity cable and connectors.

34. Slightly off station.

68. Hex file only – converted ok and no bottles were fired.

78. bottle #1 tripped by accident – don’t prepare CHE file

82. An extra bottle was fired at surface and likely sampled, but no sample number was assigned so all are said to be from Niskin 16, sample 297.

84. 15 min to settle on station –position & time in header look ok.

98. Bottle fired accidentally – no CHE file to be prepared.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  |  |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  |  |  | 0506 |  | Yes |  | Yes |  |  |  |
|  | Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Sensor |  |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  |  |  | 2023 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  |  |  | 1763 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  |  |  | 5013 |  | 2Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  |  |  | 953DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| PAR |  |  |  |  | 4565 |  | 16Mar2011 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  |  |  | 3685 |  |  |  |  |  |  |  |  |
| pH sensor |  |  |  |  | 692 |  | 2Mar2018 |  | Factory |  |  |  |  |
| Surface PAR |  |  |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |
| Pressure Sensor |  |  |  |  | 0506 |  | 13Mar2018 |  | Factory |  |  |  |  |
| Altimeter |  |  |  |  | 43281 |  | ? |  | Factory |  |  |  |  |