## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 March 2023 | ### Added HPLC data. S.H. |
| 17 March 2021 | Shallowest CTD data available added to surface samples in BOT files. G.G |
| 13 July 2020 | Corrected sample number in file 2018-035-0013.bot. G.G. |

## PROCESSING NOTES

Cruise: 2018-035

Agency: IOS, Ocean Sciences Division, Sidney BC

Party Chief: Belton M.	 		Platform: Neocaligus

Location: Strait of Georgia		Project: Strait of Georgia Zooplankton

Chief Scientist: Young K. 		Date: 27 April 2018 –2 May 2018

Processed by: Germaine Gatien

Date of Processing: 10 October 2018 – 18 October 2018

Number of original HEX files:	28

Number of CTD files: 28			Number of BOT files: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176,SBE pH sensor 0852 and pressure sensor 464.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with a good description of the protocols used for sampling.

A 1.7L Niskin was attached about 3m above the CTD for deep water samples. Samples at 5m and 10m were sampled with 5L Niskin bottles. A timer was set to 2 minutes. The CTD was lowered to about 10m then raised and held at the surface until the 2 minutes was up. The full cast was then run. This approach enables easy removal of the data from the soak period by removing a set number of records.

Names were in non-standard format with station names in the file names rather than event numbers. It would simplify processing if CTD files had the event number in the name.

Header information was entered in the raw files in a format that permitted easy conversion into IOS Header format. That is very helpful.

The sampling scheme was complex as described above. It makes sense to combine the samples from a single site into a single file and the rosette log sheets indicate that was the plan. They can then be combined with CTD data to give a more complete picture of the conditions at that station.

Combining CTD and bottle analysis results was an error-prone process especially since the sample numbers are in order of collection but not in order of depth. It would help if the event log contained a column that listed the associated CTD event number for samples.

There were surface samples with no corresponding CTD data either because there was no CTD cast or because the CTD did not get close enough to the surface to provide relevant data. Because there are no pressure readings for some of the bottles, a nominal depth channel was created and values added for all bottles based on rosette sheet entries. Depth:Nominal should be chosen for plotting or exporting data if you want to include surface samples.

There were 6 deep salinity bottle samples during this cruise. The CTD salinity read lower than the bottles by a median value of 0.001psu. This comparison is subject to errors such as inefficient flushing of Niskin bottles so that bottle samples come from higher in the water column than the CTD observation; the near-bottom vertical salinity gradients were complex and of varying sign so the net error is likely small. Another potential error is due to delayed analysis which may raise the bottle salinity values through evaporation and/or adsorption; this error is not likely to cause a change of more than 0.003psu. Some of these errors would offset others, so the net errors are not expected to be very large. CTD salinity is likely good to ±0.003psu.

There was no dissolved oxygen sampling during this cruise or previous cruises using this equipment in 2017. An inter-ship comparison of 2 casts from a June 2017 cruise using these sensors with a cast from a cruise that did include calibration sampling, suggested that dissolved oxygen values were likely low by up to 4%. A comparison with an SBE911+ cast at the same site during cruise 2018-05 had a lot of scatter but suggested that dissolved oxygen was reading low by at least 5%. Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

Extracted chlorophyll samples were taken at the surface and no reliable CTD data are available from that level, so no comparison with SBE fluorescence was possible.

Bottle files were produced combining CTD data with analysis results where possible; for the casts with only surface sampling the files contain only the analysis results with nominal depth. For plotting or data extraction it is recommended that Depth:Nominal be used in order to capture the surface samples.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were non-standard; they included the station name but lacked the event number.

2. Preliminary Steps

The Daily Log, rosette logs and a spreadsheet of sampling done were obtained.

The deployment method used was as follows: The CTD was switched on, a 2-minute timer was started and the CTD was put in the water. It was taken down to 10m, up to the surface, wait at the surface at least until the 2 minutes were up. The cast was then started. If a deep salinity sample was needed a Niskin bottle was attached 3m above the CTD. The CTD was taken to 10m off the bottom where a Niskin was fired if needed. The CTD was then brought up and switched off as it was taken out of the water. This method is helpful as by removing the first 960 records, the initial soak can be removed. In 2 cases there was a delay which was noted in the log book, so more records may need to be removed for those. This is a good method.

The cruise summary sheet was completed.

The parameters in the configuration file provided were checked and those for the fluorometer were found to be wrong. The scale factor is one that is sometimes selected but was not used on the previous cruise so it seemed unlikely. A test conversion showed that it produced some large negative values at depth. The file was corrected and saved as 2018-035-ctd.xmlcon.

## 3. Conversion of Raw Data

The HEX files were renamed using event numbers based on information in the log book.

The HEX files were converted using configuration file 2018-035-ctd.xmlcon.

Plots show that the channels all produced reasonable values. The descent rates look high and fairly steady. The temperature profile looks a little odd in the northern part of the Strait but it is stable and likely real. Fluorescence is very close to 0 in deep water so that instrumental noise leads to very small negative values.

4. WILDEDIT

Since there are no obvious spikes in the data, this step was skipped.

5. WFILTER

Test were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. WFILTER was run using a cosine filter, size 5 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25.

CELLTM was run on all casts using (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values. Negative fluorescence values were replaced with pad values.

The header times are in PDT so ADD TIME CHANNEL was run to add 7 hours.

10. Checking Headers

An initial track plot looked ok.

A cross-reference list was produced and compared with the log book entries. The only discrepancy was the longitude for event #1. Comparisons with other cruises that occupied that site (GEO1) show that the header entry is most likely correct – it is likely that a 0 was missing after the decimal point.

The next step is to remove the data collected during soaks at 10m. CLIP was run to remove 900 records from each cast. While there are 960 scans in 2 minutes there were a few cases where the downcast started a few scans earlier, so 900 scans were removed. Events #63 and 75 were slightly shorter and the soak for cast 8 was just over 1 minute, so CLIP was run separately for those 3 files. Plots were made after this step and there appears to be no significant loss of data near the surface.

HEADER CHECK was run. There were some negative values in some channels; most will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of 0.5db before CLIP was run and 1.0db after CLIP. The salinity was very low for the pre-CLIP values.

When the CTD pressure drops to <0db on upcasts the fluorescence immediately drops to very low values but with occasional spikes to very high values, perhaps due to waves. Salinity drops to low values within 1 to 4s. A pressure of -0.6db is the minimum measured with most values much closer to 0. This suggests that the pressure is reasonably accurate, well within the specifications of ±1db.

##### 11. SHIFT

Conductivity

During 7 other cruise when the same sensors were used, a setting of +1.2 records was found suitable to align the conductivity with pressure. During 2018-96 that setting worked for most casts, though no setting seemed to help for a few. For this cruise applying a setting of +1.2 made the alignment between the T and C traces worse and the resulting data looked much worse in T-S plots. A setting of -0.2 records made a slight improvement.  So a shift of -0.2 records was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; no further adjustment was found appropriate.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. The same was true for this cruise with a setting of +15 records looking slightly better than +10. SHIFT was run to advance the pH signal by 15 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 7 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records near the top and bottom of many casts and to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts required some editing.

Notes of editing details were made in the headers.

T-S plots were examined and no significant problems were found; there are small unstable features but this is an area where they are expected.

##### 14. Initial Bottle Data Steps

There was no rosette available for this cruise. Samples were taken from a Niskin Bottle mounted about 3m above the CTD and then samples were taken at 5m, 10 and 15m after the CTD cast. A surface sample was taken while the CTD was in the water. To enable searching of bottle data, BOT files were prepared.

Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-035-bot-hdr.txt which will be updated as needed during processing. All of the spreadsheets included a precision study.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-035 CHL*.xls

SALINITY

Salinity analysis was obtained in spreadsheet QF2018-035SAL*.xls.

Workbook 2018-035-bottle_plus_CTD.xlsx was created. Separate sheets were created to contain the data from the event log, bottle data (salinity and CHL), CTD data at bottle depths, CTD plus bottle data, CTD plus bottle data in 6-line header format and bottle-CTD comparison (salinity only).

Because it is impossible to do all the sampling in a single cast, there are often 2 or more events noted in the Daily Science log book for a single site. But it makes most sense to combine them in a single bottle file, corresponding to a single sample/rosette log sheet. Generally the first event number listed on a rosette log sheet corresponded to the CTD event. The analysts used the first event number on the rosette logs. For clarity, it might help to make an explicit entry on the log sheet indicating to which CTD event the bottle data belong. This approach would also allow for the possibility that the sampling might have to be done out of order at some time in future.

To select CTD data to go with the bottle samples, a list of depths sampled (including those for the 6 deep bottles) was prepared and edited and bin-averaged CTD files were thinned to those depths. Data were exported and copied to a separate sheet of the workbook and records were removed for which there was no sampling. This is the most error-prone part of this task.

Those data were added to the bottle data in the appropriate worksheet. The data had to be rearranged to match the bottle samples as there are no CTD files to match some of them and the order of sampling was not always the same. For the cases where events were combined the time was selected from the CTD cast. Nominal depths were entered since there are no CTD data for the surface samples.

On one spreadsheet the data were organized with a 6-line header which was saved separately as a CSV file, 2018-035-bottle-plus_CTD_6linlehdr.csv and that was converted to IOS files, 1 per cast.

For some casts there are only CHL samples, no CTD data. For others, there are samples from 0m but no CTD data were available from that depth.

##### 15. Compare

Salinity Comparison

A spreadsheet in workbook 2018-035-bottle_plus_CTD.XLS was used to do a comparison between the bottle salinity and CTD salinity. The CTD was found to be lower than bottle salinity by an average of 0.001psu ranging from being low by 0.005psu to being high by 0.007psu. There was a delay of almost 2 months in analysis which may raise bottle salinity values by 0.002 to 0.005 due to adsorption and/or evaporation. That would make the CTD data look lower than it really is. Another factor is incomplete flushing of Niskin bottles so that the water in the bottles comes from higher in the water column. Generally that would lead to bottles having lower salinity water so that the CTD looks higher than it really is, but for this cruise the vertical gradients were complex, so it is unclear what effect poor flushing would have. For 2 casts there was a large salinity gradient reversal near the bottom and those are the 2 casts for which the CTD salinity looks like it is reading lowest.

It is likely that the delayed analysis effect is at least partly offset by the errors due to flushing for 4 of the casts and the average for those shows the CTD salinity to be within 0.001psu of bottles. The other 2 bottles indicate it is low by about 0.005psu but incomplete flushing effects would be added to delayed analysis effects, so the CTD is likely reading better than that. Overall an estimate that salinity is good to ±0.003psu looks good. This is good correspondence.

For more details see the “salinity-comparison” worksheet in file 2018-035-bottle_plus_CTD.xlsx.

Fluorescence

Chlorophyll sampling was done only at 0m so no CHL-Fluorescence comparison was possible.

16. Other calibration checks

Sensor History – These sensors were used for 8 previous cruises since they were last calibrated. No problems were found in the pressure. There was little useful calibration sampling but comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – The only excursion from the local climatology was a section of cast #22 which had  temperature values above the historic minimum between 7and 15db. On a T-S plot this feature looks stable. The cast was close to shore and that area may not be well represented in the climatology. The feature is likely real. There is no evidence of a calibration problem.

Post-cruise calibrations – None were available.

17. CALIBRATE

CTD Salinity will not be recalibrated.

There was no dissolved oxygen sampling but the results of a study on the same sensor during 2018-05 will be applied to these data.

The pressure does not need recalibration.

CALIBRATE was run using file 2018-035-recal1.ccf  to apply the following correction to the Dissolved Oxygen:

CTD Dissolved Oxygen Corrected = 1.05* CTD Dissolved Oxygen

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show high values (115 to 130%) in the north-east section of the Strait of Georgia and low values (70 to 95%) in the area south of Vancouver, the lowest being those are in the Gulf Island area. The central and north-eastern sections have values in the mid-range (95% to105%).

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Data Processing Notes:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There were 6 deep salinity bottle samples from this cruise. The CTD salinity

values were lower than the bottles by a median value of 0.001psu which is well

within the expected accuracy. However, delay in analysis may have led to increased

bottle salinity values that would make the CTD appear to read lower than it really

is. As well, incomplete flushing of Niskin bottles may lead to bottle salinity

differing from in situ waters in an unpredictable way; local vertical salinity

gradients were complex with some large reversals so the errors may be of variable

sign. Some of these errors would be compensating. So the CTD salinity may be a

little less accurate than it appears in the comparison, but is likely within

+/- 0.003psu.

A comparison with an SBE911+ cast during a previous cruise also showed that

the salinity was likely within expected accuracy.

No recalibration was applied to the CTD salinity.

There was no dissolved oxygen sampling during this cruise or previous cruises

using this equipment in 2017. An inter-ship comparison of 2 casts from a

June 2017 cruise using these sensors with a cast from a cruise that did include

calibration sampling, suggested that dissolved oxygen values were likely low by

up to 4%. A comparison with an SBE911+ cast at the same site during a February

2018 cruise had a lot of scatter but suggested that dissolved oxygen was reading low

low by at least 5%.

Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples.

For details on the processing see document: 2018-035_Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

##### 20. Final BOT file preparation

To enable searching of bottle data, BOT casts were created that contain sample data and CTD data from the preceding downcast. The event number will match the CTD cast.

File 2018-035-6linehdr  was updated by thinning final CTD files and exporting data to a spreadsheet so as to capture the effects of recalibrating DO data and deriving DO in mass units. The new values did not affect the bottle comparison since CTD salinity was not changed.

This process would be easier if the samples were gathered in either order of depth or reverse order of depth, or if that is not practical, if the event log spreadsheet contained nominal depths.

The spreadsheet file was converted to IOS Header files for each cast.

The time and date are present as channels as these cannot be changed directly into header entries.

CLEAN was run to add START and END time. The END TIME is identical so the START time so it will be removed later. CLEAN was also used to enter 0 flags where the flag channels are empty and to remove channels with only pad values.

REMOVE was run to remove the DATE and TIME channels.

SORT was used to order the files by Depth:Nominal and then Sample Number. (They already are in that order, but this was run just in case there was an error.)

HEADEDIT was used to add comments and to remove the END time since it is the same as START TIME.

The final files have extensions BOT.

The standards check was run and the only warnings concerned non-standard channel names. Those were corrected.

A cross-reference list and header check were run on the BOT files. No problems were found.

Plots were made of all BOT casts. No problems were found.

Finally all data from BOT files were extracted to a spreadsheet and compared to the rosette sheets and no problems were found.

CRUISE SUMMARY

| Cruise ID#: 2018-035 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 27 April 2018 End: 2 May 2018 |  |  |  |  |  |  |
| Location: Strait of Georgia Zooplankton |  |  |  |  |  |  |
| Party Chief: Belton M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2018-035

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 7Nov2015 | Factory |  |  |
| Conductivity | 1764 | 22Dec2015 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 11Nov2015 | Factory |  |  |
| pH | 0852 |  |  |  |  |
| Pressure | 0464 | 28Feb2017 | Factory |  |  |