## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19 Feb. 2025 | Corrected sal units and channel name: HPLC:c2-MGDG. G.G. |
| 27 March 2023 | Added HPLC data. S.H. |
| 18 March 2021 | Replaced bad Fluorescence value in 2018-036-0038.ctd and corrected surface records in CHE files 50 and 72. G.G. |
| 17 March 2021 | Shallowest CTD data available added to surface samples in BOT files. G.G |
| 8 July 2019 | Corrected sample # in file 2018-036-0038.bot |

## PROCESSING NOTES

Cruise: 2018-036

Agency: IOS, Ocean Sciences Division, Sidney BC

Party Chief: Belton M.	 		Platform: Neocaligus

Location: Strait of Georgia		Project: Strait of Georgia Zooplankton

Chief Scientist: Young K. 		Date: 27 April 2018 –2 May 2018

Processed by: Germaine Gatien

Date of Processing: 18 October 2018 – 22 October 2018

Number of original HEX files: 28

Number of CTD files: 28		Number of BOT files: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176,SBE pH sensor 0852 and pressure sensor 464.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with many comments about problems encountered.

A 1.7L Niskin was attached about 3m above the CTD for deep water samples. Samples at 5m and 10m were sampled with 5L Niskin bottles. A timer was set to 2 minutes. The CTD was lowered to about 10m then raised and held at the surface until the 2 minutes was up. The full cast was then run. This approach enables easy removal of the data from the soak period by removing a set number of records.

Header information was entered in the raw files in a format that permitted easy conversion into IOS Header format. That is very helpful.

Early in the cruise it was discovered that the configuration file in use was for a different CTD. Another configuration file was found with the correct parameters except that the fluorometer scale factor was wrong. Because the sensor serial numbers were wrong in the hex files conversion failed using the corrected con file until the data were uploaded again with correct serial numbers.

The sampling scheme was complex as described above. It makes sense to combine the samples from a single site into a single file and the rosette log sheets indicate that was the plan. The analysis data can then be combined with CTD data to give a more complete picture of the conditions at that station.

Combining CTD and bottle analysis results was an error-prone process especially since the sample numbers are in order of collection but not in order of depth.

There were surface samples with no corresponding CTD data either because there was no CTD cast or because the CTD did not get close enough to the surface to provide relevant data. Because there are no pressure readings for some of the bottles, a nominal depth channel was created and values added for all bottles based on rosette sheet entries. Depth:Nominal should be chosen for plotting or exporting data if you want to include surface samples.

Some records contain only CTD data; these are associated with phytoplankton sampling.

There were 6 deep salinity bottle samples during this cruise. The CTD salinity read higher than the bottles by a median value of 0.002psu. However, incomplete flushing of Niskin bottles may lead to bottle values being lower than in situ values. Thus the CTD may look like it is reading higher than it really is. The CTD salinity data are considered +/- 0.003psu which is well within the expected accuracy for this equipment. No recalibration was applied to the CTD salinity.

There was no dissolved oxygen sampling during this cruise or previous cruises using this equipment in 2017. An inter-ship comparison of 2 casts from a June 2017 cruise using these sensors with a cast from a cruise that did include calibration sampling, suggested that dissolved oxygen values were likely low by up to 4%. A comparison with an SBE911+ cast at the same site during cruise 2018-05 had a lot of scatter but suggested that dissolved oxygen was reading low by at least 5%. Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

Extracted chlorophyll samples were taken at the surface and no CTD data are available from that level, so no comparison with SBE fluorescence was possible.

The temperature and salinity values often fell outside the climatology for spring, but the one cast in June fell well within the summer climatology. So either late spring values are not well represented in the climatology or late spring 2018 ocean conditions were unusually summery. The climatology is 10 years old.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were non-standard; they included the station name but lacked the event number.

2. Preliminary Steps

The Daily Log, rosette logs and a spreadsheet of sampling done were obtained.

The deployment method used was as follows: The CTD was switched on, a 2-minute timer was started and the CTD was put in the water. It was taken down to 10m, up to the surface, wait at the surface at least until the 2 minutes were up. The cast was then started. If a deep salinity sample was needed a Niskin bottle was attached 3m above the CTD. The CTD was taken to 10m off the bottom where a Niskin was fired if needed. The CTD was then brought up and switched off as it was taken out of the water. This method is helpful as by removing the first 960 records, the initial soak can be removed. This is a good method.

The cruise summary sheet was completed.

Data from the first day were sent to IOS for initial assessment and it was discovered that the configuration file provided was incorrect for this CTD. Eventually a corrected file was found but conversion could not be completed without either editing each file to change the serial numbers in the headers, or the files needed to be uploaded again with the correct sensor numbers. The new configuration file had an error in the fluorometer scale. This was corrected and the file was saved as 2018-036-ctd.xmlcon.

## 3. Conversion of Raw Data

The HEX files were converted using configuration file 2018-036-ctd.xmlcon.

Plots show that the channels all produced reasonable values. The descent rates look high and fairly steady. The temperature profile looks a little odd in the northern part of the Strait but it is stable and likely real. Fluorescence is very close to 0 in deep water so that instrumental noise leads to very small negative values.

4. WILDEDIT

There were a few spikes in the data, so WILDEDIT was run on the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

5. WFILTER

Test were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. WFILTER was run using a cosine filter, size 5 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. One cast was missing its latitude header; that was added and the file was reconverted.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

10. Checking Headers

An initial track plot looked ok.

A cross-reference list was produced and compared with the log book entries. There were no significant discrepancies though there are small differences in positions.

The next step is to remove the data collected during soaks at 10m. CLIP was run to remove 850 records from each cast. While there are 960 scans in 2 minutes there were a few cases where the downcast started a few scans earlier, so 850 scans were removed. Event #1 had a shorter wait so 700 scans were removed and cast #72 had a delay so that 2000 scans needed to be removed to avoid including data form the initial drop.  Plots were made after this step and there appears to be no significant loss of data near the surface.

The header times are in PDT so ADD TIME CHANNEL was run to add 7 hours.

HEADER CHECK was run. There were some negative values in fluorescence and pressure; most will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of 0.6db before CLIP was run and all the associated salinity values were very low.

It is difficult to check the accuracy of the pressure sensor but examination of the end of casts as the pressure goes through 0 shows fluorescence and conductivity dropping rapidly as pressure varies from -0.3db to +0.3db. Those values are well within the specifications of ±1db. The minimum pressure recorded was -0.4db.

##### 11. SHIFT

Conductivity

During 7 other cruise when the same sensors were used, a setting of +1.2 records was found suitable to align the conductivity with pressure. During 2018-96 that setting worked for most casts, though no setting seemed to help for a few. For 2018-035 applying a setting of +1.2 made the alignment between the T and C traces worse and the resulting data looked much worse in T-S plots. A setting of -0.2 records made a slight improvement. Tests on the data from 2018-036 show similar results to 2018-035, so a shift of -0.2 records was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. There are parts of profiles where a larger or smaller offset is seen in dissolved oxygen relative to temperature. Given these are likely reflective of different response rates in large gradients, alignment is unlikely to improve this. No further adjustment was made.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. The same was true for this cruise with a setting of +10 records looking slightly better than +15. SHIFT was run to advance the pH signal by 10 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 7 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records near the top and bottom of many casts and to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts required some editing.

Notes of editing details were made in the headers.

T-S plots were examined after this step and no significant problems were found; there are small unstable features but this is an area where some instability is expected.

##### 14. Initial Bottle Data Steps

There was no rosette available for this cruise. For 6 casts samples were taken from a Niskin Bottle mounted about 3m above the CTD. Samples were also taken at 5m and 10m after the CTD cast. A surface sample was taken while the CTD was in the water. To enable searching of bottle data, BOT files were prepared.

Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-036-bot-hdr.txt which will be updated as needed during processing. All of the spreadsheets included a precision study.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-036 CHL*.xls

SALINITY

Salinity analysis was obtained in spreadsheet QF2018-036SAL*.xls.

Workbook 2018-036-bottle_plus_CTD.xlsx was created. Separate worksheets were created to contain the data from the event log relevant to Niskin sampling, bottle data (salinity and CHL), CTD data at bottle depths, CTD plus bottle data, CTD plus bottle data in 6-line header format and bottle-CTD comparison (salinity only).

Because it is impossible to do all the sampling in a single cast, there are often 2 or more events noted in the Daily Science log book for a single site. It makes most sense to combine them in a single bottle file, corresponding to a single sample/rosette log sheet. The first event number listed on a rosette log sheet corresponded to the CTD event. The analysts used the first event number on the rosette logs.

There was phytoplankton sampling from bottles from which there was no CHL or SAL sampling, so lines for those samples need to be added to the sampling list on the bottle data worksheet.

To select CTD data to go with the bottle samples, a list of depths sampled in increasing order (including those for the 6 deep bottles) was prepared and edited and bin-averaged CTD files were thinned to those depths. Data were exported and copied to a separate sheet of the workbook and records were removed for which there was no sampling and lines were added for the surface samples. This is the most error-prone part of this task.

Those data were added to the bottle data in the appropriate worksheet. The data had to be rearranged to match the bottle samples as there are no CTD files to match some of them and the order of sampling was not always the same. For the cases where events were combined the time was selected from the CTD cast. Nominal depths were entered since there are no pressure data for the surface samples.

On one worksheet the data were organized with a 6-line header which was saved separately as a CSV file, 2018-036-bottle-plus_CTD.csv and that was converted to IOS files, 1 per cast.

##### 15. Compare

Salinity Comparison

One sheet in workbook 2018-036-bottle_plus_CTD.XLS was used to do a comparison between the bottle salinity and CTD salinity. The CTD was found to be higher than bottle salinity by an average of 0.003psu, median of 0.002psu ranging from being low by 0.001psu to being high by 0.009psu. Analysis was done promptly so negligible error is expected due to adsorption and/or evaporation. There likely is some error due to incomplete flushing of Niskin bottles so that the water in the bottles comes from higher in the water column. For 4 casts that would lead to bottles having lower salinity water so that the CTD looks higher than it really is. In 2 cases there is a slight decrease in salinity near the bottom so bottle contents could be either higher or lower than ambient conditions depending on how poor the flushing was. For the two casts with the CTD looking highest compared to bottles the vertical salinity gradient near the bottom was also highest, with salinity 5m above the firing level being lower by more than  0.010psu. Thus it is likely that the salinity error is smaller than indicated by the comparison. An estimate that salinity is good to ±0.003psu looks appropriate.

For more details see the “salinity-comparison” worksheet in file 2018-036-bottle_plus_CTD.xlsx.

Fluorescence

Chlorophyll sampling was done only at 0m so no CHL-Fluorescence comparison was possible.

16. Other calibration checks

Sensor History – These sensors were used for 10 previous cruises since they were last calibrated. No problems were found in the pressure. There was little calibration sampling available but comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. There was salinity sampling during 2018-035 that suggest salinity was within 0.003psu. There was no dissolved oxygen sampling for any of the cruises. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Most of the casts had temperatures above the climatology maxima in the top 10 to 50db and below that for many. A few were entirely outside the climatology. Salinity values were lower than the climatology minima in the same way. The final cast fell entirely within the climatology and was the only one that took place in June, so that it was compared with summer values rather than spring. So either the climatology does not contain data from late May or late May 2018 had unusual conditions. The excursions do not appear to be indicative of calibration problems.

Post-cruise calibrations – None were available.

17. CALIBRATE

CTD Salinity will not be recalibrated.

There was no dissolved oxygen sampling but the results of a study on the same sensor during 2018-05 will be applied to these data.

The pressure does not need recalibration.

CALIBRATE was run using file 2018-036-recal1.ccf  to apply the following correction to the Dissolved Oxygen:

CTD Dissolved Oxygen Corrected = 1.05* CTD Dissolved Oxygen

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

There are a few negative fluorescence values at depth so CLEAN was run to replace those with pad values.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 80 to 135%. The highest saturation rates were in the north-east section of the Strait of Georgia and Baynes Sound with the lowest values in the Gulf Island area.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Data Processing Notes:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There were 6 deep salinity bottle samples from this cruise. The CTD salinity values

were higher than the bottles by a median value of 0.002psu. However, incomplete

flushing of Niskin bottles may lead to bottle values being lower than in situ values.

Thus the CTD may look like it is reading higher than it really is. The CTD salinity

data are considered +/- 0.003psu which is well within the expected accuracy for

this equipment. No recalibration was applied to the CTD salinity.

There was no dissolved oxygen sampling during this cruise or previous cruises

using this equipment in 2017. An inter-ship comparison of 2 casts from a

June 2017 cruise using these sensors with a cast from a cruise that did include

calibration sampling, suggested that dissolved oxygen values were likely low by

up to 4%. A comparison with an SBE911+ cast at the same site during a February

2018 cruise had a lot of scatter but suggested that dissolved oxygen was reading low

low by at least 5%.

Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples.

For details on the processing see document: 2018-036_Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

##### 20. Final BOT file preparation

To enable searching of bottle data, BOT casts were created that contain sample data and CTD data from the preceding downcast. The event number will match the CTD cast.

The 6-line header worksheet from file 2018-036-bottle-plus_CTD.csv was updated by thinning final CTD files and exporting data to a spreadsheet so as to capture the effects of recalibrating DO data and deriving DO in mass units. The new values did not affect the bottle comparison since CTD salinity was not changed.

The spreadsheet file was converted to IOS Header files for each cast.

The time and date are present as channels as these cannot be converted directly into header entries.

CLEAN was run to add START and END time. The END TIME is identical so the START time so it will be removed later. CLEAN was also used to enter 0 flags where the flag channels are empty and to remove channels with only pad values.

REMOVE was run to remove the DATE and TIME channels.

SORT was used to order the files by Depth:Nominal and then Sample Number. (They should already be in that order, but this was run just to be sure.)

HEADEDIT was used to add comments and to remove the END time since it is the same as START TIME.

The final files have extensions BOT.

The standards check was run and the only warnings concerned non-standard channel names. Those were corrected.

A cross-reference list and header check were run on the BOT files. A problem was found with cast #38 which had a Pressure channel but no CTD data. The spreadsheet was corrected to remove the pressure entry and that cast was reconverted and reprocessed.

Plots were made of all BOT casts. No problems were found.

Finally all data from BOT files were extracted to a spreadsheet and compared to the rosette sheets and no further problems were found.

PARTICULARS – notes from logs

All times are local in the raw files.

Request sent to IOS for spare Niskin bottles – most tubing/rubbers are loose and most need bolts/butterfly nuts.

1. CTD left on until 1644 PDT

4. Big wire angle on CTD cast. Stopped @58m on way down to straighten the wire.

6. Need spare Niskin bottles – most tubing/rubbers are loose and most need bolts/butterfly nuts.

16. Approximate time – didn’t get positions

17. Down to 130m then after going up a bit, realized didn’t got down far enough, so went down to 230m, then up as usual.

38. Sample number 21 changed to 9021 because of duplication.

46. Messenger didn’t deploy at 10m, so redeployed.

69. Approximate 10° wire angle.

78. Approximate 15+° wire angle.

CRUISE SUMMARY

| Cruise ID#: 2018-036 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 26 Mayl 2018 End: 1 June 2018 |  |  |  |  |  |  |
| Location: Strait of Georgia Zooplankton |  |  |  |  |  |  |
| Party Chief: Belton M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2018-036

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 7Nov2015 | Factory |  |  |
| Conductivity | 1764 | 22Dec2015 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 11Nov2015 | Factory |  |  |
| pH | 0852 |  |  |  |  |
| Pressure | 0464 | 28Feb2017 | Factory |  |  |