## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 20 March 2025 | Updated channel names & formats in TOB files. |
| 5 March 2019 | Correction to nutrient data in CHE files for events 75, 77, 81, 88. |

## PROCESSING NOTES

Cruise: 2018-039

Agency: OSD

Location: WCVI / Strait of Georgia

Project: La Perouse

Party Chief: Robert M.

Platform: Sir Wilfrid Laurier

Date: 7 May 2018 – 18 May 2018

Processed by: Germaine Gatien

Date of Processing: 9 August 2018 – 31 August 2018

Number of original HEX files:  	96 (1 test)	Number of CTD files:  (89 SBE911, 6 SBE25)

Number of rosette files: 	59	   	Number of bottle casts processed: 59

Number of original TSG files: 	2		Number of processed TSG files:  2

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the primary pump, a SeaPoint Fluorometer (#3641) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613) and an altimeter (#62355).

CTD #334 with strain gauge pressure sensor (#482) was used without a rosette. Attached were a WetLab Eco fluorometer (#2215) and a SBE 43 DO sensor (#766)

A thermosalinograph (SeaCat 21 S/N 3274) was mounted with a remote temperature sensor and a SeaPoint fluorometer (#3654). There was no flow meter.

Seasave version 7.26.2.13 was used for acquisition.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+ #425.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

As described by the chief scientist in the cruise report there were serious problems with Niskin bottles and lanyards with reports of tops not being seated properly and leaks in spigots. Chips were noted on one Niskin and it broke during a later cast. The comparison of CTD dissolved oxygen and salinity with bottle samples was affected by this problem, so recalibration was based on the results of cruise 2018-26 that followed 2018-39 and used the same equipment. The problems with Niskin bottle leaks also made it impossible to estimate accuracy for the dissolved oxygen sensor mounted on the CTD.

For more information on these problems see the chief scientist’s post-cruise report: “PAC2018-039_Science Cruise Report view.pdf” in the DOC folder.

The chief scientist reported that the sounder was not operating well during this cruise. This was confirmed by finding some unreasonable values particularly during deep casts. An estimate of water depth was made by adding the maximum depth sampled and the bottom altimetry reading. In cases where the water depth entry differed from the estimate by more than 4m and where the altimetry profile shows the values are reliable, the estimate was used to replace the water depth entry. This was done for 16 casts including most casts deeper than 400m that sampled within 15m of the bottom. The entries for other deep casts are suspect but altimetry is not available or not reliable enough to make a better estimate.

Events for which the entry was adjusted were: 37, 40, 45, 46, 48, 50, 75, 89, 91, 92, 105, 106, 148, 154, 171 and 180.  Bottle files were also adjusted.

There were 7 casts using an SBE25 without a rosette, one of which was a test so has not been archived. It was very helpful that the headers had station name, water depth, latitude and longitude entered so that the information was automatically entered into IOS headers when the data were converted to IOS Header format. The pressure sensor was not working properly with values staying constant for about 1 second, then changing to a new value. There are variations within a group but only at the noise level. This problem has been noted before and was supposed to be fixed when the sensor was last sent to the factory for service. Pressure was filtered to produce a smoother profile, but the quality of data is lower than usual. Alignment and descent rates could not be examined in the usual way. Other variables do vary in the normal way.

There was a test cast during which the SBE25 was used for a 50db cast immediately before a full cast was run with the SBE911.

- The top 20m were well-mixed and the comparison of the two casts at that level suggests that salinity from the SBE25 is high by about 0.003psu and temperature low by about 0.002C°. Below 30db the vertical gradients are larger and the differences between the 2 sensors vary in sign and are larger, in general. This likely reflects the poorer resolution of the SBE25 and the problems in the pressure sensor may exaggerate this.

- The SBE25 fluorometer data were quite close to the SBE911 in the well-mixed layer.

- Dissolved oxygen from the SBE25 is lower at all levels except in a narrow band around 25db where the vertical DO gradient was very high. The SBE DO sensors generally drift to lower values and the one on the SBE25 appears to be low by about 6%.

SBE25 DO values were recalibrated by multiplying values by 1.06and SBE25 pressure was recalibrated by adding 2.3db. The salinity and temperature were not recalibrated as the comparison was not extensive enough to justify such a small correction.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

There were 2 Thermosalinograph files. There was no flow meter. There was conflicting information about the calibration of the intake temperature. It may be high by up to 0.07C°. The comparisons had a lot of scatter but TSG salinity was reasonably close to loop samples and CTD values.

# PROCESSING SUMMARY

PART I – SBE911

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, ammonium, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors was checked. They were all used during 2 cruises, 2018-001 and 2018-029, since they were last calibrated.

- The calibration control file was checked. The only error was in factor E in the dissolved oxygen sensor parameters. The value entered was 0.032 – the default value is 0.036, but a value of 0.0363 was entered based on hysteresis tests during 2018-01. The corrected file was saved as 2018-039-ctd911.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2018-039-ctd.xmlcon. Depth was included.

The ROS file for event #1 was deleted since it was a test only with no sampling.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers but none were found.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-039-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-39chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-039chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2018-39oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2018-039oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2018-39SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 10 to 20 days after collection. The files were simplified and saved as 2018-039sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2018-39_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-039nuts.csv. The file was converted to individual NUT files.

AMMONIUM

NH4 data were obtained in spreadsheet 2018-39 QF NH4*.xlsx. This includes a precision study. The file was simplified and saved as 2018-039NH4.csv. This file was converted to NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few problems were found:

- Many DO samples were flagged due to problems with top caps leaking or not being seated right or other leaks. Some other samples from bottles with problems were not flagged so flags were attached to all samples where such problems occurred. The comment was adapted from OXY: to ALL:. Appropriate flags were added to the other variables.

- Event #37 - 3 bottles were fired but samples came from only Niskin #1, so the lines for Niskin #2 and 3 were removed from the SAMAVG file and the merge with the samples was rerun..

- Event 36 sample #114 nutrient values look out of line in profile. The analyst agreed and padded the values.

After the corrections the merge process was repeated.

The incidence of top caps not fitting and/or spigot leaks is very high, so a study was made to see what can be learned.

| Niskin # | Integrity problems noted by DO analyst | # of times fired | # of times DO sample |
| --- | --- | --- | --- |
| 1 | 3 | 59 | 38 |
| 3 | 2 | 52 | 39 |
| 5 | 1 | 41 | 40 |
| 9 | 2 | 25 | 24 |
| 10 | 3 | 22 | 10 |
| 11 | 11 | 20 | 14 |
| 14 | 1 | 9 | 9 |
| 16 | 3 | 9 | 9 |
| 17 | 1 | 8 | 6 |

Bottles #1, 3 and 5 did quite well given the number of uses.

Bottle #11 was the worst with 11 cases of the DO analyst noting leak problems out of 20 firings, only 14 of which had DO sampling. The bottle itself broke and was replaced before event #121 but the trouble continued, so there must have been problems with the lanyards as well. Bottle #11 was noted as being chipped well before it broke.

Most of the problems are with odd-numbered bottles, but bottle #16 failed on 3 out of 9 uses.

Most salinity sampling was from bottles with DO sampling as well, so the salinity does not offer extra evidence about whether leakage occurred at a significant level for bottles with no DO sampling.

A check was made of the rosette sheets to ensure no comments about leaky bottles were present for cases where there was no DO sampling.  Problems were noted for samples #297 and #446 so those were noted in the CHL, NUTS and NH4 samples, as required. Merges were rerun.

##### Compare

Salinity

Compare was run with pressure as reference channel.

Six outliers with differences between bottle and CTD salinity >0.02psu were removed from the comparison. Most come from close to the surface with only 1 below 20m and the standard deviations in the CTD salinity are >0.002psu for all of them and in most cases much higher than that. So the large differences are likely due primarily to incomplete flushing of Niskin bottles in the presence of large vertical salinity gradients. For sample #471 the bottle had already been flagged 3 due to the top cap not being seated properly and a spigot leak. No further flags are justified by the comparison.

Bottles fired at the bottom were excluded as well as cases where differences were >0.01 or standard deviations in the CTD salinity was >0.001 or pressure <160db. The remaining bottles show the primary salinity to be high by an average of 0.0012psu and the secondary to be low by 0.0027psu. There is slightly more pressure dependence in the primary salinity but more temporal variability in the secondary.

The deep bottom bottles look little different from others in the fit, but the shallower one (~300m) stands out with the CTD looking higher than expected which would happen if the bottle salinity was lower due to incomplete flushing.

For full details for the COMPARE run see file 2018-039-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

To see if the problems with leaks would severely affect the comparison, all bottles flagged for that reason were excluded. The pattern shows that most differences were not notably out of line, but where they were the CTD was reading lower than the bottle by more than expected. We expect a leak to cause higher DO in a bottle than expected; only one such bottle was below the OMZ where the effect would be more complex.

When bottles are excluded that have a standard deviation in the CTD DO >0.1 and further outliers are excluded based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0698 + 0.0494 (R2 = 0.97)

This is a much higher slope than found in February during cruise 2018-001, or in April during 2018-029.

The only major outliers were either from very close to the surface or had a large standard deviation in the CTD DO data or had been flagged due to bottle integrity problems.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises. Fluorescence is somewhat higher relative to CHL for higher CHL values than usually seen.

For full details of the comparison see file 2018-039-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-039-ctd.

The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present and the traces look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. That setting was used during 2018-01 and 2018-026.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. The differences were very noisy so the variability in the results is not surprising. The salinity differences are not far from the difference implied by the results of the bottle comparison, ~0.004psu. There is too much noise and the casts are not deep enough to see if the pressure dependence in temperature and salinity that was noted during 2018-01 continued.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0059 | 300 | ~0 | -0.0004 | -0.004 | High, V.Steady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |
| 2018-039-0044 | 1000 | -0.0009 | -0.0005 | -0.0047 | High, XNoisy |
| “ | 1900 | -0.0008 | -0.0003 | -0.0036 | High, XNoisy |
| 2018-039-0088 | 1000 | -0.0003 | -0.0004 | -0.0043 | High, VNoisy |
| “ | 1900 | -0.0006 | -0.0003 | -0.0030 | High, VNoisy |
| 2018-039-0110 | 1000 | -0.0002 | -0.0006 | -0.0065 | High, VNoisy |
| “ | 1900 | -0.0003 | -0.0004 | -0.0040 | High, VNoisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The SBE25 casts were processed up to the point of header checks so all data could be included.

The header check was run and no problems were found.

The cross-reference check was run. A few errors were found and fixed:

- The format of the station name was adjusted for event #2.

- For event #6 it looks like there was a delay after the CTD file was started before the cast began as the time and positions are significantly different from the log entries. A text editor was used to enter the log entries in the CLN file.

The cruise track was plotted and added to the end of this report including the SBE25 casts.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. There were many cases where the check value was very high.

The chief scientist reported that the sounder was not operating well during this cruise. This was confirmed by finding some unreasonable values particularly during deep casts. An estimate of water depth was made by adding the maximum depth sampled and the bottom altimetry reading. In cases where the water depth entry differed from the estimate by more than 4m and where the altimetry profile shows the values are reliable, the estimate was used to replace the water depth entry. This was done for 16 CLN files including most casts deeper than 400m that sampled within 15m of the bottom. The entries for other deep casts are suspect but altimetry is not available or not reliable enough to make a better estimate.

Events for which the entry was adjusted were: 37, 40, 45, 46, 48, 50, 75, 89, 91, 92, 105, 106, 148, 154, 171 and 180. For those casts that included bottle sampling, the SAMAVG files were also adjusted.

A surface check was run and shows an average surface pressure for the cruise was 3.0db which is reasonable for the Laurier. There are a few cases of pressures as low as -0.35db but they all occur at the ends of casts when the pumps were off so it is hard to judge from salinity whether the CTD was in water or air. In one case the pumps were on but the pressure is varying between positive and negative values rapidly so the CTD may have been dipping in and out of the ocean. There is insufficient evidence to make a judgment about the pressure accuracy, though it may be reading slightly low.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. The primary traces were extremely noisy towards the end of the cruise.

SHIFT was run twice on all SBE911 casts using -0.5 records for the primary conductivity and -0.4 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 2018-01 and 2018-029. The primary salinity was found to be high by ~0.0017 and the secondary low by ~0.0007 during the first cruise. During 2018-029 they were found to be low by 0.001psu and 0.0057psu. The differences from 2018-01 were thought to be due at least partly to incomplete flushing of Niskin bottles in the presence of large vertical gradients. The dissolved oxygen was recalibrated using slope 1.0283 and offset 0.0491 which was the result of the comparison for 2018-01. The comparison for 2018-029 was not trusted due to likelihood of Niskin bottles not flushing well.

Experience after this cruise –

During the cruise that followed this one the primary salinity was found to be low by 0.0012psu but there was a different temperature sensor in use. The secondary salinity was found to be low by 0.0039psu.

The fit for dissolved oxygen depended on what data were included. The fit that looked most reliable had a slope of 1.0276 and offset 0.477. Slightly higher slopes were found when only deeper samples were included.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions for temperature were a few cases of high values near the surface on the LBP line. There were a number of excursions in salinity including high values around 50m on the inshore half of the LBP line, low values at the bottom at station LC4 and a few low near-surface values. None of these suggest calibration problems.

Repeat Casts – The only repeat casts using this sensor were too shallow to be useful. There was also a repeat cast to compare the SBE911+ with the SBE25+.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary conductivity was extremely noisy towards the end of the cruise. The bottle comparison contains only a few bottles from that period so this was noted in COMPARE. There is slightly more pressure dependence in the primary fit of differences in COMPARE. During 2018-01 there was some concern about the primary temperature sensor and it was replaced before the cruise that followed this one. Overall, the secondary sensors look more reliable than the primary.

The choice of which T-S pair to choose for editing and eventually archiving is not obvious. The primary salinity appears to be closer to bottle values. However, the primary conductivity has more spikes than the secondary data so they will not require as much editing and there is concern about the primary temperature sensor at high pressures. The secondary was also chosen when these sensors were used during 2018-001 and 2018-029.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. For some casts highly variable descent rates including some speeds >1.5m/s resulted in noisy salinity. All files required some editing and many required heavy editing due to very noisy descent rates, including complete reversals of direction.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some unstable features but those are near the surface and no instrumental cause was obvious, so they may be real. One bad section of salinity was found in file #44; the bad values were padded using a text editor.

Bad data were found below 500db in the fluorescence channel of cast #89 as noted by the chief scientist. CTDEDIT was run again to remove the bad FL values and the output file was copied to *.EDT.

No further editing was applied.

##### Recalibration

The secondary salinity was found to be lower than bottles by ~0.0027psu. Since there were problems with Niskin bottle seals, it is likely that the bottles contain some water with lower salinity than the ambient values at the firing level. So the CTD would appear to be reading higher than it really is. During 2018-26 which followed this cruise, the CTD was found to be reading low by 0.0039psu. This is a more reliable result as there were many fewer bottle problems.

Similarly for the dissolved oxygen the comparison for this cruise was out of line with other cruises using the same equipment, so the results of 2018-26 will be used for recalibration.

File 2018-039-SBE911-recal.ccf was prepared to add 0.0039psu to channel Salinity:T1:C1 and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0276 + 0.0477

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The secondary salinity was found to be high by 0.0012psu. Given the problem of leaks we do expect that the CTD salinity will be higher, on average, than the bottles, but this is due to the bottle values being too low.

See file 2018-039-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the corrections were applied properly. When data are excluded based on residuals, the CTD DO to be low by ~4.2%, with an offset ~0. While the CTD appears to be reading low, this is expected due to leakage of higher DO into bottles.

See file 2018-039-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

Normally a final check is made by comparing downcast DO with upcast bottle DO. Given the problems with Niskin bottle leaks this step was skipped as the bottle data is not reliable.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

The chief scientist reported that the sounder was not operating well during this

cruise. This was confirmed by finding some unreasonable values particularly during

deep casts. An estimate of water depth was made by adding the maximum depth sampled

and the bottom altimetry reading. In cases where the water depth entry differed

from the estimate by more than 4m and where the altimetry profile shows those

values are reliable, the estimate was used to replace the water depth entry.

This was done for 16 casts including most casts deeper than 400m that sampled within

15m of the bottom. The entries for other deep casts are suspect but altimetry is not

available or not reliable enough to make a better estimate.

Events for which the entry was adjusted are: #37, 40, 45, 46, 48, 50, 75, 89, 91,

92, 105, 106, 148, 154, 171 and 180.

Conductivity, Transmissivity, Fluorescence:URU:Seapoint and PAR data are nominal

and unedited except that some records were removed in editing temperature and

salinity.

There were serious problems with Niskin bottles and lanyards with reports in the logs

of tops not being seated properly and leaks in spigots. Chips were noted on one Niskin

and it broke during a later cast. The comparison of CTD dissolved oxygen and salinity

with bottle samples was clearly affected by these problems, so recalibration was based

on the results of cruise 2018-26 that followed 2018-39 and used the same equipment.

These problems also made it impossible to make an estimate of accuracy for the dissolved

oxygen sensor mounted on the CTD.

For more details on the problems with bottles see the chief scientist's post-cruise

report "PAC2018-039_Science Cruise Report view.pdf" in the DOC folder.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on the CTD data processing see document: 2018-039_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found. This was repeated after the SBE25 data were completed so all could be reported together.

Profile and T-S plots were examined and no problems were found. There are some unstable features in T-S space, but those are small and possibly real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~85% to 165% with the lowest values in Juan de Fuca Strait and Haro Strait, while values were highest at stations LB01 to LB06 and E1; those were also areas of high fluorescence. Values offshore were generally between 96% and 107%. The values look reasonable for the area and season. During the 2017 spring La Perouse cruise most offshore values were between 102% and 106%

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2018-039-bottles-final.xls. The entries were compared with the rosette log sheets. A few problems were found:

- The NH4 data for event #88 were missing because they had been entered as event 85. Processing steps were repeated for that cast and the NH4 spreadsheet was amended.

- The chlorophyll data for event #77 were missing because they had been entered as event 72. Processing steps were repeated for that cast and the CHL spreadsheet was amended.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

PART II – SBE25

##### Seasave

This step was completed at sea. There were 7 files, but the last one was run as a test to compare with the SBE911 so not needed for archiving.

##### Preliminary Steps

The cruise summary sheet was completed.

A test cast was sent from sea and it was discovered that there was an error in the fluorescence parameters in the con file. The pressure resolution also looked poor.

The fluorometer scale factor and dark value were corrected in the configuration file used at sea and the file was saved as 2018-039-ctd25.xmlcon.

##### Conversion of Raw Data

All files were converted using configuration file 2018-039-sbe25-ctd.xmlcon.

Plots were examined and all expected channels were present. The traces look ok except for the pressure which has poor resolution. Values only change every 8 scans. This problem has been noticed in the past and continues despite a return to the factory for service.

The deployment method was to lower the CTD to about 10m and return to the surface, a process taking about 1.6 minutes. Removing 800 scans will remove the soak period from most casts.

The pressure at which conductivity drops suddenly during soaks and upcasts was -2db ±1db. This is outside the ±1db accuracy quoted by SeaBird for the pressure sensor. The pressure sensor was not resolving well with values only updated every 8 scans. Adding 2.3db to the pressures looks appropriate.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

There are frequent small steps in the pressure and some small reversals. Tests were run applying a variety of filters and the best choice was found to be a Windows cosine filter size 9.

WFILTER was run on pressure, depth, temperature and conductivity with cosine file width 9.

##### ALIGNCTD

Tests were run to find the best setting to align dissolved oxygen with temperature traces. The best results were found when the DO channel was advanced by 3s.

ALIGNCTD was run to apply a 3s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

The station names positions and water depths are missing from the headers, so they were added using a text editor.

The start times are bad, but not in a predictable way. Two are off by only about 15 minutes but others are off by many hours. The upload times appear to be what was selected. The actual event numbers are in the headers but in a corrupted format. A text editor was used to enter times to match the log.

CLEAN was run to add event numbers.

##### Checking Headers

At this point the CLN files were placed in the same folder as the 911+ files so that header checks could be done together.

CLIP was used to remove 800 initial records from the soak period to ensure that data are only kept from after the pumps came on and the sensors were able to equilibrate.  Plots were examined and show the number of records removed were appropriate

Track plot and Header Check were run together with the SBE911 data and no problems were found.

The surface check shows an average of -1.9db on files before CLIP was applied and -0.8db afterwards. The associated salinity values are extremely low before CLIP and much higher afterwards.

CALIBRATE was run using calibration control file 2018-039-recal-sbe25.ccf to add 2.3db to the pressure channel and 2m to the depth channel. Note that it was intended to add 2.3m to the depth so a further correction of +0.3m was made later in processing to make the cumulative correction +2.3m.

After that step the average surface pressure was 1.5db and pressure values as the CTD appears to leave the water range from -0.5db to +0.5db.

##### SHIFT

Fluorescence

The fluorometer on the SBE25 was not pumped so alignment was not needed.

Conductivity

It has been found during other cruises using this equipment that no setting improved most casts. The casts from this cruise are not suitable to test this. SHIFT was not applied to conductivity.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; the alignment looks good. SHIFT was not run on DO.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to remove corrupted by shed wakes; salinity was cleaned lightly in a few casts. Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined; there are a few near-surface unstable features but they are small and may well be real.

##### Other calibration checks

Sensor History – This is the first use of this pressure and conductivity sensors since they were recalibrated in May 2018. Unfortunately, the pressure sensor was supposed to be fixed, but was not.

The dissolved oxygen sensor has only been used once since it was last serviced and that was on a Boots camera so there was no calibration information.

Comparison of repeat casts – None were available.

Historic Ranges – All temperature data were within the climatology while there were 2 small patches of low salinity around 20m for the casts at B7 and LBP8. The same excursion was seen in the LBP8 cast using the SBE911+ so this is assumed to be a real feature, not due to instrumental problems.

Post-cruise calibrations – None were available.

Inter-comparison of the 2 CTDs – There was an inter-comparison between the two CTDs during casts #85 and 86. The SBE25 cast went to only 50db. The top 20m were quite well mixed vertically while the gradient was high from about 20 to 30db and below that the gradient is moderate.

- Temperature from the SBE25 was lower than the SBE911 by about 0.002C° at 16db and higher by about 0.05C° at 35db. Below 35db there is a section where the SBE25 is lower than the 911.

- Salinity from the SBE25 was higher than the SBE911 by about 0.0035psu at 16db and lower by 0.0057psu at 35db. Similar to temperature the differences vary in sign below 30db.

- Fluorescence values from the SBE25 were lower at 16db by about 0.08ug/L (~2%) and higher at 35db by 0.04ug/L (~8%). Because both fluorescence signals are very noisy, these values are not suitable to use for recalibration.

- The SBE25 Dissolved oxygen was consistently lower than that from the SBE911. At 16db it is low by about 5.8% and at 35db it is lower by about 6.2%. We expect these sensors to drift lower with time.

The results near the surface likely reflect the calibration reasonably well and the differences in temperature and salinity are reasonably small. The results at 35db probably reflect the SBE25’s inability to resolved gradients as well as the 911+. The known pressure resolution problems with the SBE25 may also account for some of the variability. The DO sensor has likely drifter lower by about 6%.

##### CALIBRATE

The only SBE25 variable for which we have sufficient evidence to justify recalibration is dissolved oxygen. It is likely low by 6%, so file 2018-039-sbe25-recal.ccf was prepared to multiply all DO values by 1.06 and to add 0.3m to depth to correct an error in the first recalibration of depth.

##### Filter

This step was skipped since the fluorescence data are not spiky.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Oxygen:Dissolved:SBE Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units.

Saturation rates were calculated and plots show near-surface saturation from 104% to 150%. The highest values were at E1, C1, C2 and C3, close to the inshore part of the LB line where similar values were seen from the 911+ CTD. For the test cast the value was close to that using the SBE911+. So the DO recalibration produced reasonable values.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

The pressure sensor was not operating normally during this cruise. Values were only

updating about once per second. Other variables updated normally. Pressure

was recalibrated by adding 2.3db based on observations of "in water" conductivity

values during upcasts when pressure was <0db.

Conductivity, Fluorescence and Transmissivity data are nominal and unedited except that

some records were removed in editing temperature and salinity.

There was no calibration sampling for the SBE25 casts, but there was a test cast run to

compare with a cast using an SBE911+. Salinity, temperature and fluorescence from the

2 CTDs were reasonably close; the comparison was not extensive enough to justify

recalibration or a statement of accuracy. Channel Oxygen:Dissolved:SBE was recalibrated

because values were consistently lower than the SBE911 sensor by about 6%.

For details on the processing see document: 2018-039_Processing_Report.doc.

The standards check was run and no problems were found.

At this point the CTD files were moved to the same folder as the SBE911+ files so that quality checks could be run on the combined set, excluding cast #85 which was only run as a test.

A cross-reference listing was produced for the whole cruise.

A header check was run including files from both CTDs. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PART III - TSG

##### Thermosalinograph Data

There were 2 hex files. There was no flow meter.The sampling interval was 5s.

a.) Checking calibrations

The configuration files were checked and are correct and did not change through the cruise. One was saved as 2018-039-tsg,xmlcon.

b.) Conversion of Files

The 2 files were converted to CNV files using file 2018-039-tsg,xmlcon.

##### c.) WILDEDIT

Plots show discrete single-point spikes in the salinity channel, so tests were run using a variety of settings to remove them using WILDEDIT. The best results were found by using parameters:

Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 25

Keep data within this distance of the mean = 0

The routine was run on both files.

The results show that the single-point spikes are removed, but a few problems remain.

- In one case there is a very large drop in values that lasts for about a minute – it looks like the flow might have stopped since the lab temperature rises while the intake temperature falls. These data will be padded later using CTDEDIT.

- In some other cases the spikes involve more than one scan but there is evidence of a small gradual drop in salinity over 5 to 20 scans and then a large drop for 1 or 2 scans. WILDEDIT removes the large drop, but not the gradual one. These gradual changes are not large, <0.02psu and they could be real. Some of these variations may be real, some due to small bubbles. In any case there are few such drops noted and they are small. So while the overall salinity measured by the TSG may be a little low due to bubbles, we do not expect a large overall effect.

##### d.) Conversion to IOS Header Format and Initial Quality Checks

The CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots showed that the salinity signal had only 1 significant spike in the first file and none in the second.

A track plot was produced and added to the end of this document. There are symbols every 2 hours.

Next data from the TSG and CTD were prepared in order to do a comparison.

The TSG files were averaged over 12 scans using BIN AVERAGE (with standard deviations included) and then REMOVE was used to remove channels not needed for the comparison with CTD files.

CTD data were thinned to 5m and then exported to a spreadsheet that was saved as 2018-039-ctd5-tsg-comp.xlsx.

The TSG files were opened in EXCEL and the files were reduced to the times of CTD files. There were 87 cases of overlap, 81 from the SBE911+ and 6 from the SBE25.

Similarly, TSG data were found to match the loop samples.

e.) Comparison of T, S and Fluorescence from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2

The temperature rose in the loop by an average of 0.21C°, standard deviation of 0.04C°.

As expected heating in the loop decreases as intake temperatures rise and approach the ship temperature.

- Flow Rate There was no flow meter.

- TSG vs CTD

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. When all casts are included the median differences between latitude and longitude are both 0.0000°. There were significantly larger differences during cast #11 for positions and other variables. That was an SBE25 cast. Because the quality of the CTD data is lower for the SBE25 those lines were moved to a separate page from the SBE911 comparison.

For the SBE911+ casts the median difference in latitudes and longitudes were both 0.0000° with no differences >0.0002° for latitude and 0.0006° for longitude.

Temperature, salinity and fluorescence were then compared for all SBE911+ casts. The median and standard deviations are given below using all the casts and then using only the 20 with the lowest standard deviation (over 12 records) in the TSG data for the relevant variable.

Note that in the table the standard deviation refers to the differences in the fit, not the variability in the TSG data.

|  |  | Tint - Tctd | Tlab-Tctd | Sal TSG-Sal CTD | TSG FL/CTD FL |
| --- | --- | --- | --- | --- | --- |
| all data | median | 0.0979 | 0.3111 | -0.0172 | 0.839 |
|  | std dev | 0.3286 | 0.3343 | 0.6488 | 0.225 |
| 20 “best” cases | median | 0.0677 | 0.2684 | -0.0016 | 0.920 |
|  | std dev | 0.2784 | 0.3776 | 0.0972 | 0.158 |

    - The intake temperature was found to be high by ~0.1 C° using all data or by 0.07C° using the 20 with low standard deviation. The standard deviation in the fit is very high. Examination of some casts with low vertical temperature gradients near the surface does not suggest that the difference is caused by the TSG drawing water from higher in the water column than 5m unless it came from much higher, above the level of CTD observations. When used during 2018-001 this system did not have an intake thermistor and the lab temperature was found to be higher than the CTD by about 0.23C°. We would expect heating in the loop to be at least 0.23C° in February so the TSG temperature was likely close to or lower than the CTD at that time. If the intake thermistor has drifted by 0.07C° then heating in the loop would be about 0.27C° which seems high, though if the flow rate was lower that is quite possible. Another possible explanation is that the CTD temperature was bad, but that is unlikely since the 2 CTD sensors are in good agreement.

    - The lab temperatures are higher than the CTD by 0.31C° or 0.27C°. That implies heating in the loop by about 0.2C° if we trust the intake temperature, or higher if we don’t.

    - The TSG salinity was found to be lower than the CTD salinity by 0.017psu using all data and 0.002psu using the lowest standard deviations. The standard deviation was very high for this comparison. The casts with differences <0.01psu were all from deeper, offshore waters where the near-surface waters tend to be better mixed.

In the plot below the fit of TSG fluorescence against CTD fluorescence is shown in black while the line in blue shows the fit if the two were equal. This shows that the two fluorometers are quite close at low values.

For the SBE25 casts the median difference in latitudes was 0.0000° and for longitudes was 0.0001° with no differences >0.0014° for latitude and 0.0056° for longitude.  TSG intake temperatures were higher than CTD temperatures by a median of 0.123C° but the standard deviation was 0.166C°. The lab temperature was higher than the intake temperature by ~0.23C°. The TSG salinity was higher than the CTD salinity by a median of 0.04psu with a standard deviation of 0.19psu. The ratio of TSG fluorescence to CTD fluorescence ranged from 0.63 to 1.13 with the higher ratios associated with lower fluorescence.

For more details see 2018-039-ctd5-tsg-comp.xls.

- Loop vs Rosette Bottles

Differences were calculated between samples from the loop and from the rosette samples closest to 5m. There were only 6 points of comparison available.

    - For salinity the loops were lower than the rosette by a median of 0.009psu with a standard deviation of 0.02psu. The differences ranged from being lower by 0.049 to higher by 0.0017psu. One sample had been flagged due to a leak. The median is not very helpful with so few samples. The 5 unflagged samples had loop salinity lower than the rosette salinity by an average of 0.019psu with a standard deviation of 0.12psu. It is notable that 2 of the 6 samples had the loop salinity low by about 0.004 and 1 high by 0.002psu. The two that were most out of line were from near-shore casts where the vertical gradients may be high enough to make a small difference in depth significant. It seems likely that the loops are within ±0.004psu of the rosette samples.

    - CHL from the loop was lower than the rosette samples by a median of 0.01ug/L with a standard deviation of 0.11ug/L. If the flagged sample is excluded and an average calculated, the loop CHL is lower by an average of 0.024ug/L with a standard deviation of 0.024ug/L.

There were 3 cases with loop values higher and 3 lower.

For more detail see 2018-01-loops-tsg-rosette-comp.xls.

- Loop Bottle - TSG Comparisons

There were 18 loop CHL samples and 17 loop salinity samples, of which 11 were taken underway.

    - TSG Fluorescence was mostly higher than CHL at low CHL values and mostly lower for CHL>1ug/L. This sort of pattern is typical of these fluorometers. The median ratio of FL/CHL was 0.86ug/L with a standard deviation of 0.45ug/L. A fit of FL vs CHL had a slope of 0.62 when forced through the origin. The ratio FL/CHL was a little lower for the underway sampling but CHL was higher for those data as well, so this not significant.

    - Salinity from the TSG was higher than that from the loop by a median of 0.009psu but the standard deviation was 0.033psu. The salinity read significantly higher when underway than when stopped but the standard deviation was also higher.  However, the two cases with the lowest standard deviation in the TSG salinity have TSG salinity lower than the loop samples; one is while stopped and one underway. So the picture is far from clear.

For more detail see 2018-01-loops-tsg-rosette-comp.xls.

- Calibration History

The temperature and conductivity sensors were recalibrated in December 2017.  The CTD was used during 2018-001 in February. There was no intake temperature available for that cruise. The lab temperature was found to be higher than the CTD by 0.23C°. Salinity was found to be high by ~0.01C°. For the Tully cruises in February and June of 2017 the heating in the loop was 0.18C° and 0.19 C°, respectively. The loop may be longer for the Laurier or the flow rate lower.

Conclusions

- The time and position data agree well between the CTD and TSG.

- The comparison between the SBE 25 CTD and the TSG is too noisy to be useful.

- The comparison between the SBE 911+ CTD and the TSG is also noisy but given many more points of comparison we can have a little more confidence in it.

- Alignment was not run before these comparisons, so the secondary temperature will be leading the primary temperature. This should be random in sign. It would lead to noisy comparisons but not a systematic error.

- There was no flow meter but the differences between the intake and lab temperature was fairly steady except for small variations explained by intake temperature variations and/or rapid changes in temperature.

- The lab temperature is higher than the intake temperature by ~0.20C° and higher than the CTD by ~0.27 C°. The comparison with the intake is about what we would expect for heating in the loop at this time of year on the Tully. (In February and June of 2017 the heating in the loop was 0.19C° and 0.18 C°, respectively.) The loop may be longer for the Laurier or the flow rate lower which would increase heating.  During 2018-001 on the Laurier in February the lab temperature was higher than the CTD by 0.23 C°. This was assumed to be due to heating in the loop but there could have been drift in the lab thermistor as well so this is not strong evidence. We would expect less heating in May. If the intake temperature is too high and the lab temperature is correct, then heating in the loop would be larger than expected, but certainly not impossible.

- The TSG intake temperature is higher than the CTD temperature by about 0.068C°.  There is no evidence of a significant error in the CTD temperature nor is there any indication that a mismatch in depths can explain the difference. So large an error due to temperature calibration drift is unusual, but there may be other explanations such as damage or where and how the external thermistor was installed. But if we assume that there is an error, then either heating in the loop is more than we expect or the lab temperature also has a significant error; the latter seems unlikely.

- The TSG salinity is lower than the CTD salinity by about 0.0016psu using the data with the least noise in the TSG record. This as close as we can expect to get, especially if small bubbles are present as they would lead to slightly lower TSG salinity. The loop salinity appears to be slightly lower than the CTD as well. If salinity is close to the CTD, then the lab temperature must be reasonably accurate.

- The loop salinity was likely within 0.004psu of the rosette, but there are few samples and too much scatter to be confident about this. Extracted CHL from the loop compared well with the rosette for higher CHL values. For lower CHL the differences are larger but random in sign.

- The TSG salinity is higher than loop samples by an average of 0.009psu with a standard deviation of 0.033psu. When the ship was stopped it is higher than loops by about 0.0085psu and while underway it was higher than loops by 0.0148psu, though the noise in the TSG salinity was also higher while underway.  Having the TSG salinity reading high does indicate that the lab temperature is unlikely to be reading high.

- The TSG fluorometer is reading a little lower than the CTD fluorometer, overall. It reads higher than the CTD sensor at low chlorophyll values and lower at high values.

- There are too many uncertainties to justify recalibration of the TSG data. If the intake temperature is too high by 0.07C°, then heating in the loop is higher than expected or the lab temperature is also too high. The salinity comparison is evidence that the lab temperature is not the problem. If the intake temperature is correct then the CTD temperature is lower than expected and we have evidence to suggest that is not the case.  A warning will be added to the headers to say that the intake temperature may be reading high by as much as 0.07C°.  If a post-cruise calibration later shows the intake temperature to be significantly high, this can be revisited.

e.) SHIFT

SHIFT was used to align the primary and secondary temperature. Checks were made to see by how many scans the secondary temperature leads the lab-based variables. This varied from 13 to 21 scans with some suggestion that the lead increased partway through the first file, presumably due to the flow rate increasing. The primary traces are smoother than the secondary due to mixing so this is approximate. SHIFT was run to align the secondary temperature by -16 records. Plots show this step was effective.

f.) Editing

Each file was opened in CTDEDIT to check that there are no obvious bad data. In file #1 10 points were padded where it appears the flow might have stopped; this step leaves the time and positions intact. (records 348 – 357.) No further editing was applied.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

The files were bin-averaged over 6 scans, thus reducing the interval to 30s.

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference,  Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats. Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found except that there were station names in the files. Those indicate where the files started, but are meaningless for these data, so they were removed.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

Particulars

PAR on for all casts done with SBE911+

Casts with bottle fired out of order: 36, 37, 59, 81, 86, 94, 114, 135, 154, 164.

1. Test cast – 24 Niskins closed but no sampling. Bottle file not prepared.

37. Niskins 11 and 24 not needed.

45. Soft bottom poor sounder depth.

48. Meant to be “out of order” but Niskins got closed in order so 3 Niskins closed at wrong depth.

85. Test cast with SBE25 to compare with SBE911. Don’t archive.

89. Weird noise in fluorometer ~500m.

91. One Niskin closed at wrong depth and 2 with no sampling (open spigot).

92. Depth from plan – no good read from sounder.

101. Niskin 11 not needed. No other bottles fired. Bottle file not prepared.

110. ~1 mile from station – high winds – not keeping station well.

114. No sample from Niskin11 – bottle broke.

151. Fluorometer data looked spiky. Cleaned connectors, applied connector lubricant.

157. Misnamed as 158 on rosette sheet.

164. Niskins all closed out of order and at the wrong depth. Check rosette sheet.

182. Fluorometer spiky.

CRUISE SUMMARY - CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  |  | CTD Calibration Sheet Competed? |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0443 |  | Yes |  |  | Yes |  |  |  |  |  |
| 2 | SEABIRD | 25 |  | 0334 |  | No |  |  | No |  |  |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  | Post Cruise |  |  |  |  |
| Name |  |  | S/N |  | Date |  | Location |  |  | Date |  | Location |  |  |
| Temperature |  |  | 2710 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Conductivity |  |  | 2128 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Temp. |  |  | 4752 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Cond. |  |  | 2399 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 7Aug2017 |  | Factory |  |  |  |  |  |  |  |
| SBE 43 DO sensor |  |  | 0997 |  | 2Sep2017 |  | Factory |  |  |  |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |  |  |  |
| SeaPoint Fluor. |  |  | 3641 |  |  |  |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 31Aug2017 |  | Factory |  |  |  |  |  |  |  |
| Altimeter |  |  | 62355 |  | 21Mar2016 |  | Factory |  |  |  |  |  |  |  |
| Calibration Information - 0334 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  |  | Post Cruise |  |  |  |
| Name |  | S/N |  |  | Date |  |  | Location |  |  | Date |  | Location |  |
| Temperature |  | 2968 |  |  | 22Feb2017 |  |  | Factory |  |  |  |  |  |  |
| Conductivity |  | 2173 |  |  | 23Feb2017 |  |  | Factory |  |  |  |  |  |  |
| ECO Fluorometer |  | 2215 |  |  | 7Mar2017 |  |  | Factory |  |  |  |  |  |  |
| SBE43 Oxygen |  | 773 |  |  | 1Mar2017 |  |  | Factory |  |  |  |  |  |  |
| Pressure |  | 482 |  |  | 28Feb2017 |  |  | Factory |  |  |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/21/3274

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3274 | 16Dec17 | Factory |  |  |
| Conductivity | 3274 | 16Dec17 | Factory |  |  |
| Temperature SBE38 | 271 | 3Dec2015 | Factory |  |  |
| SeaPoint Fluorometer | 3654 |  | Factory |  |  |