## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 20 March 2025 | Updated channel names & formats in TOB & Loop files. GG & SH |
| 2 July 2024 | Added EXPORTS carbon data. Corrections to DO flags, HPLC analysis comments, DMS/DMSP units, added missing flag channel and update of channel table format. GG |

## PROCESSING NOTES

Cruise: 2018-040

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 11 September 2018 – 28 September 2018

Processed by: Germaine Gatien

Date of Processing: 9 January 2019 – 21 March 2019

Number of original HEX files: 65	Number of CTD files: 62

Number of rosette files:	 64		Number of bottle casts processed: 62

Number of original TSG files: 19	Number of processed TSG files:  15

# INSTRUMENT SUMMARY

CTD #0585 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#1483 from events 1-47 and #997 from event 48 to 126) on the primary pump, a SeaPoint Fluorometer (#3641) on the primary pump, a Biospherical QSP-400 PAR sensor (#70613) and an altimeter (#62355).

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully ASUS.

The deck unit was a Seabird model 11+ #425.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in excellent order with comments about problems encountered. Sampling notes were provided which are a great help in processing data.

For DO sensor #1483 there were many spikes in the dissolved oxygen channel above 500m in both downcast and upcast. The spikes were much smaller in downcasts, ≤0.01mL/L, while upcast spikes were as large as ±0.05mL/L. No spikes were noted during bottle stops. This affects events #1-47.

There was a lot of  noise in the temperature and salinity data from the primary channels during this cruise and the previous one. It may be that putting both the fluorometer and dissolved oxygen sensors on the primary pump accounts for this. In the past there has usually been just one of these sensors on each pump. While this may lead to better data from the secondary channels that does mean that if there is a problem with any of the secondary sensors we have to rely on data of poorer quality from the primary system.

The CTD pumps were not turned on for event #5; a CHE file was prepared with pumped CTD channels removed.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

For sensor #1483

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.15 mL/L from 100db to 200db

±0.10 mL/L from 200db to 500db

±0.03 mL/L below 500db

For sensor #0997

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 125db

±0.2 mL/L from 125db to 500db

±0.05 mL/L from 500db to 1000db

±0.25 mL/L from 1000db to 1500db

±0.05 mL/L below 1500db

There were 16 Thermosalinograph files that covered the time of the cruise. Each file covers one day. The flow in the loop was turned off during parts of the files from September 13th and 27th. Only times and positions were included in the files when the flow rate was 0. The file for September 12th was not archived since the ship moved little and there was no flow in the loop. Fluorescence data were bad throughout the cruise so that channel was removed from the thermosalinograph files.

Comparisons were made between the TSG data and co-incident CTD casts, rosette samples and loop samples. These indicate that the intake temperature is as close to the CTD temperature as can be expected. The TSG salinity was lower than CTD and loop values by about 0.03psu to 0.04psu which could be due to the presence of small bubbles in the loop. Calibration drift is unlikely to be this high as this system was calibrated at the factory recently.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity and DMS data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Since their last calibration these sensors have only been used on 2 casts during 2018-025 and during all of 2018-034 except that the DO sensor used from events #48 to 126  has been used on many other cruises including most of cruise 2018-025.

- The calibration control files were checked. The DO sensor was changed at cast 48, so there are 2 different con files. The files were saved as 2018-040-ctd1.xmlcon and 2018-040-ctd2.xmlcon. The E factor in the DO sensor for file 2018-040-ctd2.xmlcon was updated based on tests during cruise 2018-026. No other changes were found necessary.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2018-040-ctd1.xmlcon and 2018-040-ctd2.xmlcon. Depth was included.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

A few alterations were made to BOT files:

- Event #36 – The data from Niskin #17 were removed from the BOT file as it was fired by mistake and there was no sampling there.

- Event #37 – Niskin 17 did not fire at 60m so Event #38 was run to get that sample. To create a single bottle file, the bottle number was changed in event #38 from 1 to 10 and named 37b, then JOIN was used to join the 2 segments called 37a and 37b. The resulting file for event 37 is out of pressure order, so the 60m data were moved to the beginning of the file.

- Event #67 – Niskin 1 failed to close at 5m, so event #68 was used to get Niskin 1. File 67.bot was renamed 67a.bot and the bottle numbers were revised from 1 to 9 to 2 to 10. File 68 was changed to 67.b. Join was used to combine parts b and a, to produce a new file 67.

Temperature and salinity were plotted for all BOT files to check for outliers and casts 1, 42, 48, 62 and 72 had a few in primary salinity. CTDEDIT was used to clean those outliers and to remove a few records from event #72 due to some bad values in all T and S near 20 and 25db. The output files were copied to *.BOT. The event numbers were missing, so the BOT files were renamed BOT1 and CLEAN was rerun to add event numbers and to ensure header values are correct.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. Many sample numbers had to be removed because they were used for the Trace Metal rosette or were not used at all. Niskin 1, cast 63 was removed from the list.

Niskin 8, cast 92 was removed from the list and sample 718 was used for Niskin 19.

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The records from Niskin 1, cast 63 and Niskin 8 cast 92 were removed from the SAMAVG files.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-040-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2018-040 CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2018-040chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF 2018-040 OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2018-040oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file QF2018-040SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 15 to 76 days after collection. The files were simplified and saved as 2018-040sal.csv. That file was then converted to individual SAL files.

The IOS Salinometer #68572 malfunctioned after a few samples had been analyzed. Those samples were flagged 2 by the analyst. Other samples were analyzed using another salinometer of the same type at AML Oceanographic.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2018-040*.xlsx. This includes a precision study. The file was simplified and the saved as 2018-040nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in spreadsheet DMS summary (2018-040).xls. Details on analysis are in file 2018-040DMS report.doc which includes duplicate analysis. An error in station name was found – cast 44 was the DMS cast, not 43. This file was converted to DMS files.

DMSP

DMSP data were obtained in spreadsheet 2018-40 dmsp QF summary.xls. This file was simplified and converted to DMSP files.

Final salinity data were not yet available so a temporary file was prepared and the MERGE process was run in order to run compare on DO data to see if there was a need to update factor E for DO sensor 1483. Later these steps were repeated with the final salinity data.

The SAL, CHL, OXY, NUT, DMS and DMSP files were merged with CST files in 6 steps.

After the 6th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Some problems were found:

- Nutrient samples 63-73 are identified as from event #4. The rosette file shows no nutrient sampling for event #4 nor for event #5. The rosette sheet for event #7 shows sampling  for samples 63-73. The event # was changed to 7.

- Events #23 & 58 have many types of sampling shows on the rosette log but only nutrients were found. It is assumed that the other samples were taken for non-OSD participants.

- Event #28 – DO sample 192 has a pad value but no flag or comment. The analyst added flag 1 as titration data was lost due to a computer crash.

- Cast #44 – there was an error in the sample #s in the addsamp file. This was fixed.

- Cast #53 – Salinity sample #363 missing –analyst found it.

The DO analyst noted 9 cases of leaks from Niskin bottles – in 8 of those cases there are other samples that need a 3 flag added (Samples 97, 198, 204, 267, 362, 379, 469, 648.  Those flags were adjusted. The comment is already entered as applying to ALL.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were 3 extreme outliers:

- Event #10 – Sample 76 is supposed to be a surface sample but has salinity that matches CTD salinity at about 90db. The conductivity on the analysis logsheet matches the spreadsheet entry. The salinity looks like the CTD salinity for sample 74 at the bottom of the cast. Nutrients and DO are not out of line so this does not look like a misfire. The analyst decided to reject the value.

- Event #48 – Sample 330 is supposed to be a surface sample but has salinity that matches CTD salinity at about 900db. The conductivity in the analysis spreadsheet differs from that in the spreadsheet. When corrected the salinity looks fine.

- Event #70 – Sample 473 is supposed to be a surface sample but has salinity that matches CTD salinity at about 900db. The conductivity in the analysis spreadsheet differs from that in the spreadsheet. When the salinity looks fine.

After excluding those outliers there is more scatter in the fits than usual. Cases where pressure is <350db were excluded as those depths are more likely to be affected by poor flushing of Niskin bottles and outliers with differences >0.01psu, there is still more scatter than usual:

|  | Average difference | Standard deviation |
| --- | --- | --- |
| Primary CTD Sal – Bottle Sal | +0.0004psu | 0.0025psu |
| Secondary CTD Sal – Bottle Sal | +0.0028psu | 0.0023psu |

The average differences are close to those observed during the previous cruise, 2018-34, but the standard deviations from the earlier cruise were much smaller at 0.0007psu. This may be partly due to analysis being delayed for this cruise which may introduce random errors and to 2 different salinometers being used.

The differences between primary and secondary salinity is 0.0025psu which is in line with the observations of differences between the 2 channels during downcasts as detailed in section 9.

A few studies

- The scatter in the fits is larger than expected. To see if there is a pattern to outliers plots were made that excluded differences >0.003psu from the average difference. As expected the outliers in the top 100db almost all have the CTD salinity looking low, which is the usual situation due to incomplete flushing of Niskin bottles in the presence of larger vertical gradients. Below that there are 7 outliers on the negative side; delay in analysis would lead to the CTD looking low due to evaporation and adsoprtion raising salinity values in bottle samples. There are 10 cases with the CTD looking too high by more than 0.003psu. Two of those have noisy CTD data. The other 8 were investigated. None were from the bottom where incomplete flushing can lead to CD values looking a little high. Two outliers came from the average of duplicates with one value very close to the CTD salinity. So one each of the duplicates for samples 309 and  315 were chosen and the outliers disappears for both; one of those outliers was on the low side and one on the high side. No explanation could be found for the other outliers.

- Cast 48 had a full deep profile. Variability from 1000db down is higher than expected and the pattern is the same for both sensor pairs, though average values differ. The range of differences is about 0.01psu at depth. The standard deviations in the CTD salinity channels is low.

- P4 early and late. The comparisons are very different with the late P4 cast looking like both CTD salinity channels were lower compared to bottles than they were earlier. The early cast samples were analyzed within 4 weeks while the others waited about 9 weeks and they were analyzed on different instruments at different bath temperatures. If there was significant evaporation and/or adsorption of samples the bottle salinity would rise and that would make the CTD salinity look lower than it really is. That might explain the Station P4 difference but not other stations sampled after the first P4 occupation. They should also look lower. The late P4 data look notably different from the early one in T-S space but not at the depths from which samples came.

- Time-series plots show a lot of variability with no clear patterns.

The average differences are close to those of earlier cruises. Delay in analysis could lead to some scatter to the low side. The salinometer is not one with which the analyst had any experience and the lab conditions may have varied with time. As we have more confidence in the results of cruise 2018-034, and the average results are similar to this one, it looks like 2018-034 results are the best choice for recalibration purposes.

When outliers larger than 0.002psu from the average were removed and 2 duplicate outliers replaced with sample “b” as detailed above the results were:

|  | Average difference | Standard deviation |
| --- | --- | --- |
| Primary CTD Sal – Bottle Sal | +0.0003psu | 0.0015psu |
| Secondary CTD Sal – Bottle Sal | +0.0029psu | 0.0013psu |

For full details for the COMPARE run see file 2018-040-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The data for each DO sensor were assessed separately.

For DO #1483

Problems with spikes were noted in the DO sensor at sea affecting upcasts more than downcasts and mostly above 500m. Three casts were examined that have clear spikes in the profiles to see if this problem was seen during bottle stops:

- Cast #7 – Spikes in the downcast from 10db to 60db and the upcast above 50db. Spikes stop during bottle stops.

- Cast #11 – Shallow casts. Spikes both down and up but much smaller in downcasts. No spikes during bottle stops. Downcast spikes ≤0.01mL/L while upcast spikes were as large as ±0.05mL/L.

- Cast #36 – Spikes are seen in downcasts but they are small. During upcasts they are larger There were many spikes of size ±0.1mL/L at about 400db but none during the stop at 500db. We expect some noise in the signal; for sensor #997 the noise level is about ±0.005mL/L.

A general examination of the standard deviations in CTD DO during stops does not suggest a problem. So whatever the source of spikes it appears that it did not have a significant effect during stops.

Outliers were gradually removed based on residuals. There were no severe outliers and no obvious pattern in the minor outliers except that many had fairly noisy CTD DO. A few had been flagged by the analyst for minor problems.

When bottles are excluded that have a standard deviation in the CTD DO >0.1 and further outliers are excluded based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0605 + 0.0376 (R2 = 0.97)

This is very close to the result during 2018-034.

Data below 1000db were excluded from the fit to see if there are signs of hysteresis. The 2 deepest measurements are lower than the trendline but they also have noise in the CTD DO. A value from 2500db is out of line in the opposite direction. There may be a little hysteresis but the evidence is weak, especially in light of problems noted with the sensor. There was only 1 cast deep enough to be affected significantly if there was hysteresis.

For DO #0997

There were 3 major outliers:

- Event #53, 50m, sample 382 – difference 0.939 – This appears to be a case of samples 382 and 383 being reversed in the DO spreadsheet.

- Event #53, 24m, sample 383 – difference -1.360  - As above. Both differences are in line if the DO values are exchanged.

- Event #116, 5m, sample 972 – difference 0.022 – already flagged 4 because there was no water seal and there were bubbles in the sample.

After those outliers were excluded plus others based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0237 + 0.0477 (R2 = 0.83)

Data below 1000db were excluded from the fit to see if there are signs of hysteresis and there was none. The deep points fall close to the shallow ones in a plot of differences versus SBE DO. The hysteresis parameter E had been updated based on earlier deep-water cruises.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises. CHL was mostly very low so the fluorometer is mostly higher than CHL.

##### Conversion of Full Files from Raw Data

All files were converted using 2018-040-ctd.xmlcon.

The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. As noted at sea there are spikes in the dissolved oxygen traces during the upcast above 500db until the DO sensor was changed.

As noted in the log the pumps were off during cast #5. The log mentions only the downcast, but the pumps were not on for the upcast either. The only files available from cast #5 would be pressure, transmissivity and PAR and since it is was night-time PAR is very low in both. CTD casts for that site were run before and after this cast, so the downcast will not be processed further.

Casts #4, 5 and 6 were all run at station JF2 and started roughly every 40 minutes. During the 1st cast the minimum transmissivity is about 80%/25cm with no notable mid-depth minimum. During the 2nd cast there is a distinct layer around 90db  with values down to 75%/25cm; below that values are similar to the 1st cast. During the 3rd cast the minimum is even lower and slightly deeper. For pumped channels only the 1st and 3rd casts can be compared; temperature, salinity and dissolved oxygen show some vertical displacement but no significant changes in the shape of the profiles, but around 90-100db they look somewhat different on T-S plots which likely corresponds to the transmissivity feature. Both casts #4 and 6 have noisy DO signals especially above 50db.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

There were some spikes in conductivity that were removed. The spikes in dissolved oxygen channel were wider, with too many records to be removed with this step.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. The differences are all small though the salinity differences grow slightly with time. There is little pressure dependence.

The salinity differences are very close to the difference implied by the results of the bottle comparison, ~0.0026psu.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-025-0217 | 300 | +0.0001 | -0.0002 | -0.0014 | High, F Steady |
| 2018-034-0038 | 1000 | -0.0006 VN | +0.00013 | +0.0019 | High, Noisy |
|  | 1900 | -0.0000 | +0.00010 | +0.0017 | “ |
| 2018-034-0091 | 1000 | -0.0006 | +0.00011 | +0.0018 | High, Noisy |
|  | 1900 | -0.0000 | +0.00012 | +0.0017 | “ |
| 2018-034-0113 | 1000 | -0.0001 | +0.00010 | +0.0014 | High, X.Noisy |
|  | 1900 | -0.0001 | +0.00010 | +0.0015 | “ |
| 2018-040-0048 | 1000 | -0.0004 | +0.0001 | +0.0021 | High, V Noisy |
|  | 1900 | -0.0001 | +0.0002 | +0.0022 | “ |
|  | 2900 | 0 | +0.0002 | +0.0022 | “ |
| 2018-040-0070 | 1000 | -0.0004 | +0.0002 | +0.0024 | High, Noisy |
|  | 1900 | -0.0002 | +0.0002 | +0.0023 | “ |
|  | 2900 | -0.0001 | +0.0002 | +0.0023 | “ |
|  | 3900 | -0.0001 | +0.0002 | +0.0025 | “ |
| 2018-040-0085 | 1000 | -0.0006 | +0.0002 | +0.0027 | High, V Noisy |
|  | 1900 | -0.0003 | +0.0002 | +0.0027 | High, V Noisy |
|  | 2900 | -0.0001 | +0.0002 | +0.0028 | High, Noisy |
|  | 3900 | +0.0001 | +0.0002 | +0.0027 | High, Noisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run. The only error found was the station name in event #80. It was changed from P36 to P26.

The header check was run. There are transmissivity values of 0 during cast #112; they occurred at about 3db during the soak. (Those values were not seen in the file after DELETE was run.)  No other unusual results were noted.

Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. Changes were made to 5 casts. The altimeter header entry was removed from casts 35, 44, 102 and 116 – spikes had been misinterpreted by the algorithm. For cast #74 the bottom depth was replaced with a sounder value found in the comments in the log.

For the bottle files the altimeter header entry was removed from casts 4, 77, 93 and 94 – these casts had no bottle sampling near the bottom while the header gives the impression it did. For cast #74 the bottom depth was replaced with a sounder value found in the comments in the log.

A surface check was run and shows an average surface pressure for the cruise was 2.9db which is reasonable for the Tully. All readings are associated with in-water salinity values.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts to assess whether the settings used to align conductivity and temperature  for cruise 2018-034 were appropriate for these data. Judging by the effect on salinity as seen in T-S space, it looks like those settings are the best choice.

SHIFT was run twice on all SBE911 casts using -0.6 records for both the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned cast #68 which had only surface data and will not be processed further.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature and conductivity sensors and DO sensor #1483 were used for just 2 casts during 2018-025 and for all of cruise 2018-034. During 2018-034 the primary salinity was found to read higher than bottles by about 0.0009psu and the secondary by 0.0025psu; DO was recalibrated using a slope of 1.0602 and offset of 0.0529. DO sensor #997 was used during many cruises including 2018-026 when there was an extensive comparison with bottles; DO was recalibrated using slope 1.0276 and slope 0.0477.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature data fell within the local seasonal climatology. West of P17 the only climatology available was cruder and covered the whole year. There were some cases where the base of the mixed layer was slightly warmer than the climatology maximum and above 50m water was fresher than the minimum. (Note that the late summer Line P cruise was run later in the year than usual.)

Repeat Casts – There were many repeat casts including some deep enough for a useful test of repeatability. Casts 80 and 85 at P26 started about 4 hours apart and sampled to at least 3000m. At about 2800m differences in temperature and salinity along lines of constant σt were ~0.002ºC and ~0.0002psu. This is good repeatability.

.Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There was more noise in the T-S plots using primary channels with more unstable features than the secondary. The primary salinity was closer to bottles. Neither salinity shows significant pressure or time dependence. During 2018-034 the secondary sensors were selected for editing and that was done for this cruise as well.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some small unstable features but those are near the surface and no instrumental cause was obvious, so they are likely real.

##### Recalibration

Pressure tests during 2018-034 showed no recalibration was needed. There is no evidence from this cruise to suggest otherwise.

Given the limitations of the salinity comparison for this cruise and given the average differences were close to those from cruise 2018-034, the results from 2018-034 will be used to recalibrate salinity.

File 2018-040-SBE911-recal.ccf was prepared to recalibrate as follows:

- For casts 1-47 - subtract 0.0025psu from channel Salinity:T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0602 + 0.0529

- For casts 48-126- subtract 0.0025psu from channel Salinity:T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0237 + 0.0477

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The secondary salinity was found to be high by an average of 0.0004psu. See file 2018-040-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the corrections were applied properly. When data are excluded based on residuals, removing roughly the same data the results were:

- For sensor #1483 CTD DO was high by an average of 0.015mL/L

- For sensor #997 it was high by an average of 0.002mL/L.

Standard deviations for both were about 0.02mL/L. The better results for the 2nd sensor may be due to better flushing of Niskin bottles further offshore. Both results are satisfactory and show the recalibration worked well.

See file 2018-040-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.001mL/L (standard deviation of 0.093mL/L) for CTD #1483 and ~0.037mL/L (standard deviation of 0.089mL/L) for CTD #0997 . Looking at the differences versus pressure shows that the sensor DO tends to be on the high side above 1000m for both sensors and slightly low below oxygen minimum zone, which is likely due to incomplete flushing of Niskin bottles.

No further recalibration is justified. See 2018-040-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------------

Conductivity, Transmissivity, Fluorescence and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

For sensor #1483

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.15 mL/L from 100db to 200db

±0.10 mL/L from 200db to 500db

±0.03 mL/L below 500db

For sensor #0997

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 125db

±0.2 mL/L from 125db to 500db

±0.05 mL/L from 500db to 1000db

±0.25 mL/L from 1000db to 1500db

±0.05 mL/L below 1500db

For details on the CTD data processing see document: 2018-040_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are a few odd features in the profiles:

- Cast #59 had a section of noisy Fluorescence data from 849 to 910db with values unrealistic for those depths. Those values were padded. No other variables show unusual features at that depth.

- Cast #76 has noisy transmissivity at depth but since the values are stated to be nominal and the general shape may be of interest, the data were not padded.

- Cast #85 has a noisy patch at about 1000m; as for #76 it was left unedited.

- Cast #116 and a spike in transmissivity at the bottom but this is likely real.

There are some unstable features in T-S space, but those are small, near the surface and possibly real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from ~40% to 110% with the exception of cast #38. That cast stood out as odd with a very high % saturation compared to all other casts and in particular to 4 other casts at station P12. Whereas the other casts had well mixed DO in the top 20m, there were very high gradients and unbelievably high values at that level for #38. This is very interesting because this cast was run only to fire a bottle and the 10m soak had not been done. This is further evidence that care is needed to soak the CTD to obtain good near-surface data. Cast #38 will not be processed further.

The values in Juan de Fuca Strait ran from 40 to 60%; in that region vertical mixing is likely to be strong.  Offshore values were mostly between 100% and 105% except for station P4 where values ranged from 108 to 110%. The offshore values are typical of values found there and suggest that the DO sensor calibration is reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #5 the pumped CTD channels were also removed.

For casts with no PAR sensor mounted (42, 48, 53-61, 70, 74, 76, 80-85), the PAR channel was removed.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A special note was added to file #5 to explain why CTD channels were removed and the information about analysis was removed, since there was no analysis done at IOS.

Data were exported from the CHE files to file 2018-040-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

A new thermosalinograph was in use for this cruise. It was used for the first time during cruise 2018-026 in June 2018 and has been used for 2018-025 and 2018-034. It did not support having a fluorometer attached, so a new method was developed to enable combining the TSG output with that from a fluorometer. There were 19 files produced but 2 were empty and one had no updates to positions so was likely a test at the dock. The files each contain data for a single day.

There were loop samples, a flow meter and intake thermistor. The intake is at about 4.5m.

The files have extensions RAW but are in csv format, so the files were opened in EXCEL and saved as CSV files. In opening use DELIMITED, deselect TAB, select COMMA and OTHER - *.

It was necessary to choose TEXT for the time on the 2nd page of the text import wizard.

The columns in the RAW files are:

Date UTC

Time UTC

Identifier – not needed in processing

Remote temperature (SBE38)

Lab (TSG) temperature (SBE45)

Conductivity

Flow in ℓ/minute

Fluorescence in V

Latitude  - decimal format

Longitude – decimal format

Checksum, - not needed in processing

The spreadsheet was adjusted as follows:

- 2 lines of headers were added – channel names and units.

- The file was then saved in CSV format – the name is not critical at this point.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column M. Then a concentration value was calculated in column F using offset -0.04 and scale 13.1 as determined in the most recent recalibration of the fluorometer.

- Copy/Special Paste was used to save the values and then the voltage channel was removed.

- A file break column was filled with the cruise #-data/time info from the original file name so the name of the converted file will have format YYYY-###-YYYYMMDD-HHMMSS.

- Time is a problem – when converting from RAW choose TEXT but once open in EXCEL set Time Format  to HH:MM:SS and save the file again. (Any time the file is opened in EXCEL the time format may have to be set again before saving.)

The files were then converted to IOS format with header info added.

CLEAN was run to rest the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. The time gets set to Start Time whether or not you choose that on the final form. A record number was also added to enable averaging (for use in comparison to CTD files).

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

DERIVED QUANTITIES was run a second time to derive sigma-T but the wrong temperature was used and this variable is not usually archived, so it will be removed later.

An initial run produced odd salinity traces that were due to a truncated conductivity format. After that as fixed the salinity traces looked excellent.

REORDER was run to move the Julian date channel so it is not chosen for time-series plots

a.) Plots

A track plot was produced and added to the end of this document.

Time-series plots were examined on screen. Note that for the first file TIME ZERO had to be changed to the beginning of the day in order to get the plot routine to work.

- There is a file for September 11th but the ship did not move, there was no flow in the loop and the temperature difference between lab and intake was very large. This file will not be processed further.

- There is a file for September 12th that has no flow, the temperature difference looks very large and the intake temperature looks unreasonably smooth. The ship was moving so it will be kept to enable a complete track plot. But temperature, salinity, conductivity and sigma-T should be removed.

- There is a file for September 13th that also has no flow for the first hour, but mostly it looks ok so will just need a little editing.

- There is a lot of noise in the plots of the difference between lab and intake temperature. In fairly quiet sections of the plots it varies from ~0.2Cº to ~0.35Cº except right at the end of the final file, Sept. 27th, when it becomes significantly larger ranging from 0.5Cº to 1Cº.

- The fluorescence trace shows a steady increase in values from about 3.5ug/L to values that were at the maximum for the sensor. At the end of the last file the voltage was 5.0004. There was some variability suggesting a real signal, but clearly there is a background drift that renders these data useless. Even the data from the first day appears to have unreasonably high values compared to the CTD with the CTD having a reading of about 1ug/L while the TSG has 3.5ug/L. Offshore CHL and CTD fluorescence values are very low, as expected, while the TSG gets steadily higher. While the calibration could have drifted that does not explain the high voltage. Either the sensor failed or there was a build-up of biological fouling.

- Salinity traces look good with a few spikes to the low side that are likely due to bubbles, but they are no larger than 0.015psu and some could be due to real changes, so editing will not be applied to them.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db. These were exported to a spreadsheet which was saved as 2018-040-ctd-tsg-loop-rosette-comp.xlsx. There were 59 CTD casts that overlapped with the TSG record but for the first case the TSG loop flow was zero, so that was removed from the comparison.

Bin Average was used to average TSG files over 12 records (1 minute) and standard deviations were calculated.

Head Edit was used to change the format of the standard deviations to a format that can be exported to a spreadsheet.

The data needed for comparison with the CTD data were then exported to a spreadsheet (Time, positions, intake and lab temperatures, salinity, fluorescence and flow rates plus the standard deviations for all except time and positions.)

Date were selected for the times of CTD casts and then added to file 2018-040-ctd-tsg-loop-rosette-comp.xlsx.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no position differences >0.0004º and the median differences were both 0.00000º. So the time and positions in the TSG files are reliable and the matches are good between the two systems. Small errors are caused by averaging and imperfect time matches.

c.) Comparisons

- Comparison of T and S from TSG and CTD data and Rosette samples

|  | Latitude Difference | Longitude Difference | TSG Intake Temp - CTD Temp | TSG Lab Temp - CTD Temp | TSG Sal-CTD Sal | TSG FL/ CTD FL |
| --- | --- | --- | --- | --- | --- | --- |
| median | 0.0000 | -0.0000 | 0.0039 | 0.2763 | -0.0303 | 28.210 |
| average | 0.0000 | -0.0000 | 0.0237 | 0.3272 | -0.0337 | 27.818 |
| stdev | 0.0001 | 0.0001 | 0.1584 | 0.1933 | 0.0181 | 19.863 |
| max | 0.0002 | 0.0003 | 0.6307 | 1.2152 | 0.0259 | 75.872 |
| min | -0.0001 | -0.0004 | -0.6060 | 0.1570 | -0.1010 | 1.199 |

The intake temperature is higher than the CTD temperature by a median of 0.004 C°, with a lot of variability in the differences. But the outliers are all from near-shore comparisons. When only events #16-114 are included standard deviation drops from 0.1584 C° to 0.0259 C°.

|  | Lat Diff | Long Diff | TSG Ti-CTD Temp | TSG Tlab - CTD Temp | TSG Sal-CTD Sal | TSG FL /CTD FL |
| --- | --- | --- | --- | --- | --- | --- |
| median | 0.0000 | 0.0000 | 0.0032 | 0.2706 | -0.0286 | 29.612 |
| average | 0.0000 | 0.0000 | -0.0002 | 0.2723 | -0.0311 | 29.045 |
| stdev | 0.0001 | 0.0001 | 0.0259 | 0.0582 | 0.0094 | 19.109 |
| max | 0.0002 | 0.0001 | 0.0230 | 0.4579 | -0.0120 | 75.872 |
| min | -0.0001 | -0.0004 | -0.6060 | 0.1570 | -0.0508 | 2.585 |

The Intake temperature is higher than the CTD temperature by a median value of 0.003C° and standard deviation of 0.026 C°.

The Lab temperature is higher than the CTD temperature by a median of 0.27C° and standard deviation of 0.058C°. This implies heating in the loop of about 0.267 C° when the ship is stopped. The heating appears to be roughly related to intake temperature as expected, since there will be little heating when the intake temperature is close to the ambient temperature of the loop environment.

To see how the differences vary while moving, time-series plots were examined. There is no obvious difference between times when the ship was stopped and when it was underway. Most of the variability in loop heating appears to be due to intake temperature rather than whether the ship is moving or not. The noise level in the loop heating is much noisier while moving which is likely due to real variations.

The TSG salinity is lower than the CTD salinity by a median value of 0.029psu and standard deviation of 0.009psu. However, a plot of differences against time shows a drift with time. This could be weather related with more bubbles later in the cruise, or it could indicate calibration drift. It is noted that the differences are fairly steady from about P3 to P17. Rough seas could lead to more bubbles in the loop.

TSG Fluorescence values are obviously unreliable starting out somewhat higher than the CTD fluorescence but by the end of the cruise being about 60 times the CTD values. There is a sharp increase between events #30 and 37, but even before that the values are far out of line with CHL levels.

Problems have been noted in the past that might have been due to calibration drift, but that cannot explain these data since the voltage shows the same increase reaching 5V in offshore waters with low CHL. So either the fluorometer is not performing well or possibly over time there was a problem with contamination of the loop or the connection to the fluorometer.

- Comparisons with loop samples

There were 24 loop Salinity and CHL samples. TSG values were extracted from the AVG Files for the times of the loops.

As noted in the comparison with CTD data, the TSG fluorescence data look bad and get worse through the cruise with values of TSG being greater than those of the extracted CHL loop samples by a ratio of from 4.6 to 256.1.

The salinity differences are also consistent with what was found in the CTD comparison with differences becoming larger as time went on, ranging from being low by from 0.01psu to 0.07psu when 2 outliers are excluded. However, from events 27 to 63 the differences are quite steady at about 0.03psu (with the exception of 1 outlier). The median difference shows the TSG to be reading low by 0.034psu. The differences late in the cruise are larger but all occurred while the ship was underway. The differences found when the ship was stopped show less variability than those taken while underway but that may reflect the greater range of conditions sampled while underway. Whether the variation in differences is due to sea conditions or calibration drift is not obvious. The larger of the 2 outliers came from Juan de Fuca Strait where the near-surface gradient was probably large so a small mismatch in depths could be significant.

Finally, rosette samples from about 5m were compared with the TSG salinity and chlorophyll. The loop CHL values were about 85% those of the rosette samples and salinity samples were within 0.005psu with a median difference of 0.0006psu. This is good agreement especially given there is averaging of both sources and likely some vertical and temporal offset.

For details on comparisons see document 2018-040-ctd-tsg-loop-rosette-comp.xls.

d.) Calibration History

- During 2018-026 this system was used for the first time. The intake temperature sensor read higher than the CTD temperature by a median of 0.03C°. The lab temperature was higher than the intake temperature by more than expected based on previous experience with heating in the loop on the Tully. The salinity traces were quite smooth with only a few very small spikes. Salinity read lower than the CTD. The differences were not sufficiently consistent to allow recalibration. The data were archived but with a warning that there was insufficient calibration sampling to assess quality.

- The results from 2018-025 suggest that there were problems with flow in the loop.

- During 2018-034 Salinity was lower than the CTD by a median of 0.030psu. Lab temperatures were higher than CTD temperatures by a median of 0.66 C°. The intake temperatures were high by a median of 0.072 C° using all data or 0.058 C° in a “quiet section”. Fluorescence from the TSG was lower than the CTD fluorescence except in areas where CHL was low. The comparisons have a lot of scatter and recalibration was not attempted. The data were archived but with a comment about the intake temperature being further from the CTD data than expected and heating in the loop being higher than when the SBE21 was in use. Because of the limitations only 3 decimal places will be shown for temperature and salinity.

e.) Conclusions

1. The TSG clock worked well.

2. There were 58 points of comparison with CTD. The comparisons were similar to or much better than found in previous uses with the exception of the TSG fluorometer.

3. The intake temperature was higher than that of the CTD by a median value of 0.0032C° (standard deviation 0.026C°) which is as small a difference as can be expected.

4. The difference between the lab temperature and the intake temperature (i.e. heating in the loop) was 0.276C° with a standard deviation of 0.058 when the ship was stopped. Heating in the loop is dependent on the intake temperature. When moving there is a lot of noise in the differences but average values look similar. This result suggests a more consistent flow in the loop than in previous uses.

5. Salinity was lower than that from the CTD by a median of 0.032psu and lower than loop samples by a median of 0.034psu, which is similar to the result of the previous cruise. The differences from the loop were somewhat higher late in the cruise and were gathered when the ship was moving. So it could be that there was some calibration drift or that conditions were rougher leading to more bubbles in the loop which lowers the TSG salinity.

6. Fluorescence data look bad and that channel should be removed.

7. Editing is only needed to remove temperature and salinity data from all of the Sept. 12th file and from the initial part of the file from Sept. 13th and the last few records of the file from Sept. 27th when the flow was not on.

8. The chlorophyll and salinity loop samples compare well with the rosette samples.

g.) Editing

CTDEDIT was used to pad data in channels Temperature: Lab, Temperature: Intake and Salinity:T0:C0 from parts of the files from Sept.13th and Sept. 27th with 0 flow rate.

h.) Recalibration

No recalibration was applied.

i.) Preparing Final Files

REMOVE was used to remove channel Pressure, Fluorescence:URU:Wetlabs, Temperature:Difference and Record #. For the file from Sept. 12th the Temperature:Lab, Temperature:Intake, Conductivity, Sigma-T and Salinity:T0:C0 channels were also removed since there was no flow in the loop.

Files were opened in Ultraedit to do further padding after REMOVE.

REMOVE was run a second time to remove channel Sigma-t.

There were some problems with header entries for times that were resolved by running ADD TIME CHANNEL again to set the start time to UTC 2017/12/31 00:00:00. (The record time was taken from the separate DATE and TIME channels.)

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

Sampling System

---------------

The seawater is drawn from a 20-cm x 56-cm standpipe, flush with the hull

and located 4.5 metres below the water line, and pumped approximately 50

metres through a 5-cm stainless steel insulated pipe up to the main lab.

Between the intake and the pump is a 24 x 24-cm stainless steel strainer

(1/2-cm grid). The pump is a Moyno model 2L6 stainless steel progressive

cavity pump with flow and pressure control.

TSG Data Processing

-------------------

Comparisons were made between the TSG data and co-incident CTD casts, rosette

samples and loop samples. These indicate that the intake temperature is high by

roughly ~0.003C° and salinity is low by roughly 0.03psu.

Fluorescence data from the TSG were removed from the files because the values were

unbelievably high and drifted higher through the cruise.

No recalibration was applied as the differences found may be due to factors other

than calibration drift.

See report 2018-040_Processing_Report.doc for more details.

The files were saved as 2018-040-DDMM-HHMMSS.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and the salinity channel in file #13 was found to have improper pad values; this was fixed.

The file for September 12th will not be archived since it has only positions and time and the ship moved very little during the time the TSG was running.

##### Loop File

The Chief Scientist provided file 2018-040 Loop log.xlsx which included event numbers, sample numbers and what was sampled.

That file was used to prepare file 2018-040-loop data.xlsx by adding times and dates, latitude and longitude based on the log entries, selecting the end time of samples taken during rosette casts.

Data from the analysts were then added to this file which was saved as 2018-040-loop data.xlsx.

The sampling method column was entered as USW.

The columns were arranged in the order required for the 6-line header used to prepare the loop file.

The dissolved oxygen samples from the loop were not taken as part of the core sampling so will not be included in the loop file.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2018-040-che-surface.csv.

The Oxygen:Dissolved and Oxygen:Dissolved:SBE channel in mass units were included and Draw Temperature.

Data from below 7m were removed as well as records with no analysis result.

A sample method column was added. ROS was entered for the method.

The Start Time was copied into a second column and the first was formatted for date and the second for time.

Times were corrected for rosette samples to match the end of casts.

The data were sorted on event number, then pressure.

A model 6-line header was added to the file and columns were rearranged as needed.

That file was saved as 2018-040-surface-6linehdr.csv.

The loop data file were added to the 6-line header file.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on event # and pressure.

Plots were made of positions and times and 2 errors were found and corrected:

- In the log Loop 78 is said to be from 50° 59.99’ but that should be 49° 59.99’.

- In the log Loop 60 is said to be from 08:16 on Sept. 17. The position indicates it was between 14:43 and 19:03 on the 17th and closer to the first of those times. So it is likely that the time is in PDT which would correspond to 15:16UTC on the 17th. A check of the TSG file confirms that the log position matches the time 15:16UTC.

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

Reorder was not needed as the channel names were already in the usual order.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2018-040-surface.loop. The track plots look reasonable and plots of temperature and salinity versus longitude looks reasonable.

General Comments

PAR off: 42, 48, 53, 58, 59, 61, 70, 74, 76, 80, 83, 84, 85.

Casts with bottle fired out of order: 23,37/38,58,67,89,90,92,94,105,109,114,122.

Sample numbers not used: 287, 660-662, 685-696.

Particulars

5. Pumps off – No CTD file prepared. Pumped channels removed from bottle file.

11. Niskin 1 closed at the wrong depth.

15. Station name wrong in header– change to P3

24. Severe spiking in DO in upper 600m of upcast but not during stops.

25. Depth written in log is not good – header depth is ok.

33. bl file empty – one created manually. Depth in file is not good – log reading 2336 is good – header corrected.

34 to 126. cruise number corrected from 2018-40 to 2018-040

36. Niskin 17 closed by mistake, no sample, line not needed in CHE file.

37. Niskin 17 didn’t close at 60m.

38. Run just to close Niskin 17 at 60m. No down 10 and back. Join to 37 bottle file.

48. DO sensor changed but old con file used.

50. New DO sensor and new con file.

59. Fluorometer cable cleaned and pins re-lubed.

62. Station name P18-19 should be PAR – header corrected.

63.Niskin 1 closed just for bulk water – no sample #, not needed in CHE file.

67. Niskin 1 not closed at 5m. Join with 68

68. Cast to close Niskin 1.

90. Depth in file corrected to match log entry.

92. Niskin 8 closed by mistake at 250m, so the 200m sample (#718) was taken from Niskin 19. Niskin 8 not needed in CHE file.

124. Niskin 21 failed to trip

126. Niskin 21 failed to trip again. CTD did full downcast, then came up to 80m, returned to 160m then up to surface.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0585 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 5724 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 4434 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5725 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 1118 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 7Aug2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  | 1483 |  | 30Sep2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  | 997 |  | 30Sep2017 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3641 |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0585 |  | 18Sep2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  | 21Mar2016 |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 2Aug17 | Factory |  |  |
| Conductivity | 0620 | 2Aug17 | Factory |  |  |
| Temperature SBE38 | ? | ? | Factory |  |  |
| WETStar Fluorometer | 953 | 5May2017 | Factory |  |  |