## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2018-059

Agency: PBS, Ecosystem Sciences Division, Nanaimo BC

Chief Scientist: Neville C. 		 	Platform: Sea Crest

Location: Strait of Georgia			Project: Strait of Georgia Juvenile Salmon

Date: 21 June 2018 –5 July 2018

Processed by: Germaine Gatien

Date of Processing: 3 April 2019 –  5 April 2019

Number of original HEX files: 31	Number of CTD files: 	18

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176, SBE pH sensor 0852 and pressure sensor #464.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with lots of comments. There was no bottle sampling.

Header information was not entered in the raw files; they were added based on the log book entries.

The raw data from events 89 to 170 (12 files) could not be found.

No calibration sampling was done during this cruise and there is little information from other cruises that have used this equipment. Comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. There was salinity sampling during 2018-035 and 2018-036 that suggest salinity was within 0.003psu. There was no dissolved oxygen sampling for any of the cruises. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05. That result was used to recalibrate the dissolved oxygen from this cruise, but this must be considered a rough approximation. We do expect these sensors to drift towards lower values with time.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were standard except that the standard changed after they were collected to include a 3-digit cruise number.

2. Preliminary Steps

A copy of the Daily Log was provided.

There was no description of the deployment method.

The cruise summary sheet was completed.

No configuration file was provided but one used for a previous cruise contained the correct information for the sensors used and there was no time between the cruises for a factory visit. It was saved as 2018-059-ctd.xmlcon.

There were 31 CTD casts recorded in the log book but only 19 raw files provided and one of those (#28) contained only a few surface data points. A note was sent that some files might be missing but might still be in the laptop but no files from this cruise could be found. The missing files are between events #89 and #170.

## 3. Conversion of Raw Data

Con file 2018-059-ctd.xmlcon was used to convert the files.

After conversion the file names were changed to the new standard format.

Examination of files showed that the CTD was lowered to about 2m and held there for between 1 and 3 minutes.

Plots show that the channels all produced reasonable values. The descent rates look moderate but quite variable. Temperature and conductivity look normal with no large spikes.

The CTD was left running after cast #25; it was on for 14 hours. The file is large but the data in it are odd. There is a sensible downcast. The upcast is interrupted by a sudden change to near-zero pressure and then return to a deep value. There follow several other strange segments. Since there is a full downcast, the upcast will be removed. Program Section was used to choose the data from scans #0 to 2800.

4. WILDEDIT

There were no significant spikes in the data, so WILDEDIT was not run.

5. WFILTER

Tests were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. WFILTER was run using a cosine filter, size 5 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

10. Checking Headers

None of the casts had positions added. A spreadsheeet was provided with the positions, event numbers, station names and water depths. This was saved as a CSV file and format were adjusted to enable merging with the CLN files. The file was saved as 2018-59 header-mrg.csv.

The positions were those of the NET files that followed the CTD casts, so positions were adjusted to match the CTD. The differences were mostly insignificant.

Routine Merge CSV Files to Headers was run to add that information.

A track plot was produced and an error was discovered in the position of event 59, station 34 – it should be 124° not 125°.

A cross-reference list was produced and checked against the log. It was found that file 2018-059-0043 should be called 2018-059-0041 and 2018-059-0054 should be 2018-059-0053. Those errors were fixed and the merge was repeated and the track plot looks fine.

Track plots were added to the end of this report (showing event #s and station names).

While the log times are in UTC, the data files are in PDT. ADD TIME CHANNEL was used to add 7 hours to the header times. The times are within 1 or 2 minutes of the log times.

The next step is to remove the data collected during soaks at 10m. Plots were made to determine how many records needed to be removed and CLIP was run to remove records from each cast. Plots were made after this step and adjustments made until started at an appropriate level and eliminated most of the soak period.

HEADER CHECK was run. There were some negative values in fluorescence and pressure; most will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of 0.6db before CLIP was run and all the associated salinity values were < 22psu.

To check the accuracy of the pressure sensor plots were made of the end of casts. Conductivity dropped to near-zero when pressure was between -0.4db to +0.8db with most values within 0.4db of zero. Those values are well within the specifications of ±1db. The minimum pressure recorded before removal of the soak data was -0.425db.

##### 11. SHIFT

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. The alignment generally looks good, though as usual there are parts of profiles where a larger or smaller offset is seen in dissolved oxygen relative to temperature. Fine-tuning alignment is unlikely to improve both. No further adjustment was made.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. The same was true for this cruise with a setting of +10 records looking slightly better than +15. SHIFT was run to advance the pH signal by 10 records.

Conductivity

For the 7 cruises immediately after the last factory calibration a setting of +1.2 records was found suitable to align the conductivity with pressure. For 2018-96 that setting worked for most casts, though no setting seemed to help for a few. For 2018-035 and 2018-036 applying a setting of +1.2 made the alignment between the T and C traces worse and the resulting data looked much worse in T-S plots. A setting of -0.2 records made a slight improvement. For cruise 2018-93 which immediately preceded this cruise a setting of -0.2 records worked best and for 2018-104 which came later a setting of +0.6 looked best.

The 2018-059 data mostly look good without alignment and tests show no improvement with any setting except for casts #62 and #66 which did look better with a +0.6 record shift.

SHIFT was applied to conductivity only for casts #62 and 66 for which a shift of +0.6 records was applied.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records near the top and bottom of many casts and to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts were edited lightly.

Notes of editing details were made in the headers.

T-S plots were examined after this step and no significant problems were found; there are small unstable features but this is an area where some instability is expected.

16. Other calibration checks

Sensor History – These sensors were used for 13 cruises before this one since they were last calibrated. They were also used on 2 cruises that have been processed and occurred after this cruise (2018-073 and 2018-104). No problems were found in the pressure. There was little calibration sampling available but comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. There was salinity sampling during 2018-035 and 2018-036 that suggest salinity was within 0.003psu. There was no dissolved oxygen sampling for any of the cruises. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – All data fell within local climatology for the 10 casts for which it was available.

Post-cruise calibrations – None were available.

17. CALIBRATE

CTD Salinity and pressure will not be recalibrated.

There was no dissolved oxygen sampling but the results of a study on the same sensor during 2018-05 will be applied to these data.

CALIBRATE was run using file 2018-059-recal1.ccf to apply the following correction to the Dissolved Oxygen:

CTD Dissolved Oxygen Corrected = 1.05* CTD Dissolved Oxygen

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

There are a few slightly negative fluorescence values at depth so CLEAN was run to replace those with pad values.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 75 to 140%. The lowest saturation rates were in narrow channels (stations 22 and 24) where the water was very well mixed vertically. The highest values were in Jervis Inlet (stations 5 and 6). All other casts had values between 110% and 125%  These values look reasonable but tell us nothing useful about the DO calibration.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There was no salinity calibration sampling. Previous cruises using this equipment

had limited sampling but suggest that salinity is within 0.005psu.

There was no dissolved oxygen sampling during this cruise and most other cruises

using this equipment since it was last serviced. There was an inter-ship comparison

in June 2017 for 1 cast only that suggested that dissolved oxygen values were

likely low by up to 4%. A comparison with an SBE911+ cast at the same site during

a February 2018 cruise was limited to the top 50m but suggested that dissolved

oxygen was reading low by at least 5%.

Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples. It is recommended that

extracted chlorophyll values be used where available.

For details on the processing see document: 2018-093_Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files The header check was rerun and turned up no errors.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PARTICULARS –

An email was received saying that data were missing and might still be on the laptop. No such data could be found.

25. The CTD was left on overnight. File is huge and upcast is very odd – looks like bits of downcast got copied into other parts of the file. File was clipped to include only the initial downcast.

41. Saved as 2018-059-0043.hex -  should be called 2018-059-0041.hex

53. Saved as 2018-059-0054.hex -  should be called 2018-059-0053.hex

59. Station 34 – longitude is wrong in log - it should be 124° not 125°.

CRUISE SUMMARY

| Cruise ID#: 2018-059 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 21 June 2018 End: 5 July 2018 |  |  |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2018-059

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 7Nov2015 | Factory |  |  |
| Conductivity | 1764 | 23Feb2017 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 11Nov2015 | Factory |  |  |
| pH | 0852 | 8Mar2017 | Factory |  |  |
| Pressure | 0464 | 28Feb2017 | Factory |  |  |