## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 20 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2018-089

Agency: OSD

Location: Boundary Pass / Juan de Fuca Strait  / Haro Strait

Project: Moorings, Ocean Protection Plan, Marine Environmental Quality Program

Party Chief: Vagel S.

Platform: Vector

Date: 13 April 2018 – 18 April 2018

Processed by: Germaine Gatien

Date of Processing: 22 October 2018 – 25 October 2018

Number of original HEX files:  	31		Number of CTD files:  (22 SBE911, 9 SBE25)

Number of original TSG files: 	1		Number of processed TSG files:  1

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the primary pump, a SeaPoint Fluorometer (#3641) on the secondary pump, a QSP 200L4S PAR sensor (#4565), an SBE pH sensor (0692) and an altimeter (#62355).

CTD #404 with strain gauge pressure sensor (#482) was used without a rosette. Attached were a WetLab Eco fluorometer (#4185), a pH sensor (0852) and a SBE 43 DO sensor (#1176)

A thermosalinograph (SeaCat 21 S/N 2488) was used with no intake temperature sensor or flow meter.

Seasave version 7.26.2.13 was used for acquisition of SBE911 data.

Software Version 1.59 was used for acquisition of SBE25 data.

The data logging computer was the Vector Laptop.

The deck unit was a Seabird model 11+ #0619.

# SUMMARY OF QUALITY AND CONCERNS

Two CTDs were used. For 9 casts when conditions were rough an SBE25 was used; for 22 casts an SBE911 was used. No rosette sampling was done.

For the SBE25 casts it was very helpful that the headers had station name, water depth, latitude and longitude entered so that the information was automatically entered into IOS headers when the data were converted to IOS Header format. The file names for the SBE25 data were puzzling as they appeared to be in standard format but did not correspond to event numbers in the log book. With the chief scientist’s help this problem was resolved and file names were changed appropriately.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples.

There was no calibration sampling. The dissolved oxygen and salinity channels for the SBE911 were recalibrated using the results of cruise 2018-026. SBE 25 salinity was not recalibrated based on evidence from other cruises before and after 2018-089 that suggest it is within ±0.003psu.

The SBE 25 dissolved oxygen values were extremely low as were DO surface saturation rates. Similar results were found during the cruise that preceded it, 2018-029, when there were nearby casts using an SBE911 to confirm such values were very unlikely to be real. The sensor behaved as expected during the 2 cruises that followed this one, so the problem was likely to be an issue with how the sensor was deployed during the 2 back-to-back cruises. The DO channel was removed from the SBE25 casts.

Both CTDs indicate that salinity was below the climatology minima in the top 5 to 50m in the western Strait of Georgia and the offshore casts. Temperatures were all within the climatology.

There was 1 thermosalinograph file. There was no flow meter, loop sampling or intake temperature sensor. Recalibration was based on comparison with 2 different CTDs and the history of the sensors. Temperature and salinity values must be considered approximate and formats were chosen to reflect that.

# PROCESSING SUMMARY

PART I – SBE911

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book was obtained.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- SBE911: There are processed data available for 2 cruises, 2018-001 and 2018-029, before this mission and 2 after it, 2018-039 and 2018-026.

- SBE25: There are processed data available for many cruises before this one and 2 after it, 2018-035 and 2018-036.

- The calibration control files were checked. There were no errors in the SBE25 file. The only errors in the SBE911 file were a missing date for transmissivity and a missing update to factor E in the dissolved oxygen sensor parameters for the SBE911. The default value is 0.036, but a value of 0.0363 was entered based on hysteresis tests during 2018-01. The corrected file was saved as 2018-089-sbe911-ctd.xmlcon. For the SBE911 the PAR and pH sensors were removed for casts 11-42, but only one xmlcon file is needed – the pH and PAR channels can be removed later.

##### Conversion of SBE911 Files from Raw Data

The file names were changed to the new standard, with a 3-digit mission number.

All files were converted using file 2018-089-sbe911-ctd.xmlcon. File 2018-089-0039 would not convert at first due to some format errors in the headers in the HDR and HEX files. Once those were fixed, conversion worked.

The Tau correction was chosen but not hysteresis correction since there all sampling was above 300m. Depth was included in the conversion.

A few casts were examined and all expected channels are present and the traces look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. That setting was used for other 2018 cruises using this sensor.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 2 casts to find differences between the pairs of temperature, conductivity and salinity channels. The differences were very noisy and the casts very shallow, but the results are similar to other cruises before and after this one that used the same equipment.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0072 | 1000 | -0.0005 | -0.0002 | -0.0023 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0059 | 300 | ~0 | -0.0004 | -0.004 | High, V.Steady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |
| 2018-039-0044 | 1000 | -0.0009 | -0.0005 | -0.0047 | High, XNoisy |
| 2018-039-0088 | 1000 | -0.0003 | -0.0004 | -0.0043 | High, VNoisy |
| 2018-039-0110 | 1000 | -0.0002 | -0.0006 | -0.0065 | High, VNoisy |
| 2018-089-0001 | 260 | -0.0002 | -0.0004 | -0.004 | High, Steady |
| 2018-089-0035 | 235 | +0.0004 | -0.0004 | -0.004 | High, Steady |
| 2018-026-0031 | 1000 | -0.0007 | -0.0004 | -0.004 | High, XNoisy |
| 2018-026-0062 | 1000 | -0.0002 | -0.00035 | -0.004 | High, XNoisy |
| 2018-026-0072 | 1000 | -0.0004 | -0.0003 | -0.004 | High, VNoisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Before running the header checks the SBE25 casts were processed up to the same point so all data could be included. See Part II, sections 19 to 28 for details.

An initial track plot turned up an error in the longitude for cast #38. This was corrected in the IOS and CLN files.

The header check was run and no problems were found.

The cross-reference check was run.

- The times usually differ from log times by a few minutes but the difference is not consistent so this does not appear to be a problem with the computer clock. The file times are set when the computer file is started so they often lead the log times which were likely recorded as the downcast started. Starting files as close to deployment time as possible is recommended to get accurate positions, times and water depths.

- For event #1 the station name was incomplete, so that was fixed.

The cruise track was plotted (2 versions – event 3 and station names) and added to the end of this report; these include the SBE25 casts.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet (available for SBE911 data only). A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. There were many cases where the check value was very high because the header depth was recorded before the cast. Log book values looked more accurate so were used to replace the headers for 10 casts; 1-4, 12-13, 16, 33, 35 and 41.

A surface check was run and shows an average surface pressure of 1.8db for the SBE911 casts which is reasonable for Vector casts.

For the SBE25 casts the average was -0.4db before CLIP with very low salinity values; the range was from -0.6 to +0.5db, so within the manufacturer’s specifications. The average was 1.2db after CLIP.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel with the usual advance of +24 records. Plots show that the fluorescence vertical offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. The traces were extremely noisy towards the end of the cruise.

SHIFT was run twice on all SBE911 casts using -0.5 records for the primary conductivity and -0.4 records the secondary, the same choices as made for 2018-029. Salinity was recalculated for both channels.

pH

Tests were run on 3 casts to determine the best setting to align pH and temperature. A setting of +40 records looks best. SHIFT was run applying a shift of +40 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 2018-01 and 2018-029. The primary salinity was found to be high by ~0.0017 and the secondary low by ~0.0007 during 2018-01. During 2018-029 they were found to be low by 0.001psu and 0.0057psu, respectively. The secondary sensors were chosen for archiving for both cruises. The most reliable of the comparisons was that of 2018-01.

Experience after this cruise –

The sensors were used during 2018-039 and 2018-026. During 2018-039 the primary salinity was found to be high by 0.0012 and the secondary was low by 0.0027. During 2018-026 the primary temperature sensor was different; the primary and secondary salinity values were found to be low by 0.0012 and high by 0.0039psu, respectively. The secondary sensors were chosen for archiving for 2018-039 and the primary for 2018-026.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature data fell within the climatology. Salinity data in the top 5 to 30db were lower than the climatology minima in the Juan de Fuca casts. None of these suggest calibration problems especially since the SBE25 casts showed the same pattern.

Repeat Casts – There were no repeat casts; many were close together but were too shallow to be useful. None of the SBE25 casts were close enough to a SBE911 cast to provide a useful comparison.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The choice of which T-S pair to choose for editing and eventual archiving is not obvious. The primary salinity appeared to be closer to bottle values on 2 previous cruises but had more spikes than the secondary data. For this cruise there is no obvious difference. The secondary was also chosen for 2018-039 which followed this cruise. The primary was chosen for 2018-026 but that was with a different temperature sensor. Secondary channels were selected for editing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. Many unstable features were left unedited as there was no obvious instrumental cause and this is a region of active mixing so instable features may well be real. All files required some editing.

The edited files were copied to *.EDT.

##### Recalibration

Recalibration was based on the results of 2018-26 as it had good calibration sampling.

File 2018-089-SBE911-recal.ccf was prepared to add 0.0039psu to channel Salinity:T1:C1 and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0276 + 0.0477

CALIBRATE was applied to the EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all SBE911 casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

For casts #11 to 42 channels PAR and pH:SBE were also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

Conductivity, Transmissivity, Fluorescence, PAR and pH:SBE data are nominal and

unedited except that some records were removed in editing temperature and salinity.

PAR and pH sensors were only in use during events #1 to #6.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

There was no calibration sampling during this cruise. Dissolved oxygen and

Salinity:T1:C1 were recalibrated based on the results of cruise 2018-026.

For details on the processing see document: 2018-089_Processing_Report.doc.

The Standards Check routine was run and one problem was found and corrected.

The Header Check was run; no problems were found. This was repeated after the SBE25 data were completed so all could be reported together.

Profile and T-S plots were examined and no problems were found. There are some unstable features in T-S space, but those are in Haro Strait so likely real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were about 80% in the Gulf Islands area, typical of areas of active vertical mixing. Values between 95% and 105% were found in western Juan de Fuca Strait which is typical of more open ocean values.

PART II – SBE25

##### Seasave

This step was completed at sea. There were 9 files.

##### Preliminary Steps

The cruise summary sheet was completed.

##### Conversion of Raw Data

All files were converted using configuration file 2018-089-sbe25.xmlcon.

The file names were incorrect having event numbers 1035 to 1043 instead of 22-30 and 38. The names were corrected after conversion. The mission numbers were also changed to the new 3-digit standard.

Plots were examined and all expected channels were present.

The deployment method was to lower the CTD to about 10m and return to the surface, a process taking approximately 2 minutes. Plots show that removing 1000 scans will remove the soak period from most casts without also removing some downcast data.

The CTD was typically stopped at the end of the cast at about 0.2db to 0.5db. Conductivity was very low. The CTD was then brought out of the water and after about 6s conductivity fell to near-zero. Given the pumps would be off by then and that ocean spray might affect the system, it is expected that conductivity would not go to 0 immediately. So the pressure appears to be within the factory specification of ±1db.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

There are frequent small steps in the pressure and some small reversals. Tests run on SBE25 data from cruises 2018-035 and 2018-036 found the best choice to remove the reversals without over-smoothing was a Windows cosine filter size 5. That filter worked well for these casts as well.

WFILTER was run on pressure, depth, temperature and conductivity with cosine file width 5.

##### ALIGNCTD

Tests were run to find the best setting to align dissolved oxygen with temperature traces. The best results were found when the DO channel was advanced by 2.5s.

ALIGNCTD was run to apply a 2.5s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Dissolved oxygen values looked very low at the surface.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

The start times are in PDT, so Add Time Channel was run to add 7 hours to start times.

CLIP was run to remove 1000 records from each file so that data from the initial drop during the soak period do not get selected when DELETE is run. Plots were examined to ensure no useful data were removed.

CLEAN was run to add event numbers.

##### Checking Headers

At this point the CLN files were placed in the same folder as the 911+ files so that header checks could be done together.

No problems were found in the SBE25 files.

##### SHIFT

Fluorescence

The fluorometer on the SBE25 was not pumped so alignment was not needed.

Conductivity

It has been found during other cruises using this equipment that no setting improved most casts and the same seems to be the case for these casts. SHIFT was not applied to conductivity.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; the alignment looks good. SHIFT was not run on DO.

pH

This channel has been aligned by adding either +10 or +15 records in other recent uses of this equipment. For this cruise +15 records looks best. SHIFT was run to apply an advance of +15 records.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to remove corrupted by shed wakes; salinity was cleaned lightly in a few casts. Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined; casts 22 to 30 look fine. There are many unstable features in cast 38 which is the only SBE25 cast from Haro Strait where such features are likely real.

##### Other calibration checks

Sensor History – This CTD has been used many times since it was last serviced at the factory. The limited calibration sampling available suggests that the salinity is within ±0.003psu. The only dissolved oxygen calibration information available is extremely limited, based on comparison to a different CTD and in one case to a cast from a different ship in the same area. During 2018-029 the DO values were found to be very low and nearby casts using a different CTD did not support that those values could be real. Two later cruises using this sensor had normal values for DO so the problem during 2018-029 was likely a problem in deployment.

Comparison of repeat casts – None were available.

Historic Ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature data fell within the climatology. Salinity data in the top 5 to 50db were lower than the climatology minima in the offshore casts. This is a similar result to the SBE911 casts in western Juan de Fuca Strait, so this is not suggestive of a calibration problem.

Post-cruise calibrations – None were available.

##### CALIBRATE

The SBE25 salinity is likely within ±0.003psu based on other cruises, though sampling was very limited so this should be considered a rough estimate. The SBE25 dissolved oxygen data look bad. No recalibration was applied to the files.

##### Filter

A median filter, size 5, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

Dissolved Oxygen was derived in mass units.

Saturation rates were calculated and plots show near-surface saturation from 45% to 62%. These values do not look believable. Cast #38 is in Haro Strait between casts #37 and #39 which used the SBE911. The surface saturation rates were 79%, 45% and 82% in cast number order; cast #38 is way out of line with the others. Also cast #20 from cruise 208-029 (which was a week before this cruise) was close to cast #27 from this cruise. They used different CTDs and the surface saturation from the SBE911 was 81% while the SBE25 was 57%.  These confirm that the problems noted during 2018-029 with the SBE25 DO continued through this cruise. These data need to be removed.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Oxygen:Dissolved:SBE, Descent Rate and Flag channels.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

Conductivity, Fluorescence and pH:SBE data are nominal and unedited except that

some records were removed in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

There was no calibration sampling during this cruise.

Channel Oxygen:Dissolved:SBE was removed due to malfunction.

For details on the processing see document: 2018-089_Processing_Report.doc.

The standards check was run and no problems were found.

At this point the CTD files were moved to the same folder as the SBE911+ files so that quality checks could be run on the combined set.

A cross-reference listing was produced for the whole cruise.

A header check was run including files from both CTDs. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PART III - TSG

##### Thermosalinograph Data

There was 1 hex files. There was no flow meter. The sampling interval was 30s.

The serial number of the TSG was entered incorrectly in the original file. This had to be changed in the raw data file to enable conversion with the correct configuration file.

a.) Checking calibrations

The configuration file used at sea is incorrect.  A correct file was found and saved as 2018-089-tsg,xmlcon.

b.) Conversion of Files

The TSG file was converted to CNV files using file 2018-089-tsg,xmlcon.

##### c.) Conversion to IOS Header Format and Initial Quality Checks

The CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots showed no significant spikes. Salinity increased sharply near the beginning of the cruise at scan 46. It is likely flow had just been turned on. The change in salinity of 7psu between 2 scans is not something likely reflective of ocean conditions. Temperature showed a significant change as well though not as notable. Latitude did not show a concurrent shift in direction. Salinity after cast #46 is reasonably close to that measured a week earlier near the surface in Saanich Inlet.

A track plot was produced and added to the end of this document. There are symbols every 2 hours.

CTD data were thinned to 2m and then exported to a spreadsheet that was saved as 2018-089-ctd2-tsg-comp.xlsx.

The TSG files were opened in EXCEL and the files were reduced to the times of CTD files. There were 30 cases of overlap, 21 from the SBE911+ and 9 from the SBE25.

d.) Comparison of T and S from TSG and CTD data

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. When all casts are included the median differences between latitude and longitude are both <0.0001° and all differences are <0.0006°. This is good correspondence.

Temperature and salinity were then compared for all casts and separately for the 2 different CTDs.

|  | CTD Lat - TSG Lat | CTD Long - TSG Long | Temp TSG-TEMP CTD | SAL TSG - SAL CTD |
| --- | --- | --- | --- | --- |
| average | -0.00003 | 0.00005 | 0.18074 | -0.45119 |
| median | -0.00002 | 0.00008 | 0.14625 | -0.36420 |
| stdev | 0.00010 | 0.00016 | 0.09280 | 0.21661 |
| median SBE911 | -0.00006 | 0.00009 | 0.14850 | -0.36380 |
| median SBE25 | 0.00000 | 0.00003 | 0.13830 | -0.36460 |
| median excluding outliers |  | all | 0.1401 | -0.3311 |
| median excluding outliers |  | sbe911 | 0.1440 | -0.3279 |
| median excluding outliers |  | sbe25 | 0.1317 | -0.3392 |

    - The intake temperature was found to be higher than CTDs by ~0.14 C°. This difference due to warming in the loop is a little lower for the SBE25 which is likely because the water was warmer near the surface for those casts and heating is less when the water temperature gets closer to the ship temperature.

    - The TSG salinity was found to be lower than the CTD salinity by ~0.36psu using all data and ~0.33psu when outliers are excluded.

For more details see 2018-089-ctd2-tsg-comp.xls.

e.) Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the 7th use since then. All except 2018-029 and 2018-030 were Tully cruises with a longer loop than that on the Vector so that more heating is expected.

- 2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than CTD salinity and that appeared to be due to bubbles. No recalibration was applied.

- 2017-06: The temperature in the lab was higher than CTD temperatures by from 0.15Cº to 0.20Cº which looks reasonable based on expected warming in the Tully loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

- 2017-23: The temperature in the lab was higher than CTD temperatures by from 0.16Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.004Cº. The TSG Salinity read lower than the CTD salinity by ~0.29psu and that appeared to be due to bubbles. There were no loop samples. No recalibration was applied to the TSG data.

- 2017-08: The temperature in the lab was higher than CTD temperatures by a median of about 0.018Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. No recalibration was applied to the TSG data.

- 2017-09: The temperature in the lab was higher than CTD temperatures by a median of about 0.028Cº. The salinity traces was full of spikes; salinity was low by a median value of .0064psu or by 0.018psu if only the 10 casts with the lowest standard deviations in the TSG salinity were included. No recalibration was applied.

- 2018-029 (Vector) the TSG lab temperature was higher than the CTD temperature by about 0.13Cº using both 2db and 3db CTD data The TSG Salinity was lower than that from the 2db and 3db CTD data by about 0.30psu. There was no evidence of large bubbles except during one noisy patch. Salinity was recalibrated by adding 0.33Cº.

The following cruise was after 2018-089:

- 2018-030 (Vector) There were data from only 1 day with few CTD casts available for comparison. The TSG lab temperature was higher than the CTD temperature by about 0.28Cº The TSG Salinity was lower than that from the CTD by about 0.06psu. This was out of step with other comparisons so the results of 2018-029 were used to recalibrate.

f.) Conclusions

- The time and position data agree well between the CTD and TSG for both CTDs.

- There are few good calibration checks for this TSG while on the Vector. Looking at 2 other TSGs that have been used on the Vector shows that the temperature in the lab has been higher than CTDs by between 0.1 and 0.2Cº with 0.13Cº being the value noted most often, so the value found for this cruise looks reasonable. It is also close to that found for 2018-029 which occurred just 1 week earlier. So the lab temperature seems to be reliable and subtracting 0.14Cº would give a reasonable substitute for intake temperature.

- For salinity from the Vector TSG the results have been highly variable with values being low by between 0.04psu and 0.93psu with the largest differences associated with the noisiest comparisons. Most recently it was found to be low by 0.30psu during 2018-029 in April. During 2018-030 in June the results were not trusted because there was only 1 day’s data, few CTDs available for comparison and heating in the loop was out of step with previous results for the Vector. The salinity being low by 0.033psu for this cruise is reasonably close to the 2018-029 results. While the large error could be due to water being drawn from higher in the water column, this would also be expected to affect the temperature comparison and it does not. So the large salinity error is more likely due to fresh water getting into the loop or bubbles in the loop or calibration drift. In any case subtracting 0.33 from all vales is a reasonable step to get realistic values.

- A proxy for intake temperature will be created by subtracting 0.14 Cº to the lab temperature.

- Salinity will be recalibrated based on the comparison with CTD data by subtracting 0.33psu, but a note will be added to the header that this is a rough estimate.

- Temperature will be reported with only 3 significant figures and salinity with only 2 significant figures as confidence in the data does not justify the usual 4 figures. This may draw users’ attention to the warning in the headers.

g.) Editing

The ATC file was opened in CTDEDIT to remove salinity and temperature data from the first 49 scans.

h.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2018-089-tsg-recal1.ccf to subtract 0.14Cº to Temperature:Primary and to add 0.33psu from Salinity:T0:C0.

i.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference,  Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names. Formats for temperature and salinity were set to 3 and 2 decimal places only to reflect lower confidence in accuracy.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check time series and cruise track plots were examined and both looked fine.

Particulars

22-30 & 38. Casts run using an SBE25. The raw files were misnamed and had the wrong date in them.

11-42. PAR and pH were removed for rest of cruise.

CRUISE SUMMARY - CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  |  | CTD Calibration Sheet Competed? |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0443 |  | No |  |  | Yes |  |  |  |  |  |
| 2 | SEABIRD | 25 |  | 0404 |  | No |  |  | Yes |  |  |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  | Post Cruise |  |  |  |  |
| Name |  |  | S/N |  | Date |  | Location |  |  | Date |  | Location |  |  |
| Temperature |  |  | 2710 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Conductivity |  |  | 2128 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Temp. |  |  | 4752 |  | 02Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Secondary Cond. |  |  | 2399 |  | 14Sep2017 |  | Factory |  |  |  |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 7Aug2017 |  | Factory |  |  |  |  |  |  |  |
| SBE 43 DO sensor |  |  | 0997 |  | 2Sep2017 |  | Factory |  |  |  |  |  |  |  |
| PAR |  |  | 4595 |  | 16Mar2011 |  | Factory |  |  |  |  |  |  |  |
| Surface PAR |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |  |  |  |
| SeaPoint Fluor. |  |  | 3641 |  |  |  |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 31Aug2017 |  | Factory |  |  |  |  |  |  |  |
| Altimeter |  |  | 62355 |  | 21Mar2016 |  | Factory |  |  |  |  |  |  |  |
| Calibration Information - 0404 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  |  | Post Cruise |  |  |  |
| Name |  | S/N |  |  | Date |  |  | Location |  |  | Date |  | Location |  |
| Temperature |  | 2449 |  |  | 16Dec2015 |  |  | Factory |  |  |  |  |  |  |
| Conductivity |  | 1764 |  |  | 23Feb2017 |  |  | Factory |  |  |  |  |  |  |
| ECO Fluorometer |  | 4185 |  |  | 7Mar2017 |  |  | Factory |  |  |  |  |  |  |
| SBE43 Oxygen |  | 1176 |  |  | 2Mar2017 |  |  | Factory |  |  |  |  |  |  |
| pH |  | 0852 |  |  | 8Mar2017 |  |  |  |  |  |  |  |  |  |
| Pressure |  | 464 |  |  | 28Feb2017 |  |  | Factory |  |  |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/2488

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |