## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 13 Feb 2025 | Corrected file names and formats in BOT files. G.G. |
| 16 March 2021 | Shallowest CTD data available added to surface samples in BOT files. G.G. |

## PROCESSING NOTES

Cruise: 2018-96

Agency: IOS, Ocean Sciences Division, Sidney BC

Party Chief: Young K.	 		Platform: Neocaligus

Location: Strait of Georgia		Project: Strait of Georgia Zooplankton

Chief Scientist: Young K. 		Date: 3 March 2018 –9 March 2018

Processed by: Germaine Gatien

Date of Processing: 7 May 2018 – 24 May 2018

Number of original HEX files:	26

Number of CTD files: 	26		Number of BOT files: 18

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176,SBE pH sensor 0852 and pressure sensor 464.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with a good description of the protocols used for sampling.

A 1.7L Niskin was attached about 3m above the CTD for deep water samples. Samples at 5m, 10m and 20m were sampled with 5L Niskin bottles in a chain on a wire. Surface samples were collected from a 5L Niskin on a pole tripped on the opposite side of the vessel, usually while the CTD was in the water.

A timer was set to 2 minutes. The CTD was lowered to about 10m then raised and held at the surface until the 2 minutes was up. The full cast was then run. This approach enables easy removal of the data from the soak period by removing a set number of records. For 2 casts it was noted in the log that there was a delay in starting the cast but the CTD was near the surface by the 2-minute mark so the regular clipping method worked.

Names were in standard format except for a small and easily correctable error. Header information was entered in the raw files in a format that permitted easy conversion into IOS Header format. That is very helpful.

The sampling scheme was complex as described above. It makes sense to combine the samples from a single site into a single file and the rosette log sheets indicate that was the plan. They can then be combined with CTD data to give a more complete picture of the conditions at that station. It would be helpful to have an indication on each rosette sheet of the CTD event to which the samples relate.

Combining CTD and bottle analysis results was an error-prone process because the sample numbers were in order of collection rather than in order of depth. In future, if it is practical it is recommended that samples be collected in either order of depth or reverse order of depth. If that is not practical it would help to have separate lines for each sample in the event log with sample depths entered.

There were surface samples with no corresponding CTD data either because there was no CTD cast or because the CTD did not get close enough to the surface to provide relevant data. Because there are no pressure readings for some of the bottles, a nominal depth channel was created and values added for all bottles based on rosette sheet entries. Depth:Nominal should be chosen for plotting if you want to include surface samples.

There were 8 deep salinity bottle samples during this cruise. The CTD salinity read higher than the bottles by a median value of 0.0031psu. This small a difference is not considered significant because flushing of Niskin bottles is somewhat inefficient so that the water in the bottles may have come from slightly higher in the water column. A comparison with an SBE911+ cast during a previous cruise also showed that the salinity was reasonably accurate. No recalibration was applied to the CTD salinity.

There was no dissolved oxygen sampling during this cruise or previous cruises using this equipment in 2017. An inter-ship comparison of 2 casts from a June 2017 cruise using these sensors with a cast from a cruise that did include calibration sampling, suggested that dissolved oxygen values were likely low by up to 4%. A comparison with an SBE911+ cast at the same site during cruise 2018-05 had a lot of scatter but suggested that dissolved oxygen was reading low by at least 5%. Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

Extracted chlorophyll samples were taken at the surface and no reliable CTD data are available from that level, so no comparison with SBE fluorescence was possible.

Bottle files were produced combining CTD data with analysis results where possible; for the casts with only surface sampling the files contain only the analysis results with nominal depth.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. One file name had an extra digit; that was fixed.

2. Preliminary Steps

The Daily Log and a spreadsheet of sampling done were obtained. The Niskin bottle used at the surface was leaky.

The deployment method used was as follows: The CTD was switched on, a 2-minute timer was started and the CTD was put in the water. It was taken down to 10m, up to the surface, wait at the surface at least until the 2 minutes were up. The cast was then started. If a deep salinity sample was needed a Niskin bottle was attached 3m above the CTD. The CTD was taken to 10m off the bottom where a Niskin was fired if needed. The CTD was then brought up and switched off as it was taken out of the water. This method is helpful as by removing the first 960 records, the initial soak can be removed. In 2 cases there was a delay which was noted in the log book, so more records may need to be removed for those. This is a good method and the notes about variations are helpful.

The cruise summary sheet was completed.

The parameters in the configuration file provided were checked and those for the fluorometer were found to be wrong. The scale factor is one that is sometimes selected but was not used on the previous cruise so it seemed unlikely. A test conversion showed that it produced some large negative values at depth. The file was corrected and saved as 2018-96-ctd.xmlcon.

## 3. Conversion of Raw Data

The XML files were converted using configuration file 2018-96-ctd.xmlcon.

Plots show that the channels all produced reasonable values. The descent rates look high and fairly steady.

4. WILDEDIT

Since there are no obvious spikes in the data, this step was skipped.

5. WFILTER

Based on tests run for other cruises in this project using the same equipment WFILTER was run using a cosine filter, size 7 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s or 2s. The 2.5s choice produced better results. Plots were examined after this step and the results look good.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it proved best when this equipment was used during 2017-07.

CELLTM was run on all casts using (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

The header times are in PDT so ADD TIME CHANNEL was run to add 7 hours.

10. Checking Headers

An initial track plot looked ok.

A cross-reference list was produced and compared with the log book entries. The only discrepancy was the station name, water depth and position for event #19 which had been mistakenly given the information from the previous cast. That information was corrected in the header of the IOS and CLN files and in the files in the DERIVE and CONVERT folders.

The next step is to remove the data collected during soaks at 10m. CLIP was run to remove 940 records from each cast. While there are 960 scans in 2 minutes there were a few cases where the downcast started a few scans earlier, so 940 scans were removed. Events #6 and #7 had longer waits before beginning the full cast but were at the surface by the 2-minute mark. Plots were made after this step and there appears to be no significant loss of data near the surface.

HEADER CHECK was run. There were some negative values in some channels; most will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of -0.20db; values ranged from -0.6 to +0.8db. Salinity values were all very close to 0 except for the 0.8db reading. It is hard to judge pressure accuracy because pumps are not on as the sensors enter the water and as they leave there is likely water in the system and data acquisition usually stops quickly. After cast #1 acquisition was left on for several hours and the pressure reads about +0.2db. At the end of casts the fluorescence gets extremely noisy at about 0db. Conductivity drops are generally later at from -0.3 to -1db but that may be due to water in the system. Given the pressure sensor accuracy is ±1db, no recalibration of the pressure is called for.

##### 11. SHIFT

Conductivity

During 2017-07, 2017-47, 2017-48 and 2017-49 when the same sensors were used, a setting of +1.2 records was found suitable to align the conductivity with pressure. During 2018-05 no setting seemed to help. Applying a setting of +1.2 made most casts look better, but casts #1, 6, 16 and 22 were made worse overall. This appeared to be due to some T and S reversals that were made unstable by the alignment.

So a shift of +1.2 records was applied to all casts except #1, 6, 16 and 22.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; no further adjustment was found appropriate.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. A setting of +15 records looked best for these data. SHIFT was run to advance the pH signal by 15 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 7 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records near the top and bottom of many casts and records corrupted by shed wakes. It was also used to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts required some editing except for cast #9.

Notes of editing details were made in the headers.

T-S plots were examined and no significant problems were found; there are small unstable features but this is an area where they are expected.

##### 14. Initial Bottle Data Steps

There was no rosette available for this cruise. Samples were taken from a Niskin Bottle mounted about 3m above the CTD and then samples were taken at 5m, 10 and 15m after the CTD cast. A surface sample was taken while the CTD was in the water. To enable searching of bottle data, BOT files were prepared.

Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2018-96-bot-hdr.txt which will be updated as needed during processing. All of the spreadsheets included a precision study.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2018-96 CHL*.xls

SALINITY

Salinity analysis was obtained in spreadsheet QF2018-96SAL*.xls.

NUTRIENTS

Nutrient analysis was obtained in spreadsheet QF2018-96NUTS*.xls.

DIC and ALKALINITY

DIC and Alkalinity analysis data were obtained in spreadsheet QF2018-96_DIC*.xlsx.

The analysis spreadsheet data were combined in a workbook 2018-96-bottle_plus_CTD.xls in spreadsheet “Bottle data”. Because it is impossible to do all the sampling in a single cast, there are often 2 or 3 Niskin sampling events noted in the Daily Science log book for a single site. But it makes most sense to combine them in a single bottle file, corresponding to a single sample/rosette log sheet. Generally the first event number listed on a rosette log sheet corresponded to the CTD event. The analysts used the first event number on the rosette logs. However, for events 22 to 24, the log entry was 23-24. For clarity, it might help to make an explicit entry on the log sheet indicating to which CTD event the bottle data belong. This approach would also allow for the possibility that the sampling might have to be done out of order at some time in future.

To select CTD data to go with the bottle samples, a list of depths sampled (including for the 8 deep bottles) was prepared and edited and bin-averaged CTD files were thinned to those depths. Data were exported and copied to a separate sheet of the workbook, “2018-96-ctd-at-bottle-depths”, and records were removed for which there was no sampling. Those data were added to the bottle data and saved in spreadsheet “2018-96-bottle_plus_CTD”. The data had to be rearranged to match the bottle samples as there are no CTD files to match some of them. Times from the log were added to lines for which this was missing.  For the cases where events were combined the time was selected from the CTD cast. Nominal depths were entered since there are no CTD data for the surface samples.

On one spreadsheet the data were organized with a 6-line header which was saved separately as a CSV file, 2018-96-bottle-plus_CTD_6linlehdr.csv and that was converted to IOS files, 1 per cast.

For some casts there are only CHL samples, no CTD data. For others, there are samples from 0m but no CTD data was available from that depth.

##### 15. Compare

Salinity Comparison

A spreadsheet in workbook 2018-96-bottle_plus_CTD.XLS was used to do a comparison between the bottle salinity and CTD salinity. The CTD was found to be higher than bottle salinity by an average of 0.0806psu using all data; the median was 0.0284 and the standard deviation was 0.2064. Removing some outliers reduced the median difference to 0.0188. Using the 8 bottom bottles only the CTD salinity was high by a median value of 0.0031psu, an average of 0.0033psu and a standard deviation of 0.013.

We expect a poor match at the surface due to slight differences in depth of sampling in the presence of high gradients. There may also be some problems with Niskin bottles not flushing completely. Near the surface this is likely to result in the bottles containing fresher water. For the deep samples the effect is likely smaller, though some effect of poor flushing may still be present as the descent rate was mostly very steady. The CTD salinity may be slightly high, but even slightly flushing inefficiency could account for the difference so no recalibration is appropriate.

For more details see the “sal comparison” sheet in file 2018-96-bottle_plus_CTD.xlsx.

Fluorescence

Chlorophyll sampling was done only at 0m so no CHL-Fluorescence comparison was possible.

16. Other calibration checks

Sensor History – These sensors were used for 7 previous cruises since they were last calibrated. No problems were found in the pressure. There was little useful calibration sampling but comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – The only excursion from the local climatology was a 6db section of cast #52 which had salinity and temperature values below the historic minimum. On a T-S plot this feature looks stable but interleaving is noted. The feature is likely real. There is no evidence of a calibration problem.

Post-cruise calibrations – None were available.

17. CALIBRATE

CTD Salinity will not be recalibrated.

There was no dissolved oxygen sampling but the results of a study on the same sensor during 2018-05 will be applied to these data.

The pressure does not need recalibration.

CALIBRATE was run using file 2018-96-recal1.ccf  to apply the following correction to the Dissolved Oxygen:

CTD Dissolved Oxygen Corrected = 1.05* CTD Dissolved Oxygen

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show most values between 95% and 115%. Lower values were found in the Gulf Islands area at 85%. The values look reasonable.

REORDER was used to get the 2 dissolved oxygen channels together.

CLEAN was run to insert pad values for the few records with slightly negative fluorescence values.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Data Processing Notes:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There were 8 deep salinity bottle samples during this cruise. The CTD salinity

values were higher than the bottles by a median value of 0.0031psu which

is within the expected accuracy. Some of this difference may be due to that

fact that flushing of Niskin bottles is somewhat inefficient so that the

water in the bottles may have come from slightly higher in the water column.

A comparison with an SBE911+ cast during a previous cruise also showed that

the salinity was likely within expected accuracy.

No recalibration was applied to the CTD salinity.

There was no dissolved oxygen sampling during this cruise or previous cruises

using this equipment in 2017. An inter-ship comparison of 2 casts from a

June 2017 cruise using these sensors with a cast from a cruise that did include

calibration sampling, suggested that dissolved oxygen values were likely low by

up to 4%. A comparison with an SBE911+ cast at the same site during a recent

cruise had a lot of scatter but suggested that dissolved oxygen was reading low

by at least 5%.

Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples.

For details on the processing see document: 2018-96_Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no further errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

##### 20. Final BOT file preparation

To enable searching of bottle data, BOT casts were created that contain sample data and CTD data from the preceding downcast. The event number will match the CTD cast.

File 2018-96-bottle_plus_CTD.csv (which has a 6-line header) was updated by thinning final CTD files and exporting data to a spreadsheet so as to capture the effects of recalibrating DO data and applying graphical editing.

The new values were also entered into the bottle comparisons. This demands some rearranging of data to align CTD data with bottle samples.

This process would be easier if the samples were gathered in either order of depth or reverse order of depth, or if that is not practical, that they be given sample numbers in order of depth or reverse order of depth or if the event log spreadsheet contained nominal depths.

Positions for the sites with CTD sampling were taken from the CTD files but for those without CTD casts that information was added based on the spreadsheet provided.

The sheet with the 6-line header format was saved as file 2018-96-bottle_plus_CTD_6linehdr.csv.

The spreadsheet file was converted to IOS Header files for each cast.

The time and date are present as channels as these cannot be changed directly into header entries.

CLEAN was run to add START and END time. The END TIME is identical so the START time so it will be removed later. CLEAN was also used to enter 0 flags where the flag channels are empty and to remove channels with only pad values.

REMOVE was run to remove the DATE and TIME channels.

SORT was used to order the files by Depth:Nominal.

HEADEDIT was used to add comments and to remove the END time and TIME ZERO.

The final files have extensions BOT.

The standards check was run and the only warnings concerned the format for Alkalinity but it is the format chosen by the analyst.

A cross-reference list and header check were run on the BOT files.

A few problems were found:

- Some longitudes had the wrong sign. Those were fixed in the 6-line header spreadsheet and the conversion was rerun.

- The instrument summary was missing so that was added to the Header text file.

.

Plots were made of all BOT casts. One alkalinity sample that was somewhat out of line had been flagged by the analyst. No other problems were found.

PARTICULARS

1. Niskin leaked - didn’t seal – no samples. CTD left on for an hour after cast.

6. Soak time 2.5min.

7. Soak time 3 min

36. Leaky 5L Niskin

CRUISE SUMMARY

| Cruise ID#: 2018-96 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 3 March 2018 End: 9 March 2018 |  |  |  |  |  |  |
| Location: Strait of Georgia Zooplankton |  |  |  |  |  |  |
| Party Chief: Young K. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2018-96

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 7Nov2015 | Factory |  |  |
| Conductivity | 1764 | 22Dec2015 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 11Nov2015 | Factory |  |  |
| pH | 0852 |  |  |  |  |
| Pressure | 1123 | 23May2014 | Factory |  |  |