## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2018-104

Agency: IOS Ocean Sciences Division, Sidney BC

Chief Scientist: Gauthier S.	 		 	Platform: Nordic Pearl

Location: Juan de Fuca Strait and Johnstone Strait

Project: Killer Whale Foraging / Pacific Hake Survey

Date: 31 July 2018 –31 August 2018

Processed by: Germaine Gatien

Date of Processing: 29 November 2018 – 4 December 2018

Number of original HEX files: 36 (1 test)		Number of CTD files: 35

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176, SBE pH sensor 0852 and pressure sensor #464.

# SUMMARY OF QUALITY AND CONCERNS

There was no log book available, but a spreadsheet of events was provided. There were no raw files found for the final 3 casts indicated on the spreadsheet. It is known that there was damage to the CTD during the first of the missing casts so it is assumed that this led to a data acquisition failure.

The format used for header information in the raw files did not permit automatic conversion into IOS Headers. If the following format is used in the HEX files the information will be entered automatically and this will make data processing quicker and less error-prone:

** Latitude: 49 9.781 N

** Longitude: 123 32.970 W

** Station: 39

** Depth: 378 m

The files were saved with non-standard names Cast00.hex to Cast35.hex. We cannot use event #0, so these files were renamed as 2018-104-0001.hex to 2018-104-0036.hex. The digital log was amended to add an event number column and to add notes to indicate that no hex files were found for the final 3 casts. The amended file was saved as “Nordic Pearl Survey_CTD Data_complete-4Dec2018.xlsx”.

There were a number of problems with the deployment method. The CTD was typically soaked at between 2 and 4m for 2 minutes, then returned to the surface. The length of the soaks was good though SeaBird recommend that this be done at 10m. After a return to the top there was a stop that varied from 1s to 15s, then the full cast was run. If the full cast is started immediately after the end of the soak, the data in the top few metres may be poor because the shed wakes created in raising the CTD may not have had time to dissipate. A short wait, say 10s, is recommended especially in very quiet waters where ship drift is slight. In rough conditions this wait could be skipped.

A deployment method that makes data processing more efficient is to set a timer for the estimated time to do the soak and the wait at the surface, something like 2 or 3 minutes. Start the cast as soon as possible after the timer indicates so that removal of soak data can be automated by removing a set number of scans. There will be times when the method fails due to unavoidable delays or rough conditions making a wait at the top dangerous, but those can be dealt with by plot checks.

There was no calibration sampling during this cruise. Recalibration of dissolved oxygen was based on a comparison with an SBE911+ cast at the same site during cruise 2018-05 that suggested that dissolved oxygen was reading low by at least 5%. Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values. Given that these sensors tend to drift towards lower values this correction may be an underestimate.

Salinity was not recalibrated since the few calibration samples during 2017 and 2018 had a lot of scatter or suggested that the error in salinity was small.

The CTD was damaged during event #37 and no raw CTD data were found for that event or the two that followed it.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were non-standard; they had names Cast00 to Cast35.

2. Preliminary Steps

A spreadsheet with even information was provided.

The deployment method varied somewhat, but for most casts the CTD was soaked at about 4m for 2 minutes after which it was returned to the surface and the full cast was run. That may not be sufficient time to ensure that bubbles escape the plumbing and allow the salinity, fluorescence and dissolved oxygen sensors to equilibrate. There are many near-surface data that look unstable in T-S space but it is hard to judge whether these are due to insufficient soaking or reflect real conditions. Many of the sites are near-shore and/or in areas of active mixing where instabilities are common.

The time at which the full downcast started varied, so there is no simple way to remove the soak period data. It is recommended that a timer be set to a time that is sufficient for a soak and return to the surface, say 2 or 3 minutes – then start each cast when the timer indicates so that removal of soak data can be automated. There will be times when this method fails due to unavoidable delays or rough seas making a wait dangerous, but those can be dealt with by plot checks.

The cruise summary sheet was completed.

A configuration file was provided and was saved as 2018-104-ctd.xmlcon. The parameters entered were correct.

## 3. Conversion of Raw Data

The HEX files were converted using configuration file 2018-104-ctd.xmlcon.

Plots show that the channels all produced reasonable values.

From the upcasts it appears that the surface is between -0.2db and + 0.2db.

4. WILDEDIT

There were no significant spikes in the data, so WILDEDIT was not run.

5. WFILTER

Tests were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. WFILTER was run using a cosine filter, size 5 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. One cast was missing its latitude header; that was added and the file was reconverted.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

Merge Headers was used to add latitude, longitude, water depth and station names.

ADD TIME CHANNEL was run to add 7 hours to all start times.

10. Checking Headers

A cross-reference list was produced and checked against the event spreadsheet; no errors were found.

The track plots looked fine and were added to the end of this report.

The next step is to remove the data collected during soaks at 10m. CLIP was run to remove initial from each cast based on plots. Plots were made after this step and a few needed further clipping.

HEADER CHECK was run. There were some negative values in dissolved oxygen, fluorescence and pressure; most will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of 0.6db with a maximum of 1.1db after CLIP was run.

Examination of the end of casts as the pressure goes through 0 shows fluorescence and conductivity dropping rapidly as pressure varies from -0.2db to +0.2db. Those values are well within the sensor specifications of ±1db. The minimum pressure recorded was -0.5db.

##### 11. SHIFT

Conductivity

During 7 other cruise when the same sensors were used, a setting of +1.2 records was found suitable to align the conductivity with pressure. During 2018-96 that setting worked for most casts, though no setting seemed to help for a few. For 2018-035 and 2018-037 a setting of +1.2 made the alignment between the T and C traces worse and the resulting data looked much worse in T-S plots. A setting of -0.2 records made a slight improvement. Tests for 2018-093 showed no improvement with any setting.

For this cruise there was a slight improvement when SHIFT was applied by advancing the conductivity by +0.6records. That setting was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. The alignment generally looks good, though as usual there are parts of profiles where a larger or smaller offset is seen in dissolved oxygen relative to temperature. Fine-tuning alignment is unlikely to improve both. No further adjustment was made.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. The same was true for this cruise with a setting of +10 records looking slightly better than +15.

SHIFT was run on all casts to advance the pH signal by 10 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records corrupted by shed wakes – mostly near the top and bottom of casts - and to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts required some editing. Notes of editing details were made in the headers.

T-S plots were examined after this step. Some near-surface unstable features remain, but these casts come from areas where some instability is expected. The shallow soak may also account for some variability in salinity.

16. Calibration checks

Sensor History – Since they were last calibrated these sensors were used for 14 other cruises. No problems were found in the pressure. There was little calibration sampling available but comparisons with nearby SBE911 casts were possible from 2 cruises and they suggested that the salinity was reasonably accurate. There was salinity sampling during 2018-035 and 2018-036 that suggest salinity was within 0.003psu. There was no dissolved oxygen sampling for any of the cruises. Dissolved oxygen was found to be low by at least 5% when compared to an SBE911 cast during 2018-05.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Where available, profiles of temperature and salinity were examined with a 3-standrd deviation climatology superimposed. All temperature and salinity values fell within the climatology except for 3 casts (event #21, 34 and 35) for which the temperature values were above the maxima around 70 to 100m.; salinity values were below the minima at the same depths. This is a region with high variability and the climatology likely has few observations in mid-summer; moreover a 3-standard-deviation climatology is not really appropriate close to shore. The excursions are not considered indicative of calibration problems.

Post-cruise calibrations – None were available.

17. CALIBRATE

CTD Salinity will not be recalibrated.

There was no dissolved oxygen sampling but the results of a study on the same sensor during 2018-05 will be applied to these data.

The pressure does not need recalibration.

CALIBRATE was run using file 2018-104-recal1.ccf to apply the following correction to the Dissolved Oxygen:

CTD Dissolved Oxygen Corrected = 1.05* CTD Dissolved Oxygen

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

There are a few small negative fluorescence values in deep water, so CLEAN was run on the FIL files to replace those with pad values.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 70% to 120%. The variability reflects varying vertical temperature gradients with low saturation values where surface waters are quite well mixed. This test rarely offers useful results in inshore waters and that is the case here. The 2 casts farthest from shore have surface saturation values of 102% and 94%.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There was no salinity calibration sampling from this cruise. Previous cruises

had limited sampling but suggest that salinity is within 0.005psu.

There was no dissolved oxygen sampling during this cruise and most other cruises

using this equipment since it was last serviced. There was an inter-ship comparison

in June 2017 for 1 cast only that suggested that dissolved oxygen values were

likely low by up to 4%. A comparison with an SBE911+ cast at the same site during

a February 2018 cruise was limited to the top 50m but suggested that dissolved

oxygen was reading low by at least 5%.

Channel Oxygen:Dissolved:SBE was recalibrated by adding 5% to all values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples. It is recommended that

extracted chlorophyll values be used where available.

For details on the processing see document: 2018-104_Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and a few header errors were found and corrected.

The header check was rerun and turned up no further errors.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

Particulars

4. Niskin broke surface, it was then re-lowered for Niskin reset.

CRUISE SUMMARY

| Cruise ID#: 2018-0104 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 31 July 2018 End: 31 August 2018 |  |  |  |  |  |  |
| Location: Killer Whale Foraging / Pacific Hake Survey |  |  |  |  |  |  |
| Party Chief: Gauthier S. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2018-104

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 7Nov2015 | Factory |  |  |
| Conductivity | 1764 | 23Feb2017 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 11Nov2015 | Factory |  |  |
| pH | 0852 | 8Mar2017 | Factory |  |  |
| Pressure | 0464 | 28Feb2017 | Factory |  |  |