#### 2019-01

CCGS Sir Wilfrid Laurier

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### February 5 – February 22, 2019

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from stations P2, P4, P12 and P26 for DMS, DMSPD (dissolved) & DMSPT (total).

  - DMS

Fourteen water samples from various depths (300m, 200m, 150m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at each of the stations P4, P12, P16, P20 & P26.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface).  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in a fridge, in the dark and removed one at a time before analysis.

  - DMSP

Six samples for both DMSPD and DMSPT were collected (200m, 150m, 100m, 20m, 5m, surface).  The only exception to this was P2 where there were no 150m or 200m samples, hence, only 4 samples were collected. Duplicates were taken at 20m.

- Analysis

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Nitrogen for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasorb 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

3.5mL of seawater was collected directly from the niskin into a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.  Stripping efficiency was evaluated at the beginning of the cruise and was proven to be acceptable at over 97%.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by simply treating MQ water as an actual sample.  For example, in the case of DMSPD it was put through a separate funnel and for DMSPT it was added directly to the polypropylene tube.

- Data & Results

  - DMS

There were problems with DMS on this cruise.  Unfortunately, the problem was never solved and, as a result, despite samples being collected and run at stations P12 and P26 the validity of the data is questionable based on the problems encountered with contamination and reproducibility. The data from P2 and P4 is considered valid as the instrumentation issues did not occur until P12. No samples were collected at P20 and  station P16 was missed due to weather.

  - DMSP

The DMSP samples were collected for all stations with the exception of P16 and P20.  These samples will be run back in the laboratory.

- Conclusions

6.1 DMS

The problems encountered with the DMS system were difficult to trace and despite several days of troubleshooting there was no resolution.  Initially the problem started out as one of contamination and the inability to get a clean blank but as time progressed the problem seemed to more of an active site somewhere in the system that was causing the contamination and then subsequent reproducibility issues  More work will have to be done in the lab to try and find the source of the problem.

  - DMSP

No problems to report.