## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 19 March 2025 | ### Updated channel names & formats in TOB & Loop files. GG & SH |
| ### 5 September 2024 | ### Added DOC & TOC SH |
| ### September 2023 | ### Fixed HPLC analysis comments. SH |
| ### 19 March 2023 | ### Added HPLC data. J.R. |
| ### 9 June 2021 | ### Fixed DIC/TA flags. G.G. |
| ### 30 April 2021 | ### Added comments to the phosphate flag changes below. S.H. |
| 16 April 2021 | Added DIC and TA data, and changed phosphate flags for 5 samples in event 16. S.H. |

## PROCESSING NOTES

Cruise: 2019-001

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 5 February 2019 – 23 February 2019

Processed by: Germaine Gatien

Date of Processing: 9 Mary 2019 – 28 July 2019

Number of original HEX files: 91 		Number of CTD files: 91

Number of rosette files:	 77			Number of bottle casts processed: 77

Number of original TSG files: 4			Number of processed TSG files:  16

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1201DR & 1396DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (#3685), a Biospherical QSP-400 PAR sensor (#70613) and an altimeter.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully Laptop.

The deck unit was a Seabird model 11+ #425.

Two salinometers were used, both Guildline model 8400B Autosal: serial # 68572  and #69086.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in excellent order with comments about problems encountered. Sampling notes were provided which are a great help in processing data.

Niskin 21 malfunctioned throughout the cruise.

The deployment scheme for the CTD varied during the cruise:

- About 60% of the casts included a 10m soak after which the CTD was returned to the surface; archiving then began and the full cast was run.

- For Events #67 and Events #74 to #112 there was a 2-minute surface soak before archiving began. After a further short wait (20s to 100s) the full cast was then run. The salinity data appear to have equilibrated adequately after the surface soak. but for some casts there is some evidence of dissolved oxygen data having peak values at about 10m where bubbles may have been released that would normally happen during a 10m soak. This was not the case for many of these casts, but where there are DO maxima between 5 and 15m, they should be considered suspicious.

There were concerns about the transmissometer in use for the first 5 CTD casts; it was then replaced. There is no notable difference in the quality of profiles. Values appear to update at 8Hz producing “steppy” profiles.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 150db

±0.3 mL/L from 150db to 300db

±0.07 mL/L from 300db to 800db

±0.04 mL/L below 800db

There were 4 Thermosalinograph files that covered the time of the cruise. Some covered less than a day while others covered several days. The data were reorganized into single files for each date to enable easy searching. The flow rate was quite steady. The salinity and temperature data were mostly free of spikes except for some sections between Feb 11th and 14th. The ship was moving between station P22 and Dixon Entrance to escape a storm and the spikes are likely related to high sea state and/or high ship speed. There are a few gaps in the temperature and salinity data where the flow in the loop at stopped; during those gaps only flow rate, position and time data were left in the records. Fluorescence data were removed because comparisons with both loop chlorophyll and CTD fluorescence show values were much too high early in the cruise, though they gradually dropped and were close to the CTD values for the last few casts. During the use of this equipment in September 2018 the opposite pattern was seen with TSG fluorescence gradually rising to unbelievable values.

Comparisons were made between TSG data and co-incident CTD casts, rosette samples and loop samples. These indicate that the intake temperature is high by roughly 0.05C° and salinity is low by roughly 0.005psu. The error in salinity is small which implies that Temperature:Lab and Conductivity data are reasonably accurate. Recalibration was applied to all files to subtract 0.05C° from Temperature:Intake and to add 0.005psu to Salinity. Temperature and salinity are shown with 3 decimal places to reflect uncertainty in the calibration.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity, DMS and DMSP data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Since their last calibration most of these sensors have been used on 3 casts (just 2 for the DO sensor), 2 of which have been processed (1 for the DO sensor).

- The calibration control files were checked. Some problems were detected during the cruise. An error in the PAR sensor calibration was repaired before cast #7. The transmissometer was switched before cast #13. The files from events 7 and 13 were saved as 2019-001-ctd1.xmlcon and 2019-001-ctd2.xmlcon.

The deployment procedure varied through the cruise:

- From cast 1 (Saanich Inlet test cast) to cast 61 (P24), cast 69 (DIX4) to cast 73 (DIX1), cast 80 (CH3) and cast 113 (HAK1-C) to the end of the cruise:

The rosette was brought to the surface.  Pumps were turned ON.  Then the rosette was brought down to 10m and back up.  Data started to be archived then, with the rosette at the surface for 30 seconds longer.  Then the cast would start.

- For cast 67 (DIX5), and casts 74 (HECS5) to 112 (OGCH46):

The rosette was brought to the surface.  Pumps were turned ON.  The rosette would sit at the surface for approximately 2 minutes, then data were archived and the cast would start.  (No “down 10 and back”).

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2019-001-ctd1.xmlcon for casts #1 to #11 and 2019-001-ctd2.xmlcon for casts #13 to #143. Depth was included.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and the only problem was a little noise in the primary salinity during cast #54 at about 150m. CTDEDIT was used to clean those outliers. The output files were copied to *.BOT.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Bottles were removed that have not been sampled.

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

At this point lines not required due to no sampling were removed from the SAMAVG files.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-001-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2019-001 CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2019-001chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF 2019-001 OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-001oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file QF2019-001SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 15 to 76 days after collection. The files were simplified and saved as 2019-001sal.csv. That file was then converted to individual SAL files.

51 Salinity:Bottle values required recalculation due to non-linearity in Autosal s/n 68572. See the notes in the above document for more details.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2019-001*.xlsx. This includes a precision study and a comparison of nutrients collected using both polypropylene bottles and plastic tubes for 2 casts. Polypropylene bottles were used because there were not enough plastic tubes due to extra stations being added. The file was simplified, saved as 2019-001nuts.csv and converted to individual NUT files.

DMS –

DMS data were obtained in spreadsheet DMS summary (2019-001).xls. Details on analysis are in file 2019-001 DMS report.doc which includes duplicate analysis. The file was sorted on sample #. This file was converted to DMS files.

DMSP

DMSP data were obtained in spreadsheet 2019-001 dmsp QF summary.xls. This file was simplified, sorted on sample # and converted to DMSP files.

The SAL, CHL, OXY, NUT, DMS and DMSP files were merged with CST files in 6 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Some problems were found:

- Some dissolved oxygen samples had been flagged 3 because of Niskin leaks. The flags need to be applied to all other samples from the same Niskin bottles. The only other variable that had sampling from those bottles were nutrients. The nutrient flags were changed to 3, as needed.

- Events #41 shows a planned nutrient sample but none was found. Flag 9 was added.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were 7 major outliers but 5 were near-surface samples, 4 of which had high standard deviations in the CTD data. The only cases that look suspicious are:

- sample #197 from cast 44 at 1500db. The difference was 0.176 psu and the bottle value looked out of line in profile. Problems were discovered in the analysis with a few errant readings with low salinity having been included in the average. When those were deleted, the value was recalculated with correct readings. The previous value of 34.316 became 34.4576 and the difference was then 0.035psu which still looks pretty poor. The analyst changed the flag to 5 and padded the value.

- Sample #180 from cast 41 at 299db. The difference is 0.04psu. It is not out of line in profile and could be the result of poor flushing. It is not clearly an analysis issue – possibly poor flushing. There was no DO or Nutrient sampling to guide us.

After excluding those outliers and cases where the CTD salinity standard deviation is >0.0008psu and pressure <200db, the primary and secondary CTD salinity channels are low by an average of 0.0019psu and 0.0004psu. Excluding data above 1100db indicated that the CTD salinity was slightly lower relative to bottles than in the previous fit. That would suggest that incomplete flushing of bottles is not a significant effect below 200db.

We must keep in mind that not all of the data were analyzed on the same salinometer. So a comparison was made that separated the data according to which salinometer was in use. This test was run on the primary salinity only. For casts below 700db the bottle values that were recalibrated are low by an average of 0.0007psu while the ones run on the Arctic salinometer are low by an average of 0.0024. So the average difference is 0.0017psu. This may imply that the correction was slightly too small or that the Arctic salinometer was reading slightly high. However, the bottles were collected in different areas and conditions. The difference is within the quoted accuracy of either salinometer. The precision result for the combined salinometers was Sp=0.0013. So the correction appears to be reasonable but possibly a little too low.

For full details for the COMPARE run see file 2019-001-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The data for each DO sensor were assessed separately.

There were 2 outliers large enough to need investigation:

- Event #107: 58m, sample 379 – difference -1.50. The analyst had flagged the sample 4 because of titration problems. The comparison supports that the value is bad. The analyst padded the value and changed the flag to 5.

- Event #127: samples 738 (40m) and 739 (30m) are out of line, #738 is high by 0.67mL/L and #739 is low by 0.45. Profiles are very complex around these depths (30-50m) with the CTD DO showing reversals at 30 and 40m. Given incomplete flushing of Niskin bottles, the reversals can easily account for the bottle at 30m looking different from in situ waters. No flags is recommended for #737. But #738 cannot be accounted for that way even allowing for recalibration of the CTD DO. The analyst had already flagged this sample as 2 due to over-titration, so the high value is likely due to that problem. The analyst changed the flag to 4.

Data below 1000db were excluded from the fit to see if there are signs of hysteresis and there was a hint of hysteresis with the deep points falling a little lower than the shallow ones in a plot of differences versus SBE DO.  There were not as many deep casts as usual for Line P, so this is not as convincing as usual. Test conversions were done to see if there was any improvement if the value of hysteresis factor E was changed in the configuration file. The choice of E = 0.37 went too far, so E=0.365 was tried and looked ok.

After those outliers were excluded plus others based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0299 + 0.0275 (R2 = 0.96)

As a final check of hysteresis, the data between 0 and 3mL/L were divided into above 1000m and below 1000m. Because there are more points near 0 for the data above 1000m, the offset for both above and below was set equal to the one found for the whole set, 0.0275mL/L. With those restrictions, the slope of the fit above 1000m was 0.0326 and below was 0.0333. So any hysteresis is slight.

For events #74 to #112 there was a 2-minute surface soak before archiving began. After a further short wait (20s to 100s) the full cast was then run. The salinity data appear to have equilibrated adequately after the surface soak. but for some casts there is some evidence of dissolved oxygen data having peak values at about 10m where bubbles may have been released that would normally happen during a 10m soak. This was not the case for many of these casts, but where there are DO maxima between 5 and 15m, they should be considered suspicious. Few of the suspicious cases had bottle data available to compare with. But for cast #101, the CTD DO (after applying the recalibration shown above) at 10m on the downcast was ~6.806, on the upcast it was ~6.604 and the bottle value was 6.597mL/L. So the downcast DO appears to be high by about 0.2mL/L. This was the largest peak noticed and for most casts there was no obvious sub-surface maximum.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

When 4 bottles were excluded the ratio of FL/CHL versus CHL had the shape typical of these sensors starting at about 2.5 for very low CHL and dropping to about 1 when CHL approaches 1ut/L. There were no high CHL values. The 4 points rejected were ones where the CHL was very close to zero (≤ 0.02ug/L) rendering the ratio unreliable.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-001-ctd1.xmlcon for casts #1-11 and 2019-001-ctd2.xmlcon for events #13-143.

The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. The transmissometer was changed between events #11 and #13. Both transmissometers have “steppy” profiles due to low sampling rates,. Cast #13 has a smoother transmissivity profile and slightly higher values than #11, but that may well be due to moving further offshore. Supporting that explanation is the fact that towards the end of the cruise there are near-shore casts with noisier profiles and lower transmissivity.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

There was a lot of corruption by shed wakes but few large spikes with just a few points in them. One large spike in conductivity was noted and it was greatly reduced by this step.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. The differences are all small. The salinity differences are close to the difference implied by the results of the bottle comparison, ~0.0015psu.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |
| 2018-030-0050 | 300 | -0.0008 | +0.00003 | +0.0009 | High, Steady |
| 2018-030-0086 | 300 | -0.0013 | -0.00003 | +0.0008 | F.High, Noisy |
| 2019-001-0020 | 1000 | -0.0009 | +0.00007 | +0.0019 | High, XNoisy |
|  | 1900 | -0.0008 | +0.00005 | +0.0015 | High, XNoisy |
| 2019-001-0037 | 1000 | -0.0009 | +0.00008 | +0.0019 | High, XNoisy |
|  | 1900 | -0.0010 | +0.00007 | +0.0021 | Med, XNoisy |
|  | 2900 | -0.0009 | +0.00009 | +0.0020 | Med, XNoisy |
| 2019-001-0049 | 1000 | -0.0005 | +0.00010 | +0.0019 | High, XNoisy |
|  | 1900 | -0.0007 | +0.00010 | +0.0020 | High, XNoisy |
|  | 2900 | -0.0006 | +0.00011 | +0.0019 | High, XNoisy |
|  | 3900 | -0.0006 | +0.00012 | +0.0022 | High, XNoisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run. No errors were found.

The header check was run. No problems were found.

Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. Many problems were found and changes were made to the CLN files, as needed:

- No water depth was entered for the first cast. A value of 178 was deduced from the altimeter and in previous occupations of that site values of 170 to 183 have been found so 178 is reasonable.

- In a few cases the water depth entered in the log book produces a much better result (events 49, 82, 83, 130) For event 126 there is a range of depths given in the log book and using the maximum of those values has a good result while the header has the minimum. Those 5 header entries were changed.

- For a few casts the CTD did not get within 15m of the bottom and the altimetry header was removed (14, 38, 61, 85)

- For some casts the altimetry header looks fine and the check value is not very high. These are likely cases of shoaling so the water depth is variable or there was more vertical movement of the CTD so the altimetry is a little less reliable. No changes were made to the headers.

- In other cases the altimetry header looks fine and there were other casts at the same site with depths entered that lead to a better check value. (events 37, 130)

For the bottle files the same changes were made to the SAMAVG files, as required, and altimeter headers were removed in cases where the cast was deep but sampling was only from the surface. (27, 39). Also the water depth was changed for event #37.

A surface check was run and shows an average surface pressure for the cruise was 2.9db which is reasonable for the Tully; values were higher for the earlier casts when the seas were rough and somewhat lower for the near-shore casts at the end.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on several casts to assess what settings should be used to align conductivity and temperature as judged by the effect on salinity as seen in T-S space. The best settings look like the ones used during 2019-030 when the same equipment was used. The best setting for the secondary conductivity is slightly different than used for 2019-032.

SHIFT was run twice on all SBE911 casts using -0.6 records for the primary conductivity and -0.5 records for the secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature and conductivity sensors have been used on 3 cruises since they were last serviced and the DO sensor #3234 was used for 2 of those cruises. One of the cruises has not yet been processed. During 2018-030 the primary salinity was found to read low by 0.002 and the secondary by 0.001. During 2018-032 there were problems with the salinometer so the comparison was not trusted. During 2018-032 DO was recalibrated using a slope of 1.0324 and offset of 0.0096.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature from Line P fell within the local seasonal climatology. Many of the casts in Hecate Strait and Queen Charlotte Sound had some temperatures lower than in the climatology but the climatology is old and likely has few data from winter. Salinity was almost all within the climatology except for a small feature in the strait of Georgia that had values below the minimum and a section at P11 with salinity too high; the latter looks like a case of a shallower mixed layer than seen by the climatology.

There is no evidence of poor calibration of sensors.

Repeat Casts – There were many repeat casts including some deep enough for a useful test of repeatability. Casts 51 and 54 at P26 started about 2 hours apart and sampled to at least 2000m. At about 2000m differences in temperature and salinity along lines of constant σt were ~0.002ºC and ~0.0002psu. This is excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There was more noise in the secondary channels and more unstable features in T-S space than with the primary channels. The primary salinity was lower relative to the bottles, though the problems with the salinometer make that conclusion less reliable than usual. Primary channels were selected for processing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some small unstable features that are likely real. No further editing was done.

##### Recalibration

Pressure tests during 2018-030 and 2018-032 showed no recalibration was needed. There is no evidence from this cruise to suggest otherwise.

Given the problems with the salinity analysis for this cruise and given the average differences were close to those from cruise 2018-030, the results from 2018-030 will be used to recalibrate salinity by adding 0.002psu.

File 2019-001-SBE911-recal.ccf was prepared to recalibrate by adding 0.002psu to channel Salinity:T0:C0 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0299 + 0.0275

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The primary salinity was found to be high by an average of 0.0002psu. See file 2019-001-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded based on residuals, removing the same amount of data the CTD DO was low by an average of 0.0001mL/L wit a standard deviation of about 0.013mL/L. This result is excellent and show the recalibration worked well.

See file 2019-001-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When some outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.010mL/L (standard deviation of 0.038mL/L). The largest outliers come from the top 300db and many of those samples came from casts close to shore or in inlets.

Looking at the differences versus pressure shows that the sensor DO tends to be on the high side above 1500m slightly low below the oxygen minimum zone, which is likely due to incomplete flushing of Niskin bottles. All the outliers that were removed came from above 300m which includes many samples from casts close to shore. A better estimate at those levels is ±0.5mL/L down to 150db and ±0.3mL/L to 300db. From 300-800 errors are roughly ±0.07mL/L and from 800 to 4500db roughly ±0.04mL/L.

No further recalibration is justified. See 2019-001-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

For casts #37, 41-44 and 49-51 PAR was also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------------

Conductivity, Transmissivity, Fluorescence and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for are considered,

very roughly, to be:

±0.50 mL/L from 0 to 150db

±0.30 mL/L from 150db to 300db

±0.07 mL/L from 300db to 800db

±0.04 mL/L below 800db

For details on the CTD data processing see document: 2019-001_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S plots but they are likely real. No other problems were noted.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from ~83% to 101%. The lowest values came from Saanich Inlet and 2 casts at sites where such values are typical due to active mixing. In Hecate Strait values were mostly between 94% and 97%. The offshore casts had values between 97% and 101%, suggesting that the DO sensor calibration is reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

For casts #37, 41-44 and 49-51 PAR was also removed.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-001-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and it was found that PAR should have been removed from cast #54, so the steps from REMOVE onwards were repeated for that cast.

A cross-reference listing and header check were produced for the CHE files.

It was discovered later that casts #41 and 60 had comments about bottles that had not been fired; those comments were removed from the header as they do not relate to any flag in the file. A general comment was added to the header.

##### Thermosalinograph Data

A TSG45 was used for this cruise. It does not support having a fluorometer attached, so the TSG output has to be combined with that from a fluorometer. There were 4 files produced.

The first file is short with flow not on until near the end, so there are only about 15minutes of data.

There were loop samples, a flow meter and intake thermistor. The intake is at about 4.5m.

The files have extensions RAW but are in csv format, so the files were opened in EXCEL and saved as CSV files. In opening use DELIMITED, deselect TAB, select COMMA and OTHER - *.

It was necessary to choose TEXT for the time on the 2nd page of the text import wizard.

The columns in the RAW files are:

Date UTC

Time UTC

Identifier – not needed in processing

Remote temperature (SBE38)

Lab (TSG) temperature (SBE45)

Conductivity

Flow in ℓ/minute

Fluorescence in V

Latitude  - decimal format

Longitude – decimal format

Checksum, - not needed in processing

The spreadsheet was adjusted as follows:

- 2 lines of headers were added – channel names and units.

- The file was then saved in CSV format – the name is not critical at this point.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column M. Then a concentration value was calculated in column F using offset -0.04 and scale 13.1 as determined in the most recent recalibration of the fluorometer.

- Copy/Special Paste was used to save the values and then the voltage channel was removed.

- A file break column was filled with the cruise #-data/time info from the original file name. But some of the files were very large, so the file break was changed as the date changed, with files names having format YYYY-###-YYYYMMDD-HHMMSS.

- Time is a problem – when converting from RAW choose TEXT but once open in EXCEL set Time Format  to HH:MM:SS and save the file again. (Any time the file is opened in EXCEL the time format may have to be set again before saving.)

The files were then converted to IOS format with header info added. There are 19 files, 1 or 2 for each day. The gaps were very short between files when there were 2 on a single day, so JOIN was used to combine them, so there is 1 file per day.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. The time gets set to Start Time whether or not you choose that on the final form. A record number was also added to enable averaging (for use in comparison to CTD files).

REORDER was run to move the Julian date to after the Time/Date channels. Also the record # was moved to the end.

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

a.) Plots

A track plot was produced and added to the end of this document.

Time-series plots were examined on screen. Note that TIME ZERO had to be changed to the beginning of the day in order to get the plot routine to work.

The flow was off for part of the first and last files. No other problems were noted.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2019-001-ctd-tsg-loop-rosette-comp.xlsx. There were 87 CTD casts that overlapped with the TSG record but for the first case the TSG loop flow was zero, so that was removed from the comparison.

Bin Average was used to average TSG files over 12 records (1 minute) and standard deviations were calculated. REMOVE was used to remove unwanted standard deviation channels.

Head Edit was used to change the format of the standard deviations to a format that can be exported to a spreadsheet.

Each file was opened in EXCEL and reduced to the just records required for comparisons with loops and CTD casts.

The data needed for comparison with the CTD data were then exported to a spreadsheet (Time, positions, intake and lab temperatures, salinity, fluorescence and flow rates plus the standard deviations for all except time and positions.)

Date were selected for the times of CTD casts and then added to file 2019-001-ctd-tsg-loop-rosette-comp.xlsx. The line for cast #1 was removed from the comparison because there was no flow in the TSG at the time of that cast.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no position differences >0.0003º and the median differences were both 0.00000º. So the time and positions in the TSG files are reliable and the matches are good between the two systems. Small errors are caused by averaging and imperfect time matches.

c.) Comparisons

- Comparison of T and S from TSG and CTD data and Rosette samples

|  | Latitude Difference | Longitude Difference | TSG Intake Temp - CTD Temp | TSG Lab Temp - CTD Temp | TSG Sal-CTD Sal | TSG FL/ CTD FL |
| --- | --- | --- | --- | --- | --- | --- |
| median | 0.0000 | -0.0000 | 0.0506 | 0.5395 | -0.0087 | 2.04 |
| average | 0.0000 | -0.0000 | 0.0552 | 0.5533 | -0.0121 | 2.44 |
| stdev | 0.0000 | 0.0001 | 0.0730 | 0.1840 | 0.1018 | 1.63 |
| max | 0.0001 | 0.0002 | 0.3574 | 1.9275 | 0.3556 | 11.37 |
| min | -0.0002 | -0.0003 | -0.3444 | 0.1288 | -0.7329 | 0.53 |

The intake temperature is higher than the CTD temperature by a median of 0.051C°, a larger difference than we normally expect. Using the 20 casts with the lowest standard deviations in the lab temperature or using only casts from P8 to P26 produced very similar results. During 2018-040 in September 2018 the difference was 0.003C°. This might be explained by the TSG drawing water from higher in the water column, but examination of a few casts shows that the top 5m were quite well mixed and errors due to depth mismatch are likely <0.01C°. There are some outliers in the near-shore casts where near-surface gradients could be high, but for most casts the differences vary little. It is unlikely that the CTD temperature has an error as large as 0.05C° since the 2 channels are within 0.001C° of each other.

The heating in the loop had a median value of 0.49C°. Colder intake waters will lead to more heating in the loop. The fit shown below leads to an estimate of about 0.030C° if we use intake temperatures typical of 2018-040, and that is close to the observed heating of 0.28C° for that cruise. So the fit is believable. If the intake temperature is reading too high, then the heating in the loop would be greater, ~0.54C°. These estimates are very rough and take no account of flow rate variations.

The TSG salinity values are quite close to the CTD, being low by a median of 0.009psu and by 0.004psu from P8 to P26. The derivation of the TSG salinity comes from the lab conductivity and lab temperature. For 2018-040 the TSG salinity was low by a  median of 0.03psu; that difference grew with time through the cruise. The small salinity difference might imply that the lab temperature measurement is more reliable than the intake temperature.

During 2018-040 the TSG Fluorescence values were obviously rising rapidly compared to the CTD fluorescence and loop samples. During this cruise the opposite pattern is seen.

The TSG fluorescence seems unbelievably high early in the cruise and gradually  falls to values close to those of the CTD fluorometer. During event #7 the chlorophyll near the surface was ~0.9ug/L, the CTD value was ~0.5ug/L and the TSG reading was ~2.5ug/L. The highest extracted chlorophyll value found for this cruise was 1.56ug/L near the end of the cruise during event #127. The CTD fluorescence at that point was ~1ug/L while the TSG fluorescence was 0.8ug/L.  So by the end the fluorescence is more believable.

- Comparisons of loop samples with rosette samples

There were 19 loop Salinity and CHL samples, and 6 DO and nutrient samples.

- Only 1 set of the nutrient samples occurred during a rosette cast, event #23. The values are extremely close to the rosette samples with differences of 0.08uM, 0.04uM and 0.001uM for N, S and P.

- The 6 loop DO samples were higher than the rosette samples by between 0.01 and 0.08mL/L (0.2% to 1.1%).

- There were 5 cases of CHL samples from loop and rosette with an average difference of 0.016ug/L but when 1 outlier was removed it was 0.002ug/L.

- There were 6 cases of salinity samples from loop and rosette with an average of difference of 0.0143 psu but when one major outlier was excluded the average difference was 0.0005psu. The largest difference in the 5 casts was 0.002psu.

The loops samples compare well with the rosette samples.

- Comparisons of loop samples and TSG data

The 19 salinity and extracted CHL samples were compared with TSG data (averaged over 1 minute). The ratio of TSG fluorescence to CHL ranged from about 7 early in the cruise to 2 at the end. The median ratio was 4.7 using all casts and the same when Juan de Fuca casts were excluded. When only P5 to P26 casts were included the median ratio was 5.5. There is a fairly steady drop in the ratio through the cruise though the Juan de Fuca and near-shore Line P casts show considerable variability.

For salinity the TSG Salinity was lower than the loop salinity by a median value of 0.0062psu and the same value was found when Juan de Fuca casts were excluded. But using only P5 to P26 the median difference showed the TSG salinity to be low by 0.0047psu. While stopped the median difference was -0.0036 and underway it was -0.0068psu, though when 1 outlier was excluded the stopped difference was 0.005psu.

For details on comparisons see document 2019-001-ctd-tsg-loop-rosette-comp.xls.

d.) Calibration History

- During 2018-026 this system was used for the first time. The intake temperature sensor read higher than the CTD temperature by a median of 0.03C°. The lab temperature was higher than the intake temperature by more than expected based on previous experience with heating in the loop on the Tully. The salinity traces were quite smooth with only a few very small spikes. Salinity read lower than the CTD. The differences were not sufficiently consistent to allow recalibration. The data were archived but with a warning that there was insufficient calibration sampling to assess quality.

- The results from 2018-025 suggest that there were problems with flow in the loop.

- During 2018-034 Salinity was lower than the CTD by a median of 0.030psu. Lab temperatures were higher than CTD temperatures by a median of 0.66 C°. The intake temperatures were high by a median of 0.072 C° using all data or 0.058 C° in a “quiet section”. Fluorescence from the TSG was lower than the CTD fluorescence except in areas where CHL was low. The comparisons have a lot of scatter and recalibration was not attempted. The data were archived but with a comment about the intake temperature being further from the CTD data than expected and heating in the loop being higher than when the SBE21 was in use. Because of the limitations only 3 decimal places will be shown for temperature and salinity.

- During 2018-040 the TSG fluorescence was removed as it looked unreliable, starting with values close to the CTD but increasing until it was 60 times the CTD values. Salinity was lower than that from the CTD by a median of 0.032 and lower than loop samples by 0.034psu. The intake temperature was higher than the CTD temperature by a median of 0.0032C°.  The lab temperature was higher than the intake temperature by ~0.28 C°.

e.) Conclusions

1. The TSG clock worked well.

2. There were 86 points of comparison with CTD.

3. The flow rate was quite steady with values from 1.0 to 1.1 L/min.

4. The chlorophyll and salinity loop samples compare well with the rosette samples.

5. The intake temperature was higher than that of the CTD by a median value of 0.050C° with a standard deviation 0.07C°. That is a larger difference than seen in the 2 previous uses. Between P8 and P26 the median difference was 0.051 C° with a standard deviation of 0.015C°. Taking subsets and removing outliers did not make much difference. The vertical gradients do not look large enough for a mismatch in depth between TSG and CTD to account for this. There is also no obvious difference between the results for the 2 CTD deployment methods.

6. The median difference between the lab temperature and the intake temperature (i.e. heating in the loop) was 0.49C° with a standard deviation of 0.16C°.

7. TSG Salinity was lower than that from the CTD by a median of 0.0087psu overall and by 0.0035psu between P8 and P26. The TSG was lower than loop samples by a median of 0.0062psu overall and by 0.0047psu from P8 to P26. The TSG-Loop differences were somewhat higher when the ship was moving than when stopped but when a few outliers are removed from both sets, the change does not look significant.  So the TSG salinity is likely reading low by between 0.003 and 0.009psu. This is closer than found during other uses of this equipment. Small errors in the TSG conductivity or lab temperature could lead to such an error in salinity and there could be small bubbles having some effect.

8. Fluorescence data looks unbelievably high early in the cruise, compared to both CTD and loop samples. The values gradually drop. It is possible that whatever led to the extremely high values in the previous cruise was gradually being flushed out of the system, though the loop samples don’t suggest that the loop itself was the source of the problem. Even at the end, the differences are large enough compared to loop samples that the data cannot be trusted and should not be archived.

9. There are a few sections in which the flow was turned off in the first and last files. Removing temperature and salinity data is required from those sections. No other editing is necessary.

g.) Editing

CTDEDIT was used to pad data in channels Temperature: Lab, Temperature: Intake and Salinity:T0:C0 from parts of the files from Sept.13th and Sept. 27th with 0 flow rate. 
There were some spiky sections, mostly between P26 and Dixon Entrance that are likely due to high ship speed and/or high sea states, but editing such data is not possible.

h.) Recalibration

A test was done to see if lowering the TSG intake temperature by 0.05C° and then deriving salinity using the CTD conductivity, the resulting salinity is close to the CTD salinity and it was within 0.004psu.

CALIBRATE was used to subtract 0.05 from the TSG intake temperature and to add 0.005 to the TSG salinity. These are based on rough estimates but will bring the TSG data into better line with CTD and loop data. Given the large temperature correction only 3 decimal places should be included in the intake temperature, and that is appropriate for the lab temperature and salinity too.

DERIVED QUANTITIES was run again, this time to derive sigma-T using the corrected intake temperature and salinity.

i.) Preparing Final Files

REMOVE was used to remove channel Pressure, Fluorescence:URU:Wetlabs, Temperature:Difference , and record #.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

Sampling System

---------------

The seawater is drawn from a 20-cm x 56-cm standpipe, flush with the hull

and located 4.5 metres below the water line, and pumped approximately 50

metres through a 5-cm stainless steel insulated pipe up to the main lab.

Between the intake and the pump is a 24 x 24-cm stainless steel strainer

(1/2-cm grid). The pump is a Moyno model 2L6 stainless steel progressive

cavity pump with flow and pressure control.

TSG Data Processing

-------------------

Comparisons were made between the TSG data and co-incident CTD casts, rosette

samples and loop samples. These indicate that the intake temperature is high by

roughly ~0.05C° and salinity is low by roughly 0.005psu. The small error in

salinity implies that Temperature:Lab and Conductivity are reasonably accurate.

Recalibration was applied to subtract 0.05C° from Temperature:Intake and to

add 0.005psu to Salinity.

Temperature and salinity are shown with 3 decimal places due to reflect

uncertainty in the calibration.

The only editing applied was to remove some data from a few small sections where

there was no flow in the loop. The flow rate was quite steady and the temperature

and salinity traces were free of spikes.

Fluorescence data from the TSG were removed from the files because the values were

unbelievably high early in the cruise. They gradually dropped through the cruise

even though there was little variation in extracted chlorophyll and they remained

higher than those from the CTD except for a few casts at the end of the cruise.

See report 2019-001_Processing_Report.doc for more details.

The files were saved as 2019-001-DDMM-HHMMSS.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

##### Loop File

The Chief Scientist provided file 2019-001 Loop log.xlsx which included event numbers, sample numbers and what was sampled.

Earlier in the processing the loop samples were used to study the TSG calibration. To do that times were added based on the log entries. For loops taken at the same time as 5m rosettes, the times were set to the ends of casts.

TSG data were extracted for the times of the loops and this provided latitude and longitude data and sample data from analysts were added.

The spreadsheet was saved as 2019-001-Loop-Data.xlsx.

The sampling method column was entered as USW.

The columns were arranged in the order required for the 6-line header used to prepare the loop file.

The dissolved oxygen samples from the loop were not taken as part of the core sampling so will not be included in the loop file. They all came from rosette casts and values were found to be consistently higher than the rosette DO values.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2019-001-che-surface.csv. The Oxygen:Dissolved and Oxygen:Dissolved:SBE channel in mass units were included and Draw Temperature.

The Start Time was copied into a second column and the first was formatted for date and the second for time.

The times in the files are start times and the samples were actually taken near the end of the cast, so Add Time Channel was used on the full profile CLN files to calculate an END TIME. The event numbers and end times were then exported to a spreadsheet for addition to the spreadsheet. Some manipulation was required to make the data fit because most casts had several bottles fired at about 5m. Once the times were lined up properly the Start Date/Time columns were removed.

An initial line was added with channel names from a 6-line header used for a previous cruise. The data and the 6-line header orders were adjusted due to slightly different channels being available. Then the full 6-line header was inserted.

A sample method column was added. ROS was entered for the method.

Data from below 7m were removed as well as records with no analysis results.  In some cases there were no analysis results available but some samples had been taken, so the lines were left in place to accommodate additions later. (11-DOC, 38-DOC, 61-DOC) For event 70 there were no sample data but it was the only bottle fired at the surface for that cast. So it was left in place.

The data were sorted on event number, then pressure.

That file was saved as 2019-001-surface-6linehdr.csv.

It was discovered that there were date errors in the daily science log for events 39, 40 and 41.

The loop data were added to the 6-line header file.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on event #, date, time and pressure.

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

Test plots showed there were position errors in the loop samples from events 40 and 42; those were fixed and the conversion rerun. The results look much better.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2019-001-surface.loop. The track plots look reasonable and plots of temperature and salinity versus event numbers, latitude and longitude look reasonable.

General Comments

PAR off:37, 41, 42, 44, 49, 51, 54

Casts with bottle fired out of order: 37, 38, 41, 55, 60

Casts with no Niskin closed: 69, 72, 73, 113, 115, 117, 118, 119, 122, 123, 126,129

Sample numbers not used:

Particulars

1. All Niskins closed but no bottle file needed

6. PAR configuration error corrected.

7. Some Niskins not needed – only keep ones with samples.

10. Too many spigots open so no bottle file needed.

13. Niskin 14 not needed; Transmissometer replaced with #1396

25. Niskin 21 did not close – intended samples flagged 9.

29. Niskin 21 not needed

30. Niskins 20-23 not needed.

31. Loop should be #5030

33. Niskin 21 not needed.

37. Niskin 21 did not fire – no sampling planned so remove line from file.

41. Niskins 1-4 not used. Niskin 21 not used. Nut sample shown on rosette sheet for 190 but no sample found so assume not taken.

44. Niskin 21 not needed

49. Niskin 21 not needed

55. CHL sample shown on rosette sheet but no such sample indicated in CHL spreadsheet. Presume not taken.

59. Deep water gathered for storage experiments.

60. Niskins 1-4 not needed

61. Niskin 21 not needed. Niskin 24 fired at same depth as Niskin 23, but not listed on rosette log and no sampling. Removed from bottle file.

80. Held at surface after soak for >3 minutes due to ship being moved

97. Bottle 21 and 11 swapped because bottle 11 was consistently been leaking.

111. Two sampling log sheets because 1st one was incorrect. Niskins 1 and 2 closed by mistake on way down. Niskin #12 closed at surface but no water collected, no sample #.

118. Files misnamed – Fixed.

119. Stop at 105m on way down due to wire angle and stop at 83m on way up.

120. Raw files misnamed – Fixed

120. Niskin 1 closed at bottom by mistake – not needed. Sample 727 is from Niskin 2, 728 from #3, 728 from #4.

125. File should be named 124. Fixed

143. Station name Haro59 should be just 59. Fixed

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1201DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1396DR |  | 2Feb2016 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3685 |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 2Aug17 | Factory |  |  |
| Conductivity | 0620 | 2Aug17 | Factory |  |  |
| Temperature SBE38 | ? | ? | Factory |  |  |
| WETStar Fluorometer | 953 | 5May2017 | Factory |  |  |