## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2019-003

Agency: OSD

Location: Strait of Georgia / Vancouver Harbour

Project: OSD Winter Cruise

Party Chief: Page S.

Platform: Vector

Date: 21 January 2019 – 28 January 2019

Processed by: Germaine Gatien

Date of Processing: 6 June 2019 –  17 June 2019

Number of original HEX files:  	117		Number of CTD files:  115

Number of original ROS files: 	41	   	Number of CHE files: 38

Number of TSG files: 2	(None to be archived – HDR files prepared for Chief Scientist’s use.)

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#953DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (3685), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (692) and an altimeter (43281). A Biosperical/Licor Surface PAR (#20518) was mounted on the boat deck.

SeaBird SBE21 Thermosalinograph #3411in use during this cruise.

Seasave version 7.23.2 was used for CTD acquisition and 7.26.7.110 was used for TSG acquisition.

The data logging computer was water Properties #32.

The deck unit was a Seabird model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model,  8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) Kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science and Sampling Logs were generally in good order but there was no sampling log for one cast, 1 wrong event#, the Science log was missing some information and the serial # for the transmissometer was wrong in both the con file and the log book. Fortunately, there were photos of the equipment. Log entries should be based on visual checks, not the configuration file.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. Users should check extracted chlorophyll values where available.

There was a problem with the salinometer used to analyze samples for this cruise. Tests were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”.  They established a linear correction to be applied to all salinity samples analyzed on salinometer #68572 between January 2019 and March 2019. Based on those corrected bottle values and the history of the sensors, the CTD salinity values are believed to be within +/-0.002psu.

The comparison of titrated DO to CTD DO provided a good fit when using the well-mixed casts from this cruise, but those data have too narrow a range of DO values to enable extrapolation. When the offset of that fit was made to match that found during 20119-001 (when the same sensor was used), the two slopes were very close. Since the 2019-001 correction was based on a much wider DO range it was chosen to correct these data.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.40 mL/L from 0 to 10db

±0.20 mL/L from 10 to 75db

±0.10 mL/L from 75 to 200db

±0.06 mL/L below 200db

The pH:SBE:Nominal data should be used with caution; no field calibration data are available. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

The thermosalinograph data look unreliable so were not prepared for archival. Files were prepared for the use of OSD scientists but care should be taken because the temperature and salinity value are not reliable. Channel Temperature:Primary was created as a proxy for intake temperature by recalibrating TSG lab temperature and salinity based on a comparison with CTD temperature and salinity.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained. While they were generally in good order, there were some problems:

- Event 12 had bottle sampling but no rosette sheet.

- Event 57 had 17 bottles fired but that is not entered in Daily Science Log

- Event 105 – Daily Science log shows no sampling but there is a rosette sheet for 105 – it should be 106.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the dissolved oxygen , temperature, conductivity and pressure sensors was obtained.

- The configuration files used at sea had no errors. It was saved as 2019-003-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2019-003-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format. They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. A few outliers were found in cast 84 so the file was opened in CTDEDIT and Salinity:T1:C1 was edited lightly.

The output file was copied to BOT files so there is a complete set of BOT files, edited or not.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The SAM file for event #1 was removed since there was no sampling.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-003-bot-hdr.txt which will be updated as needed during processing. (* indicates date of creation for the file.)

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2019-003 chl QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2019-003chl.csv. There were no event numbers, only station names, so the event #s were added to the csv file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2019-003oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-003oxy.csv. That file was converted into individual *.OXY files.

Some bottles were flagged due to problems that will affect other samples as well, so after the bottle samples are merged the relevant flags will be added to the other samples.

SALINITY

Salinity analysis was obtained in file QF2019-003 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab 29 to 37 days after collection. The files were simplified and saved as 2019-003sal.csv.

Problems with the salinometer were detected when COMPARE was run during 2018-031 processing. The differences were very large and roughly inversely proportional to salinity. A check of cruise 2018-032 showed similar problems even though the CTD sensors were different. A study of cruise 2019-001 was also examined (for which only some of the data were analyzed on the problem salinometer) and the comparison was similar as was also found in a preliminary comparison of this cruise. Further tests were run comparing 2 Autosal salinometers, which produced a correction similar to those suggested by the comparisons for the affected cruises.

It appeared that the Autosal was producing poor results in early 2019. It had recently been serviced at the factory. Tests were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”. They established a linear correction to be applied to all salinity analyzed on salinometer #68572 between January and March 2019. The Autosal was returned to the factory in mid-March 2019. The correction was found to be:

Salinity_Corrected = 1.01135 * Salinity measured -0.3976

This correction was applied to the values in 2019-003sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2019-003*.xlsx. This includes a precision study. The file was simplified and the saved as 2019-003-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Problems found:

- Event #77 & #78 – The final bottle entered on the rosette sheet for event 77 was fired during event 78. The analysts called it 77. So the SAMAVG files for 77 and 78 were merged to enable creation of a single file for #77.MRG.

- Events 64 and 95 were mislabelled in the dissolved oxygen spreadsheet as 63 and 93, respectively.

- A sample said to be from 105 was found to be from 106. There was no sampling from a 2nd bottle, so it was removed from the SAMAVG file. There was salinity and nutrient sampling not shown on the sampling log.

- Event #124 had 2 misfires. The analysts did not include those in their files because the sample numbers had been reassigned to other bottles, so the lines were removed from the SAMAVG file.

At this point the flags that the oxygen analyst indicated should apply to all samples were changed in the other sample files and then the MERGE process was rerun.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When bottles are excluded that have standard deviation in the CTD salinity >0.0008psu, and any bottles excluded from one channel also excluded from the other, the average shows the primary salinity to be low by 0.0017psu and the secondary low by 0.0007psu. The standard deviations in the comparisons is quite high at 0.0045psu in each channel.

The difference between the comparison for primary and secondary sensors is consistent with the bottle differences noted in section 9. These values are close to those found during cruise 2019-001 which occurred shortly after this cruise and had more extensive calibration sampling. There is some expectation of poor flushing for many of these casts because the descent rates were generally very steady. However, the vertical gradients were often very low, particularly at the depth from which most samples came, so flushing errors are likely to be fairly small.

The good correspondence to 2019-001 results confirm that the bottle correction was appropriate.

The only significant outliers were found in areas of high vertical salinity gradients; either the slight vertical offset between bottle and CTD and/or Niskin bottles not flushing completely can explain the differences.

For full details for the COMPARE run see file 2019-003-sal-comp1.xls and salinometer-calibration-study.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The initial fit using all casts but excluding outliers based on residuals was:

CTD DO Corrected = CTD DO * 1.0223 + 0.0572 (R2 = 0.63)

To see how that compares with 2019-001, the offset was forced to match that cruise and the fit was:

CTD DO Corrected = CTD DO * 1.0281 + 0.0275 (R2 = 0.59)

The slope for 2019-001 was 1.0299. That cruise was mostly offshore where errors due to poor flushing are generally found to be much smaller. We would expect an inshore cruise with fairly steady CTD descent rates to have a smaller slope than in the offshore, but in this case it is only slightly smaller. This is likely due to the fact that many casts were very well mixed. That would decrease the effect of incomplete flushing of Niskin bottles, so that samples would have higher DO relative to CTD DO than would be seen in the presence of higher gradients.

When only casts with very low vertical gradients in DO were compared the fit was:

CTD DO Corrected = CTD DO * 1.0358 - 0.0133 (R2 = 0.20)

Because the range of DO values in this fit, the offset is not reliable. When the offset was forced to match the 2019-001 result the fit for the low-gradient casts was:

CTD DO Corrected = CTD DO * 1.0280 + 0.0275 (R2 = 0.19)

This is almost identical to the fit for the whole cruise when the offset was set to the 2019-001 setting.

For the well-mixed casts the choice of offset makes little difference, but for the full cast, the difference is significant. The well-mixed casts are the most reliable from this cruise but come from such a small range of DO values that extrapolation is unwise. The 2019-001 correction looks like a better choice.

CTD DO Corrected = CTD DO * 1.0299 + 0.0275

The significant outliers were examined. One had very noisy CTD DO data. The other looks suspicious:

Cast #118, Sample 377 had Bottle DO=6.220 and CTD DO (after recalibration) = 5.35. A bottle <2m above this one has DO=5.229mL/L. The sample was flagged 3 by the analyst.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The CTD fluorescence is a little higher than extracted CHL values, as is typical when chlorophyll values are low.

The plot of ratio FL/CHL versus Extracted CHL has the shape typically seen from this type of sensor with the ratio dropping as CHL rises.

For full details of the comparison see file 2019-003-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-003-ctd. The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. As usual the T, C and Sal channel pairs track well on downcasts and upcasts are much noisier. The Dissolved Oxygen, pH, transmissivity, fluorescence, pH, PAR and SPAR profiles look normal. The PAR and SPAR values are reasonably close at the surface. The altimetry was fine near the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. Note that the wrong configuration file was selected, 2019-006-0001.xmlcon. This does not affect the results because that file has the same parameters for dissolved oxygen as the correct file, and the con file is only needed to derive DO.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. None of the casts were deep enough to determine if there is any pressure dependence. The data in shaded cells are from the 2 previous cruises that used these sensors since they were last calibrated.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-030-0050 | 300 | -0.0008 | +0.00003 | +0.0009 | High, Steady |
| 2018-030-0056 | 300 | -0.0010 | -0.00002 | +0.0007 | F.High, V.Noisy |
| 2018-030-0086 | 300 | -0.0013 | -0.00003 | +0.0008 | F.High, Noisy |
| 2018-032-0053 | 370 | +0.0005 | +0.0001 | +0.0007 | High, Steady |
| 2018-032-0063 | 400 | +0.0004 | +0.0001 | +0.0009 | High, FSteady |
| 2018-032-0079 | 340 | +0.0002 | +0.0001 | +0.0010 | Mod, Noisy |
| 2019-003-0044 | 350 | +0.0004 | +0.00015 | +0.0005 | F.High, Steady |
| 2019-003-0057 | 450 | +0.0001 | +0.00015 | +0.0012 | F.High, Steady |
| 2019-003-0125 | 350 | +0.0008 | +0.00015 | +0.0009 | F.High, Steady |

The salinity differences are consistent with the differences found in the bottle comparison and similar to the previous uses.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. 3 casts were missing station names – those were added to the CLN files. A 0 value in transmissivity was found to be a surface spike only.

The cross-reference list was printed and checked against the log book. The station name was wrong for one file and was corrected in the CLN file.

At this point casts 78 and 106 were removed from the list since they were only created to capture data from a surface bottle missed in the full rosette cast that preceded each of them. No downcast file is needed.

The cruise tracks (event #s and stn names) were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±4m the readings were checked. There were 21 such casts.

- First, the bottom depths from the headers were compared with the log book; 15 casts had significant differences and using the log entry led to a check value <5m. This may be due to the header entries being made before the ship was in position or shoaling during the cast; sounder problems might also explain this. After cast #102 there was only 1 significant discrepancy between log and header.

- For the rest the altimetry was plotted and shows the header value was good.

- For the a few changing the header depth makes too small a difference to justify changing it and in 1 case it made the check value worse.

- For cast #44 the depth in the log is clearly wrong as it is less than the maximum depth sampled. While some shoaling is possible during the cast, this seems large and using a depth from when the CTD is at the bottom is reasonable, so an estimate was based on maximum depth sampled + altimetry header.

The same changes were made to 6 affected rosette files and the altimetry header was removed from rosette files for events #106 and 121 because there was only surface sampling so a reading from the bottom is misleading.

A surface check was run and shows an average surface pressure for the cruise was 2.7db which is reasonable. There were no negative values in the full casts. The lowest value was 0.03db but was just a single point. There are many values ~0.2db that have in-water values. No problems have been noted on other cruises since the sensor was last recalibrated. So pressure does not appear to be too high; it could be a little low but there is no evidence to judge that.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

pH

Tests were run on 3 casts to determine the best setting to align pH with temperature. A setting of +40 records looked best.

Conductivity

Tests were run on 4 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. Most casts were well-mixed so that alignment was not as critical as usual. There were minor differences among the choices tests but generally the best choice was -0.5 records for both channels.

SHIFT was run twice on all SBE911 casts using -0.6 records for the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The pressure sensor was used for 2 previous cruises since it was last calibrated at the factory. It was also used on a cruise in February 2019 that has been partly processed. No problems were found. The temperature and salinity sensors were used on the same cruises as the pressure sensor. For 2 of those cruises the primary salinity was found to be low by ~0.002psu and the secondary low by ~0.001. The comparison from 1 of the previous cruises was not considered reliable.

The oxygen sensor has been used once since it was last calibrated at the factory and during one cruise after this one that has been partly processed. The correction applied to the dissolved oxygen channels were 1.0324/0.0096 and 1.0299/0.0275.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions in temperature were near the bottom at a few casts that have very odd climatology ranges, presumably because very few data have been collected at this time of year. Most of the casts fell entirely within the climatology. The only excursions in salinity were some low values near the bottom at 4 sites.

Repeat Casts – There were no repeat casts and nearby casts are too shallow to provide a reasonable comparison.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary salinity appears to be closer to bottle values, but there are fewer spikes in the primary T-S plots. The primary channels were chosen for editing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. Most of the data removed are from near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise.

Some files required no editing: 20, 23, 25, 26, 34, 48, 61, 69, 71, 141, 142.

After editing T-S plots were examined for all casts. There are some small unstable features which could be real as this is an area of active mixing. No further editing was applied.

##### Recalibration

There is no evidence that pressure requires recalibration.

The primary salinity was found to be low by 0.002 for cruises 2018-030 and 2019-001 (after this cruise and including offshore comparisons). The bottle comparison for this cruise indicated it was low by an average of 0.0018psu; there was a lot of scatter in the comparison but it is close to the results of the cruises with the most reliable comparisons. The results for 2018-032 showed it low by 0.007 but the results were not trusted. Adding +0.002psu is appropriate.

For dissolved oxygen we expect a good fit using the well-mixed casts from this cruise, but those data have such a narrow range of DO values that it is unwise to extrapolate. The most reliable correction to dissolved oxygen data appears to be the fit from cruise 2019-001:

CTD DO Corrected = CTD DO * 1.0299 + 0.0275

File 2019-003-recal1.ccf was prepared to add 0.002psu to channel Salinity:T0:C0 and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0299 + 0.0275

This correction was first applied to the SAM and MRGCORBOT files.

COMPARE was rerun for salinity using the same points as in the fit used for the original COMPARE run to ensure that corrections were appropriate and they were. The primary salinity was found to be high by an average of 0.0002psu. See file 2019-003-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen. Using the same data as for the fit the average difference shows the CTD to be higher than the bottles by 0.005mL/L.  The standard deviation is 0.026mL/L. The fit looks reasonably good over the whole range with CTD salinity with values slightly low at the low end of the range and high at the high end. See file 2019-003-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When a few outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.014mL/L (standard deviation of 0.04L/L). The differences are highest at the upper end of the range where vertical gradients are high so the slow response rate of the DO sensor leads to CTD DO values that are a little too high. Incomplete flushing of Niskin bottles also lead to the CTD DO looking low. The differences are small given the limitations in the comparison method.

A plot of differences versus pressure was used to make an estimate of errors at different depths.

See 2019-003-dox-comp3.xlsx for details.

No further recalibration is necessary.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

There was a problem with the salinometer used to analyze samples for this cruise. Tests

were run and are described in document “Salinometer-68572 study_20Mar2019.xlsx”.

They established a linear correction to be applied to all salinity analyzed on salinometer

#68572 between January and March 2019. Based on those corrected bottle values and the   history

of the sensors, the CTD salinity values are believed to be within +/-0.002psu.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.40 mL/L from 0 to 10db

±0.20 mL/L from 10 to 75db

±0.10 mL/L from 75 to 200db

±0.06 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data are available.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see document: 2019-003_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S space where waters are well-mixed, but those are very small and may be real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~69% to 98%. Those values are unusually low. Many casts were very well mixed in the top 25m and we usually see low saturation for such casts as DO is mixed vertically. The highest values were from casts with a somewhat larger vertical gradient in the top 10m. These values are similar to those found during 2018-05 in February with sampling in the same regions.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-003-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files. No problems were found.

##### Thermosalinograph Data

There was 2 thermosalinograph files from this cruise.

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

The configuration file used at sea was checked and no errors were found.

b.) Conversion of Files

The cnv file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. The traces look fine with no obvious single-point spiking in salinity.

The track plot looks fine and added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 2.5db. These were exported to a spreadsheet which was saved as 2019-003-ctd-tsg-comp.xls. There were 74 casts with data that overlapped with TSG files and had data from 2-3m.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to 2019-003-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were some larger differences than usual, up to 0.0013º in latitude and 0.0022º in longitude. When 1 outlier was excluded the maximum difference in latitude was 0.0008º and when 3 outliers were excluded the maximum difference in longitude was 0.001º. The median and average differences were both 0.0000º. So the TSG clock worked well. The larger differences are presumed to be due to averaging positions over 2 minutes and slight mismatches in time.

d.) Comparison of T and S from TSG and CTD data

|  | lat diff | long diff | Ttsg-Tctd | Stsg-Sctd |
| --- | --- | --- | --- | --- |
| average | 0.00001 | 0.00005 | 0.7349 | 0.1858 |
| median | -0.00001 | 0.00000 | 0.7793 | 0.5008 |
| stdev | 0.00026 | 0.00042 | 0.1863 | 1.1876 |

The comparison showed some very unusual features:

- The TSG temperature from the lab was higher than the CTD temperature by a median of 0.78C°   (standard deviation 0.19C°)  which is much higher than has ever been recorded for this TSG system. Previous experience has found differences of between 0.12C°  and 0.23C° on the Vector. We expect more heating in winter due to cooler intake temperatures and we have no experience with this instrument in winter on the Vector. However, it was used in February 2015 on the Tully. For those parts of the Tully cruise when temperatures were in the same range as for 2019-003, the lab temperature was higher than the CTD by about 0.3C° and the loop is longer on the Tully so we expect more heating in the loop. However, flow rates can vary in both systems, so this estimates are rough.

- The temperature difference seems oddly constant, not showing the usual dependence on intake temperature. However, the temperature range is not very large, so this effect may just not be significant enough to show up in the plots.

- If the TSG is drawing from higher in the water column than the depth of the shallowest data from the CTD, then we would expect both the TSG temperature and salinity to look lower than the CTD since near-surface temperature and salinity increase with depth for those casts that are not very well-mixed. In Indian Arm and the southern Strait of Georgia, where the salinity gradients are relatively high, the TSG salinity does sometimes look lower than the CTD salinity.

- The salinity from the TSG is higher than that from the CTD by a median of 0.5psu though the standard deviation is very high at ~1.2psu. But in the areas with well-mixed casts, the median is 0.54psu with a standard deviation of ~0.07psu. Similarly the temperature difference from the well-mixed casts has a low standard deviation of ~0.027C°.

The data show no sign of some the problems that have occurred with TSG data in the past:

- Many of the casts were very well mixed near the surface, so getting the match of loop intake depth and CTD data right should not be as critical as usual.

- The time-series plots look reasonable, showing no evidence of flow being turned off, though it could be much lower than usual without that being obvious.

- The salinity is calculated using lab conductivity and temperature. If the temperature is reading too high, we would get salinity that is too low. But the opposite is true. The high salinity suggests that the temperature sensor was reading too low.

- There have been occasions when fresh water got into the loop. This could possibly explain the temperature differences, but not the salinity.

Possible explanations:

- Both T and C sensors could have malfunctioned.

- The ship was warmer than usual perhaps due to measures taken because of low outdoor temperatures. This would increase heating in the loop but does not explain the salinity difference.

- The flow rate in the loop was extremely low so that water warmed much more than usual. That does not explain the salinity difference.

- The CTD does not capture data close to the surface. If the TSG is drawing water from very close to the surface and there is surface warming, this could explain the higher TSG temperature, but if anything that should show up as TSG salinity being lower than the CTD as it would capture low salinity surface water.

- The CTD salinity could be badly calibrated. There were problems with the salinometer, but it is believed that they have been resolved and the CTD salinity compared reasonably well with bottle samples. The same CTD was used during a Line P cruise shortly after this cast and the bottle comparison from that cruise was in line with that found for this cruise.

See 2019-003-ctd-tsg-comp.xls for details.

Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and have been used for 6 other cruises in 2016 and 2017. During cruise 2017-64 differences from casts with fairly well-mixed surface waters, as judged from CTD profiles, showed the salinity to be low by a median of about 0.5psu while the temperature was high by a median value of 0.12C°.  During 2018-031 temperature was low by 0.26C which is unbelievable and salinity was low by 0.27 which got worse when it was recalculated using a recalibrated temperature value. This TSG was removed from the Vector after this cruise.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by about 0.78Cº. We would expect it to be reading higher by about 0.25Cº to 0.3Cº.

3. The TSG Salinity is high by about 0.5psu. But if the temperature used to derive that was corrected by subtracting 0.6Cº, then salinity values would become even higher.

4. While there are a number of possible explanations for the temperature differences, none can explain the salinity problem. The last time this equipment was used the temperature in the lab was found to be low, an unheard of occurrence. The salinity was found to be low by 0.27psu and even lower if the lab temperature was recalibrated.

5. The data from this TSG have produced very inconsistent results in recent uses and the instrument has not been serviced since late 2015.

6. These data cannot be trusted and will not be archived. The variations look reasonable and the differences fairly consistent, but the values are not. If a post-cruise calibration becomes available it might be interesting to try reconverting these data, but given the results of the previous use was also bad but totally different it will be hard to estimate when changes occurred.

7. Since the differences are quite consistent with time, recalibration will be applied to salinity and to produce a proxy for intake temperature. Estimating the temperature in the lab would require a guess of how much heating occurred in the loop, and the lab temperature serves no purpose for this cruise since there was no loop sampling, so that channel will be removed. Files will be prepared for the use of OSD scientists, but the data will not be archived.

f.) Editing

No editing was required.

g.) Recalibration

Add Channels was used to create Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2019-003-tsg-recal1.ccf to subtract 0.78 from Temperature:Primary, subtract 0.48 from Temperature:Lab and subtract 0.50Cº to Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan Number, Temperature:Lab and Flag channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

Sampling System

---------------

The seawater is drawn from a 5cm standpipe which is flush with the hull

and located 2.0 metres below the water line. There is 1.6m of 5cm pipe

from the hull to the Moyno 2L6 SSQ pump and the water is then pumped

approximately 26 metres through a 3.8cm stainless steel pipe up to the

main lab. The system pressure is regulated by a 3.8cm Singer 106RPS valve.

In the lab, the seawater flows directly through the Seabird model 21

Sea-cat thermosalinograph and the exhaust goes directly overboard.

TSG Data Processing

-------------------

Sampling System

---------------

The seawater is drawn from a 5cm standpipe which is flush with the hull

and located 2.0 metres below the water line. There is 1.6m of 5cm pipe

from the hull to the Moyno 2L6 SSQ pump and the water is then pumped

approximately 26 metres through a 3.8cm stainless steel pipe up to the

main lab. The system pressure is regulated by a 3.8cm Singer 106RPS valve.

In the lab, the seawater flows directly through the Seabird model 21

Sea-cat thermosalinograph and the exhaust goes directly overboard.

TSG Data Processing

-------------------

There was no loop sampling, intake temperature sensor or flow rate meter.

A comparison of temperature and salinity was made with co-incident CTD casts.

The temperature in the lab was found to be higher than the CTD casts by a median

value of 0.78 C degrees, an unusual result as heating in the loop is expected to

be on the order of 0.3C. Temperature:Primary is a proxy intake temperature

that was created by subtracting 0.78 C degrees from the lab temperature.

The TSG salinity was higher than that from the CTD salinity with a median difference

of about 0.50psu compared to CTD data from 2-3m.

Temperature:Primary - A proxy for Intake Temperature derived by recalibrating the

lab temperature using the results of a comparison with CTD data from 2 - 3m.

Salinity:T0:C0: Salinity after recalibration based on a comparison with CTD salinity.

See document 2019-003-Processing_Report.doc for more details.

-----------------------------------------------------------------------------------

WARNING: THESE DATA ARE NOT CONSIDERED SUITABLE FOR ARCHIVAL

THESE DATA SHOULD BE USED WITH CAUTION AS THE CALIBRATION OF THE

TEMPERATURE AND CONDUCTIVITY SENSORS IS NOT TRUSTED.

The files were saved as 2019-003-000*.HDR as they are not intended for archival.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars

1. Test firing of rosette – no CHE file needed.

12. Intersal Hakai

16. Sample 90 used because rosette sheets for next cast already printed using sample 40.

31. Spike in fluorescence @280m.

51. Bottles 1-6 not used due to error in rosette sheet.

61. Bottle fired  but no sampling. No CHE file needed.

78. Sample 213 collected on event #78 but on rosette sheet for 77. No downcast file required.

105/106. Rosette sheet indicates sample from #105 but no bottles fired; Science log shows it was 106. No downcast file needed for event 106 – only created to get surface sample missed on previous event.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  |  |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  |  |  | 0506 |  | Yes |  | Yes |  |  |  |
|  | Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Sensor |  |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  |  |  | 2023 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  |  |  | 1763 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  |  |  | 5013 |  | 2Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  |  |  | 983DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| PAR |  |  |  |  | 4565 |  | 16Mar2011 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  |  |  | 3685 |  |  |  |  |  |  |  |  |
| pH sensor |  |  |  |  | 692 |  | 2Mar2018 |  | Factory |  |  |  |  |
| Surface PAR |  |  |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |
| Pressure Sensor |  |  |  |  | 0506 |  | 13Mar2018 |  | Factory |  |  |  |  |
| Altimeter |  |  |  |  | 43281 |  | ? |  | Factory |  |  |  |  |