## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 18 March 2025 | ### Updated channel names & formats in TOB & Loop files. GG&SH |
| ### 5 Sept. 2024 | ### Added DOC/TOC. SH |
| ### September 2023 | ### Fixed HPLC analysis comments. SH |
| ### 26 March 2023 | ### Added HPLC data. J.R. |
| 20 January 2021 | Transmissivity:Green recalibrated to correct an error in the original conversion. S.H. |

## PROCESSING NOTES

Cruise: 2019-006

Agency: OSD

Location: North-East Pacific

Project: Line P

Chief Scientist: Robert M.

Platform: John P. Tully

Date: 2 June 2019 –  18 June 2019

Processed by: Germaine Gatien

Date of Processing: 10 October 2019 – 18 November 2019

Number of original HEX files: 	46 		Number of CTD files: 46

Number of rosette files:	 46			Number of bottle casts processed: 46

Number of original TSG files: 	2		Number of processed TSG files: 2

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (#1883DG & 1185DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (#3640), a Biospherical QSP-400 PAR sensor (#70613) and an altimeter.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

A second thermosalinograph (Seabrid 45 from MBARI Carbon Lab) was plumbed directly from a tap from the loop.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully Laptop - Silver.

The deck unit was a Seabird model 11+ #425.

A Guildline model 8400B Autosal serial # 68572 was used to analyze salinity samples.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in excellent order with comments about problems encountered. Sampling notes were provided by the Chief Scientist which are a great help in processing data.

The deployment scheme for the CTD was:

The rosette was brought to the surface.  Pumps were turned ON.  The rosette was brought down to 10m and kept there for 30 seconds.  Once back at the surface, the data started to be archived, with the rosette at the surface for 30 seconds longer.  Then the cast would start.

There were 2 transmissometers mounted during this cruise. The configuration file used at sea had the parameters reversed from the positions in which they were actually mounted. This was corrected in processing.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The primary temperature and salinity channels were chosen for most casts, but for events #17 to 23 and #27 there was a problem with the primary conductivity so the secondary channels were prepared for archival.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.3 mL/L from 0 to 10db

±0.1 mL/L from 10db to 300db

±0.08 mL/L from 300db to 500db

±0.05 mL/L below 500db

The IOS thermosalinograph stopped working on June 8th and this was not discovered until the end of the cruise. There was a second Thermosalinograph in use that operated through most of the cruise and data from that instrument were provided by the MBARI Carbon Lab. Files were for both TSGs.

Comparisons were made between the IOS TSG data and co-incident CTD casts, rosette samples, loop samples and the 2nd TSG system. The comparisons indicate that the intake temperature was high by about 0.015C°. Salinity was found to be low by about 0.057psu. Recalibration was applied to the intake temperature and salinity channels. There was no intake temperature for the 2nd system, so a proxy was created by subtracting an estimate for heating in the loop based on the IOS system.

Temperature and salinity for both TSGs are shown with 3 decimal places to reflect uncertainty in the calibrations.

The TSG fluorescence channel was removed because there was significant drift through the cruise with values much higher than loop samples and CTD fluorescence.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity, NH4, DMS and DMSP data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Since their last calibration most of these sensors have been used on 7 previous casts (just 6 for the DO sensor).

- There were 2 transmissometers mounted on the CTD and it was reported after the cruise that they were mounted in reverse order to that indicated in the configuration file.

- The calibration control files were checked. The only problem was the reversal of the 2 transmissometers as also occurred during 2019-005; they were mounted in the opposite order to that in the original con file. After that correction the file was saved as 2019-006-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2019-006-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and there were some spikes in both salinity channels, but in only one cast did those look like noise rather than real conditions. CTDEDIT was used to clean outliers in both salinity channels in cast #78 and the output file was copied to *.BOT.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-006-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF 2019-006 OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-006oxy.csv. That file was converted into individual *.OXY files.

At this point note was made of all comments in the DO file that had comments starting with “ALL:”. The rosette sheets were checked to note what other samples were taken from the same bottle so the flags will be applied to all samples from the other csv files that are affected by the DO analysts observation. Only the nutrient samples were affected.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2019-006 CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2019-006chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2019-006SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 52 to 63 days after collection. The files were simplified and saved as 2019-006sal.csv. That file was then converted to individual SAL files.

51 of the Salinity:Bottle values required recalculation due to non-linearity in Autosal s/n 68572. See the notes in the above document for more details.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2019-006*.xlsx. This includes a precision study and a comparison of nutrients collected using both polypropylene bottles and plastic tubes for 2 casts. Polypropylene bottles were used because there were not enough plastic tubes due to extra stations being added. The file was simplified, saved as 2019-006nuts.csv and converted to individual NUT files.

The flags were changed to 3 for the samples noted by the oxygen analyst as having a problem that would affect all samples.

DMS

DMS data were obtained in spreadsheet DMS summary (2019-001).xls. Details on analysis are in file 2019-001 DMS report.doc which includes duplicate analysis. The file was sorted on sample #. This file was converted to DMS files.

DMSP

DMSP data were obtained in spreadsheet 2019-001 dmsp QF summary.xls. The comments from DMSP-D and DMSP-T were combined in a single column. This file was simplified by removing blank lines, sorted on sample # and converted to DMSP files.

The SAL, CHL, OXY, NUT, DMS and DMSP files were merged with CST files in 6 steps.

After the 6th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. The only discrepancy was that event #79 had a salinity sample not shown on the rosette sheet, though a comment mentions salinity for 2 bottles.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the fits, but if data above 500db are excluded the primary fit is reasonably flat and the secondary fit is extremely flat. One of the 2 bottles that were flagged 4 was the largest outlier, so was removed from the fit. The other is no more out of line then some other bottles.  The standard deviations in the CTD salinity are not high for any of the bottles below 500db. The primary salinity is low by an average of 0.0034psu with a standard deviation of 0.0016 while the secondary salinity is low by an average of 0.0006psu with a standard deviation of 0.0015. The difference between the salinity channels are a little larger than during 2019-005 but not as large as seen in a preliminary examination of data from 2019-069 in August.

Sampling above 150db shows evidence of incomplete flushing of Niskin bottles with the CTD reading lower than bottles.

A fit against time shows differences growing with time at the same rate for both sensors, so is likely due to different depths of sampling.

The only significant outlier was already flagged 4 and it is not so far out of line as to justify a 5 flag. Other outliers are likely cases of incomplete flushing so that the values do reflect the bottle contents.

For full details for the COMPARE run see file 2019-006-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter in the fit of differences between CTD and bottle DO versus CTD DO, many of the outliers coming from Saanich Inlet samples which has high gradients known to challenge the sensor.

The fits are very sensitive to the choices made in identifying outliers. The fits generally look ok at the bottom and top of the cast, but not great at mid-depths. A fit that looks good to the eye for deep and mid-depth DO values looks poor near the surface. There are many more samples at the high and low end of the DO range. At the low end there is not a lot of variability among the values (once Saanich Inlet is excluded), but at the high end there is a lot of variability. The vertical gradients are high there and there are subsurface DO maxima at about 50db for many casts, complicating the fits. A fit that includes most of the near-surface samples was:

CTD DO Corrected = CTD DO * 1.0376 + 0.0118

When more outliers were removed based on residuals the number of near-surface bottles was greatly reduced and the fit was:

CTD DO Corrected = CTD DO * 1.0350 + 0.0138 (R2 = 0.97)	*

This looks slightly better for the high gradient zones between 3mL/L and 6mL/L and no worse elsewhere.

Cast #45 had a fairly large subsurface peak at about 40db while #65 was fairly steady in the top 50m. The fits for those 2 casts were:

#45 	CTD DO Corrected = CTD DO * 1.0392 + 0.0145

#65	CTD DO Corrected = CTD DO * 1.0336 + 0.0028

That variability does suggest that it is wisest to limit the number of near-surface values in order to get a fit that is reliable for most of the profile. So correction (*) above was selected for recalibration of the CTD DO.

The only significant outlier that is NOT from Saanich Inlet is sample 515 from cast #77. That sample was flagged 3 by the analyst as out of line; given the poor comparison with the CTD, the analyst changed the flag to 4.

A hysteresis check was done by comparing differences above and below 1000db. There were no obvious problems with deeper values in line with shallower ones in the same DO range.

A check against time shows that CTD DO tends to be a relatively lower than bottles when the ship was well offshore (casts P25 to PA-012) but when P4 is revisited the results are similar to the casts before P25. The temperatures offshore are higher than usual.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The ratio of FL/CHL versus CHL had the shape typical of these sensors starting at about 3 to 5 for very low CHL and dropping to about 1 when CHL approaches 1ug/L. The ratio drops to <1 for CHL>10ug/L. the results from the previous cruise, 2019-005 were very similar.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-006-ctd.xmlcon. The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. The T and C pairs were reasonably close during downcasts with upcasts quite noisy, as usual. Fluorescence, PAR and Dissolved Oxygen profiles looked normal. The “Green” transmissometer was generally higher than the “Red” one; the profiles had similar shapes. The altimetry looked fine near the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many SBE DO sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. Lines highlighted in grey are from other cruises using the same equipment before and after this one.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2019-001-0020 | 1000 | -0.0009 | +0.00007 | +0.0019 | High, XNoisy |
| “ | 1900 | -0.0008 | +0.00005 | +0.0019 | “ |
| 2019-001-0037 | 1000 | -0.0009 | +0.00008 | +0.0019 | High, XNoisy |
| “ | 1900 | -0.0010 | +0.00007 | +0.0021 | “ |
| 2019-001-0049 | 1000 | -0.0005 | +0.00010 | +0.0019 | High, XNoisy |
| “ | 1900 | -0.0007 | +0.00010 | +0.0020 | " |
| 2019-005-0051 | 1000 | -0.0010 | +0.00015 | +0.0028 | High, V.Noisy |
| “ | 1900 | -0.0006 | +0.00015 | +0.0025 | “ |
| 2019-005-0087 | 1000 | -0.0010 | +0.00016 | +0.0029 | High, Noisy |
| “ | 1900 | -0.0006 | +0.00015 | +0.0024 | “ |
| 2019-005-0159 | 1000 | -0.0009 | +0.00020 | +0.0034 | F.High, Noisy |
| “ | 1900 | -0.0007 | +0.00018 | +0.0030 | “ |
| 2019-006-0032 | 1000 | -0.0007 | +0.00025 | +0.0040 | High, XNoisy |
| “ | 1900 | -0.0004 | +0.00025 | +0.0035 | “ |
| “ | 2900 | -0.0003 | +0.00026 | +0.0035 | “ |
| 2019-006-0057 | 1000 | -0.0007 | +0.00027 | +0.0039 | High, Noisy |
| “ | 1900 | -0.0004 | +0.00026 | +0.0037 | “ |
| “ | 2900 | -0.0005 | +0.00026 | +0.0037 | “ |
| 2019-006-0077 | 1000 | -0.0007 | +0.00021 | +0.0032 | High, XNoisy |
| “ | 1900 | -0.0004 | +0.00021 | +0.0030 | “ |
| “ | 2900 | -0.0006 | +0.00021 | +0.0030 | “ |
|  | 3900 | -0.0005 | +0.00022 | +0.0032 | “ |
| 2019-069-0030 | 1000 | -0.0004 | +0.00041 | +0.0054 | High, Noisy |
| “ | 1900 | -0.0003 | +0.00040 | +0.0052 | “ |
| 2019-069-0051 | 1000 | -0.0009 | +0.00041 | +0.0059 | High, F Noisy |
| “ | 1900 | -0.0004 | +0.00039 | +0.0053 | “ |

The temperature differences are a little smaller than the previous cruise and it appears that a drift to smaller differences continued after this cruise judging by preliminary data from cruise 2019-069. Conductivity and salinity differences are higher and again that trend continued after this cruise.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run. No errors were found.

The header check was run. The only problem detected were some spikes in cast #1. A plot shows those occurred at the end of the upcast so are of no concern since it did not occur during any bottle stop.

Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header.  The altimetry header was removed from casts 29, 50, 53 and 94 because the CTD did not get close to the bottom and the header reading must have come from spikes. In other cases the check value is >4m so plots were made to see if the header altimetry reading is reliable and it was; it is assumed there was some change in depth between the start of the cast and the CTD reaching bottom and/or the sounder was unreliable. For station P26 a depth of 4250m was put in the headers based on historic data and because that value looks good for the one cast with altimetry data near the bottom.

Altimetry data from the SAM files were also adjusted with the header being removed from cast #26 and 95 as the bottle file contains no data near the bottom and the depths for P26 casts were changed to 4250m.

There is contradictory information concerning the pressure accuracy.

- A surface check was run and shows an average surface pressure for the cruise was 3.1db which is reasonable for the Tully.

- During 2019-001 the average surface pressure was 2.9db which is a little low for the Tully in open waters.

- The CTD was taken out of water for a few seconds during cast #49 to get a PAR reading in air. The lowest pressure recorded is 0.9db but the data associated with that reading have in-water values, though pumping may affect that since the time in air was extremely short. The transmissivity is ~0 during that period suggesting it was in a surface slick.

- During cast #10 pressures were as low as 0.3db with transmissivity looking like the sensor was in water; the pumps were off so other channels are not very informative but the secondary conductivity is very low while the primary is not, so it may have left the water briefly.

If the pressure is actually too high by 0.9db that would indicate an average surface pressure for this cruise was really 2.2db and for 2019-001 it would be 2.0db. Those seem unusually low for the Tully in open waters. For the Vector cruise in April, 2019-038, it would be 1.4db, again an unusual result. So there is some evidence that the pressure might be reading a little too high but it is not convincing and if there is an error it is <1m. No recalibration will be applied.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts to assess whether the same settings used for 2019-005 to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space) were appropriate for this cruise and they were. The best setting for both sensors was -0.7records.

SHIFT was run twice on all SBE911 casts using -0.7 records for the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature and conductivity sensors have been used on 7 cruises since they were last serviced. Two of the cruises did not have useful calibration sampling. During 2018-030 the primary salinity was found to read low by 0.002 and the secondary by 0.001. During 2018-032 there were problems with the salinometer so the comparison was not trusted. During 2018-032 DO was recalibrated using a slope of 1.0324 and offset of 0.0096. During 2019-003 and 2019-001 the primary salinity was low by ~0.002 and the secondary was low by <0.001psu. During 2019-005 the primary was low by ~0.0023 and the secondary was low by 0.0001psu.

The DO sensor #3234 was used for 6 of those cruises noted above. The fits had slopes varying from 1.025 to 1.032 and offsets from 0.0096 to 0.0275.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperatures fell within those ranges. All salinity values were also within the climatology except high salinity between 40 and 100m during cast #102. This was a repeat of cast #12 at P4; the earlier cast was within the climatology.

Repeat Casts –Differences in temperature and salinity along lines of constant σt for 2 consecutive casts at P26 were ~0.002ºC and ~0.0002psu at about 1800db. This is excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary channels were selected for editing for most casts because there was more noise in the secondary channels and more unstable features in T-S space. However, as found during 2019-005 there were some casts that had poor primary salinity, particularly near the surface. Tests were run varying the alignment of conductivity but those that improved the near-surface results led to more unstable features at depth. The descent rate does not appear to be a factor. Irregular flow rates are a likely cause.

Secondary channels were selected for events #17 to 23 and 27.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some small unstable features that are likely real. No further editing was done.

##### Recalibration

Pressure tests during previous cruises did not indicate any problem with pressure. The evidence from this cruise is confusing, so no correction will be applied.

File 2019-006-SBE911-recal.ccf was prepared to recalibrate by adding 0.0034psu to channel Salinity:T0:C0 and 0.0006psu to channel Salinity :T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0350 + 0.0138

This correction was first applied to the SAM and MRGMRH files.

COMPARE was rerun for salinity to ensure that corrections were appropriate and they were. The primary salinity was found to be low by an average of 0.00003psu and the secondary low by 0.00001psu. See file 2019-006-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded using the same points as in the original fit, the CTD DO was low by an average of 0.00003mL/L. This shows that the recalibration was done correctly. See file 2019-006-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When deep Saanich Inlet samples and a few large outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.005mL/L (standard deviation of 0.04mL/L). This is as good a fit as we are likely to achieve.

No further recalibration was applied. See 2019-006-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #17-23 and 27 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

For casts #17-23 and 27  REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

For casts #32, 45,57,65,66,69,72,74,75 and 77 channel PAR was also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the comments about processing.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from ~103% to 135%. The only values >110% came from Saanich Inlet, and stations P1 to P4. The return visit to P4 was a little lower than on the outward trip at 108% while it was 112% on the way out. These values are a little higher offshore than we usually see, so it is possible the correction has overestimated values. However, the bottle comparison does not support that conclusion with near-surface CTD values falling both high and lower than bottle values.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts except #13-23 and 27 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

For casts #13-23 and #27 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

PAR was removed from casts 32, 42, 45, 57, 65, 66, 69, 72, 74, 75, 77.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-006-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found except that a few duplicates appear not to have been taken. Those may get flagged in the final clean-up stage.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise but stopped recording on June 8th at station P20. There were data available from an MBARI TSG45 that covered the whole cruise, but no positions were available for that data set. An RMC file from the ship data was available with positions every second. The file was reduced to a position every 15s to match the TSG file and those positions were added to the MBARI file.

The IOS SBE TSG45 data were in 1 raw file.

The files have extensions RAW but are in csv format, so the files were opened in EXCEL and saved as CSV files. In opening use DELIMITED, deselect TAB, select COMMA and OTHER entering *.

It was necessary to choose TEXT for the time on the 2nd page of the text import wizard.

The columns in the RAW files are:

Date UTC

Time UTC

Identifier – not needed in processing

Remote temperature (SBE38)

Lab (TSG) temperature (SBE45)

Conductivity

Flow in ℓ/minute

Fluorescence in V

Latitude  - decimal format

Longitude - decimal format

Checksum - not needed in processing

The spreadsheet was adjusted as follows:

- 2 lines of headers were added – channel names and units.

- The file was then saved in CSV format – the name is not critical at this point.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column M. Then a concentration value was calculated in column F using offset -0.04 and scale 13.1 as determined in the most recent recalibration of the fluorometer.

- Copy/Special Paste was used to save the values and then the voltage channel was removed.

- A file break column was filled with the cruise #-data/time info from the original file name.

- Time and Date formats are a problem – when converting from RAW choose TEXT but once open in EXCEL set Time Format  to HH:MM:SS and save the file again. (Any time the file is opened in EXCEL the time format may have to be set again before saving.) Using “custom” formats seems to work best.

The file was then converted to IOS format with header info added.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. (To do this tick “Add offset from Time Zero in Days” and enter Time Zero as the last day in the previous December at 00:00:00.)  A record number was also added to enable averaging (for use in comparison to CTD files).

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end.

a.) Plots

A track plot was produced and added to the end of this document.

Time-series plots were examined on screen. There were problems with the flow rate in the first 34 minutes but the ship was not moving at the time. Those records will be removed later.

The flow rate was very steady except for one very brief time of no flow.

No other problems were noted.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2019-006-ctd-tsg-loop-rosette-comp.xlsx. There were 28 CTD casts that overlapped with the TSG record but for event #1 the TSG loop flow rate was very high, so that comparison may not reflect results overall.

The TSG file was opened in EXCEL. The median and standard deviations over 30s were calculated for intake and lab temperature, fluorescence and salinity.

The data needed for comparison with the CTD data were then exported to a spreadsheet (Time, positions, intake and lab temperatures, salinity, fluorescence and flow rates plus the standard deviations for all except time and positions.)

TSG Data were selected for the times of CTD casts and then added to file 2019-006-ctd-tsg-loop-rosette-comp.xlsx.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no differences greater than 0.0002º while the median differences were 0.00003º and 0.00004º.

c.) Comparisons

- Comparison of T, S and fluorescence from TSG and CTD data

The comparisons from event #1 are way out of line with the others which is assumed to be due to the very high flow rate. So that event was excluded from the comparisons.

|  | Tint-Tctd | Tlab-Tctd | Stsg-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| average | 0.0599 | 0.3617 | -0.0570 | 4.87 |
| median | 0.0146 | 0.3197 | -0.0574 | 4.10 |
| Std Dev | 0.1345 | 0.1495 | 0.0089 | 4.30 |
| max | 0.5972 | 0.9197 | -0.0378 | 14.17 |
| min | -0.0262 | 0.2038 | -0.0760 | 0.31 |

The intake temperature is higher than the CTD temperature by a median of 0.015C°. However, the standard deviations were very high. When 3 more outliers were excluded the intake temperature was found to be high by a median of 0.0125C° and the standard deviation was 0.0239C°.

The heating in the loop (Temp Lab – Temp Intake) had a median value of ~0.3C° which is lower than during 2019-001  and 2019-005 when heating in the loop was ~0.50C°. The results from February might be explained by cooler intake water, but for 2019-005 in May the intake temperatures were similar to those from this cruise. The standard deviation was also very high for this calculation. When 4 outliers were excluded the Lab temperature was found to be higher than the CTD by a median 0.3127C° with a standard deviation of 0.0549C° and heating in the loop (Lab Temp – Intake Temp) was found to be ~0.28C°.

The TSG salinity was low by a median of 0.057psu which is again very different from the results of 2019-005 when they were higher by 0.4psu. We expect the TSG to read a little lower due to small bubbles. The standard deviations were low in this calculation with no obvious outliers.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.3 to 14 with a median value of 4.1. When only cases with CTD fluorescence >1 are included the ratio has a median of 0.8, so the mismatch at low values is skewing the result.  The picture looks different when displayed against time. The TSG fluorescence appears to be increasing with time, while the CTD fluorescence starts high, then drops to lower values until about event #32 with a slight increase after that. There appear to be some similarities in the patterns of increases and decreases in the 2 fluorescence sources, but with a steady increase in the TSG FL values superimposed.

- Comparisons of loop samples with rosette samples

There were  just 5 cases of salinity and extracted chlorophyll sampling from the rosette that overlapped with loop samples.

The salinity differences ranged from the loop being low by 0.0024 to high by 0.0016psu with a median showing the loop salinity to be high by 0.0006psu.

The range of the ratio Loop CHL/Rosette CHL was 0.0.81 to 1.13 with a median of 0.96. The differences ranged from -0.13 to +0.06 with the loop values generally slightly lower than the rosette samples.

While few data are available these results suggest that the loop sample data are reliable.

The 6 loop DO samples were also compared with rosette samples, though this is of no significance for the TSG. The loops were closer to the rosette than seen in the past, with a median difference of  0.0mL/L, but the standard deviation was 0.03mL/L.

- Comparisons of loop samples and TSG data

There were 15 loop Salinity and CHL samples of which 10 were taken underway.

The TSG fluorescence was higher than the loop CHL with the ratio of TSG FL/Loop CHL having a median value of ~5.9. When only cases where the TSG FL >1ug/L were included the ratio was 0.93. While this might suggest that the TSG CHL is fine for higher CHL values, it could also reflect a temporal drift since most of the lower CHL values came after event #5. The temporal drift explanation is supported by the samples from events #2 and #68 which had loop chlorophyll values of 1.35 and 1.36ug/L while the TSG fluorescence values were 1.6 and 6.6ug/L, respectively.

The TSG Salinity was lower than the loop salinity by a median value of 0.06psu with a range of -0.043 to +0.075psu; while stopped the difference was 0.053psu and while moving 0.061psu. There are not enough samples to make that difference significant, but it could indicate that the TSG draws water from a little higher in the water column while in motion.

- MBARI TSG comparison

Data were available from an MBARI TSG. This recorded date, time, lab temperature and salinity but not positions. Positions were found by matching times to one of the shipboard systems. File RMC-RAW_20190603-020652.Raw and RMC-RAW_20190616-204812.RAW were opened in EXCEL and data were reduced to every 15th record to match the spacing in the TSG file. There was a gap in the RMC files starting at June 15 at 5:22, so that had to be considered in doing the fits.

Records were inserted in the TSG file to remove a gap between 2:54 and 3:12 on June 3rd. Positions were then added to the TSG file and pad values entered for temperature and salinity. Changes were needed to formats. A 6-line header was added. The MBARI file has about 4 hours of records that were not available in the ship file used to provide positions. There were a few minutes more of data in several other files, but no attempt was made to use those.

The IOS and MBARI TSG records were thinned to about every 4 minutes. The matches were far from perfect as the IOS TSG time interval was not exactly 5s (4 to 6s) presumably due to truncation of a more precise time. By the end of the comparison the time difference was about 4 minutes.

The lab temperature from the IOS TSG reads higher than that from the MBARI system by a median of 0.0049C°. The difference between salinity readings started with the IOS salinity lower than the MBARI salinity by about -0.05psu and ended at close to -0.07psu. The median difference was -0.063. The salinity difference is close to that found between the IOS TSG salinity and CTD salinity and the differences from loop samples. This suggests that the MBARI system was reasonably well calibrated and that the IOS system salinity was reading low. The lab temperature difference also suggests that <0.005psu of the salinity difference could be due to the lab temperature reading high for the IOS TSG rather than due to drifting conductivity calibration and/or bubbles in the IOS TSG system. Keeping in mind that the MBARI system may have been in a different temperature environment, we can’t determine which is correct, just that they are different. But the loop and CTD comparison do confirm that the MBARI salinity is much closer to CTD salinity. The differences between the CTD and TSG increased with time, so at least one system may have conductivity calibration drifting and the comparison of TSG with bottles suggest that the IOS TSG  conductivity was drifting though temperature variations temperature and loop heating variations complicate interpretation.

This comparison supports archiving the MBARI file as it provides salinity for more of the cruise, though it only contains lab temperature. However, the results from the IOS system provide an estimate of heating in the loop, so a proxy for intake temperature can be provided by subtracting 0.30C° from the lab temperature.

For details on comparisons see documents 2019-006-ctd-tsg-loop-rosette-comp.xlsx and 2019-006-MBARI-IOS-TSG-comparison.xlsx.

d.) Calibration History

- During 2018-026 this system was used for the first time. The intake temperature sensor read higher than the CTD temperature by a median of 0.03C°. The lab temperature was higher than the intake temperature by more than expected (~0.5) based on previous experience with heating in the loop on the Tully. The salinity traces were quite smooth with only a few very small spikes. Salinity read lower than the CTD by a median of 0.009psu. The differences were not sufficiently consistent to allow recalibration. The data were archived but with a warning that there was insufficient calibration sampling to assess quality.

- The results from 2018-025 suggest that there were problems with flow in the loop.

- During 2018-034 Salinity was lower than the CTD by a median of 0.030psu. Lab temperatures were higher than CTD temperatures by a median of 0.66 C°. The intake temperatures were high by a median of 0.072 C° using all data or 0.058 C° in a “quiet section”. Fluorescence from the TSG was lower than the CTD fluorescence except in areas where CHL was low. The comparisons have a lot of scatter and recalibration was not attempted. The data were archived but with a comment about the intake temperature being further from the CTD data than expected and heating in the loop being higher than when the SBE21 was in use. Because of the limitations only 3 decimal places will be shown for temperature and salinity.

- During 2018-040 the TSG fluorescence was removed as it looked unreliable, starting with values close to the CTD but increasing until it was 60 times the CTD values. Salinity was lower than that from the CTD by a median of 0.032 and lower than loop samples by 0.034psu. The intake temperature was higher than the CTD temperature by a median of 0.0032C°.  The lab temperature was higher than the intake temperature by ~0.28 C°.

- During 2019-001 the TSG fluorescence values started unbelievably high, but the values dropped through the cruise getting closer to CTD and loop values. The TSG intake temperature was higher than the CTD by about 0.05C°. The lab temperature was higher than the intake temperature by ~0.5C°. Salinity was lower than CTD and loop by between 0.003 and 0.009psu.

- During 2019-005 the TSG fluorescence values were about 50% of those from the CTD and ~40% of loop extracted chlorophyll samples. The pattern of variability is similar to that from the CTD. The TSG intake temperature was higher than the CTD by about 0.19C°. The lab temperature was higher than the intake temperature by ~0.5C°. Salinity was higher than CTD and loop by 0.4 and 0.34psu, respectively. The salinity is usually lower than the CTD and loop and the two TSG temperature channels are way out of line with expected values. TSG data were not archived.

e.) Conclusions re IOS TSG

1. The TSG clock worked well.

2. There were 27 points of comparison with CTD.

3. After an initial period with no flow and then very high flow, the flow rate was steady with values 1L/min to 1.1L/min.

4. The intake temperature is reading higher than the CTD by a median difference of 0.013C°.

5. TSG Salinity is reading lower than the CTD by a median of ~0.057psu and lower than loop salinity by a median of 0.061psu. The difference is slightly larger than from the early uses of this equipment but much better than in the previous 2 cruises.

6. The lab temperature is higher than the CTD temperature by a median of 0.313C° and higher than the intake temperature by a median of ~0.28C°. So heating in the loop is much lower than during 2019-005 when heating in the loop was 0.5C° when the flow rate was also ~1.

7. The history of the instrument shows the intake temperature always reading higher than the CTD but the differences were generally small except for 2019-005 when it was high by 0.20C. Heating in the loop varies greatly from cruse to cruise. TSG salinity is usually lower than CTD salinity with the only exception being 2019-005.

8. We have much to learn about this new TSG. Results to date seem fairly consistent for intake temperature and salinity with the exception of 2019-005, but heating in the loop tends to be either ~0.3C° or ~0.5C° and not between those two values though it was once ~0.6C°.

9. The TSG fluorescence is reading lower than the CTD fluorescence early in the cruise but drifts upwards so it is reading much higher overall. This appears to be due to a temporal drift rather than a dependence on chlorophyll level.  The overall pattern does not look reliable so this channel should be removed.

10. The intake temperature will be recalibrated by subtracting 0.013C° as our best estimate. The lab temperature differs slightly between the 2 systems but that may be due to slightly different environments, so no recalibration will be applied to lab temperature.

11. For the MBARI system, a proxy for intake temperature can be prepared by subtracting 0.28C° from the lab temperature.

g.) Editing

CTDEDIT was used to pad data in channels Temperature: Lab, Temperature: Intake and Salinity:T0:C0 from the first 730 records because flow rates were 0 or ~2 followed by a period of very high variability when the flow was first changed to 1. There were some other very brief periods of 0 flow, but they do not appear to have been long enough to affect data.

h.) Recalibration

CALIBRATE was use to decrease intake temperature by 0.013C° and increase salinity by 0.057psu.

Sigma-T was not derived since the errors are likely to be large as they will combine those due to using a proxy for intake thermistor and a recalibration of TSG Salinity based on comparisons that had a lot of variability.

i.) Preparing Final Files

REMOVE was used to remove channel Pressure, Fluorescence:URU:Wetlabs, Temperature:Difference , and record #.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments.

The file was saved as 2019-006-MMDD-HHMMSS.TOB. (2019-006-0603-020652.tob)

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

The MBARI file was also prepared for archiving since the IOS TSG did not cover the whole cruise. It was put through steps:

- CLEAN to add start/stop times and positions.

- ADD Time Channel to add Julian Time.

- ADD CHANNEL to add channel Temperature:Primary which was set equal to channel Temperature:Lab.

- CALIBRATE was then run to subtract 0.28C from Temperature:Primary so it can serve as a proxy for intake temperature, but this must not be considered as reliable as an intake thermistor. Fewer significant figures should be used.

- The match between the intake proxy temperature and lab salinity is not considered adequate for derivation of sigma-T.

- REORDER to rearrange channels.

- There are many small spikes towards lower values, so salinity values are likely slightly low overall.

- HEAD EDIT was used to add header comments and to correct channel names and formats. Temperature and salinity are given with 3 decimal places.

##### Loop File

The Chief Scientist provided file 2019-0006 Loop log.xlsx which included event numbers, sample numbers and what was sampled.

Earlier in the processing the loop samples were used to study the TSG calibration. To do that times were added based on the log entries. For loops taken at the same time as 5m rosettes, the times were set to the ends of casts.

TSG data were extracted for the times of the loops and this provided latitude and longitude data and sample data from analysts were added. For times after the TSG stopped positions were obtained from the log book.

The spreadsheet was saved as 2019-001-Loop.xlsx.

The sampling method column was entered as USW.

The columns were arranged in the order required for the 6-line header used to prepare the loop file.

The dissolved oxygen samples from the loop were not taken as part of the core sampling. They all came from rosette casts and values were found to be very close to rosette values, on average, with a standard deviation of 0.03mL/L. They are only available for a few casts and those all have surface rosette samples available so they will not be included in the loop file.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2019-006-che-surface.csv. The Oxygen:Dissolved and Oxygen:Dissolved:SBE channel in mass units were included and Draw Temperature.

The Start Time was copied into a second column and the first was formatted for date and the second for time.

The times in the files are start times and the samples were actually taken near the end of the cast, so Add Time Channel was run on the full profile CLN files to calculate an END TIME. The event numbers and end times were then exported to a spreadsheet for addition to the spreadsheet. Some manipulation was required to make the data fit because most casts had several bottles fired at about 5m. Once the times were lined up properly the Start Date/Time columns were removed.

An initial line was added with channel names from a 6-line header used for a previous cruise. An extra column was added for the 2nd transmissometer. The data and the 6-line header orders were adjusted due to slightly different channels being available. Then the full 6-line header was inserted.

A sample method column was added. ROS was entered for the method.

Data from below 7m were removed as well as records with no analysis results.  In some cases there were no analysis results available but some samples had been taken, so the lines were left in place to accommodate additions later.

The data were sorted on event number, then pressure.

That file was saved as 2019-001-surface-6linehdr.csv.

The loop data were added to the 6-line header file.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on event #, date, time and pressure.

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2019-006-surface.loop. The track plots look reasonable and plots of temperature and salinity versus event numbers, latitude and longitude look reasonable.

General Comments

PAR off: 32, 42, 45, 57, 65, 66, 69, 72, 74, 75, 77.

Casts with bottle fired out of order: 12, 32, 34, 53, 63, 78.

Casts with no Niskin closed: None.

TSG stopped on June 8th at station P20

Trace Metal casts (data processed elsewhere): 31, 33, 35,47, 49, 52, 60, 62, 64, 81, 82, 83, 84, 101, 103.

Particulars

1. Cruise # wrong in header.

7. Errors in depth on hand-printed labels.

7 & 10. Changed planned order so sample numbers out of order.

11. Long stop at 500m because the winch was making weird noise.

19. Station name wrong in header – should be P5. Fixed.

65. Sample numbers out of order due to change in order from that planned.

69. Water depth wrong in header – should be 3917. Fixed.

76. Station name wrong in header – should be P26. Fixed.

77. Station name wrong in header – should be P26. Fixed.

78. Niskin #10 not used; from Niskin 11 on the Niskin #s are off by 1. See correction on rosette sheet.

84. Repetition of sample numbers 587 and 588 – used 9587 and 9588 for Trace Metal samples.

95. Station name wrong in header – should be PA-012. Fixed.

2019-006

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 21Nov2018 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 16Dec2018 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3640 |  | Jan 2019 |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 2Aug17 | Factory |  |  |
| Conductivity | 0620 | 2Aug17 | Factory |  |  |
| Temperature SBE38 | ? | ? | Factory |  |  |
| WETStar Fluorometer | 953 | 5May2017 | Factory |  |  |