## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6 | Run at 2019/10/10 12:35:58 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 0.2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Salinity_T0_C0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2019-006-0006.mrg | 2019-006-0006.sam | 206 | 5.038 | 5.0297 | 0.099706 | -0.0082865 | 31.699 | 31.6927 | 0.00094052 | -0.0062809 | 31.699 | 31.6948 | 0.0013737 | -0.0041828
2019-006-0010.mrg | 2019-006-0010.sam | 288 | 5.137 | 5.1496 | 0.10815 | 0.012629 | 31.7246 | 31.7169 | 0.00042182 | -0.0077171 | 31.7246 | 31.7188 | 0.00065426 | -0.0058098
2019-006-0011.mrg | 2019-006-0011.sam | 164 | 5.079 | 5.0349 | 0.11122 | -0.04411 | 32.4977 | 32.4938 | 0.00039708 | -0.0039482 | 32.4977 | 32.4947 | 0.00039457 | -0.0030251
2019-006-0013.mrg | 2019-006-0013.sam | 42 | 1250.41 | 1250.41 | 0.11324 | -0.0014648 | 34.4814 | 34.4777 | 0.00025257 | -0.0036583 | 34.4814 | 34.4802 | 0.00025068 | -0.001194
2019-006-0013.mrg | 2019-006-0013.sam | 166 | 1250.08 | 1250 | 0.099236 | -0.08374 | 34.4813 | 34.4778 | 0.0002617 | -0.0035019 | 34.4813 | 34.4803 | 0.00025609 | -0.0010033
2019-006-0013.mrg | 2019-006-0013.sam | 37 | 1000.85 | 1000.85 | 0.11556 | 0.0026856 | 34.41 | 34.4084 | 0.0002306 | -0.0016098 | 34.41 | 34.4106 | 0.00020643 | 0.00057221
2019-006-0013.mrg | 2019-006-0013.sam | 124 | 800.054 | 800.079 | 0.11175 | 0.025208 | 34.3121 | 34.3106 | 0.0004017 | -0.0015488 | 34.3121 | 34.3125 | 0.00034369 | 0.00040054
2019-006-0013.mrg | 2019-006-0013.sam | 214 | 5.172 | 5.1918 | 0.10958 | 0.019855 | 32.3655 | 32.3621 | 0.00056541 | -0.0033608 | 32.3655 | 32.3635 | 0.00078339 | -0.0019646
2019-006-0019.mrg | 2019-006-0019.sam | 240 | 5.319 | 5.3204 | 0.091488 | 0.0013671 | 32.3419 | 32.3382 | 0.00035 | -0.0036926 | 32.3419 | 32.3401 | 0.00025866 | -0.0018425
2019-006-0021.mrg | 2019-006-0021.sam | 127 | 5.845 | 5.8051 | 0.097374 | -0.039937 | 32.2867 | 32.2771 | 0.0027959 | -0.0095749 | 32.2867 | 32.2808 | 0.0028966 | -0.0058899
2019-006-0022.mrg | 2019-006-0022.sam | 192 | 5.154 | 5.1369 | 0.11009 | -0.017104 | 32.3075 | 32.3009 | 0.0011912 | -0.0065689 | 32.3075 | 32.304 | 0.00061638 | -0.0035172
2019-006-0023.mrg | 2019-006-0023.sam | 31 | 2000.47 | 2000.48 | 0.11337 | 0.014282 | 34.5881 | 34.5831 | 0.00026923 | -0.0050163 | 34.5881 | 34.5865 | 0.00027599 | -0.0016289
2019-006-0023.mrg | 2019-006-0023.sam | 53 | 1500.42 | 1500.42 | 0.11135 | -0.00354 | 34.5153 | 34.5036 | 0.00029078 | -0.011723 | 34.5153 | 34.5063 | 0.00020375 | -0.0090103
2019-006-0023.mrg | 2019-006-0023.sam | 91 | 1250.53 | 1250.5 | 0.11084 | -0.024292 | 34.4556 | 34.4523 | 0.00028453 | -0.0033188 | 34.4556 | 34.4549 | 0.00024428 | -0.00069809
2019-006-0023.mrg | 2019-006-0023.sam | 80 | 1000.03 | 1000.02 | 0.10662 | -0.010742 | 34.367 | 34.3654 | 0.00029824 | -0.0015907 | 34.367 | 34.3679 | 0.00027498 | 0.00094223
2019-006-0023.mrg | 2019-006-0023.sam | 125 | 800.348 | 800.314 | 0.11015 | -0.033569 | 34.2758 | 34.2702 | 0.00031399 | -0.0055924 | 34.2758 | 34.2725 | 0.00025056 | -0.003315
2019-006-0023.mrg | 2019-006-0023.sam | 255 | 5.776 | 5.8206 | 0.11732 | 0.044584 | 32.3432 | 32.3382 | 0.0011345 | -0.0049667 | 32.3432 | 32.3411 | 0.001154 | -0.0021172
2019-006-0026.mrg | 2019-006-0026.sam | 233 | 6.085 | 6.0775 | 0.096007 | -0.0075064 | 32.3181 | 32.315 | 0.0010035 | -0.0031319 | 32.3181 | 32.3162 | 0.0014399 | -0.0018692
2019-006-0027.mrg | 2019-006-0027.sam | 155 | 5.341 | 5.3679 | 0.11095 | 0.026929 | 32.3383 | 32.336 | 0.00033891 | -0.0023346 | 32.3383 | 32.3374 | 0.0003025 | -0.00093842
2019-006-0028.mrg | 2019-006-0028.sam | 153 | 5.611 | 5.6232 | 0.12585 | 0.012228 | 32.3438 | 32.3408 | 0.00029712 | -0.0030212 | 32.3438 | 32.3421 | 0.00023535 | -0.0017471
2019-006-0032.mrg | 2019-006-0032.sam | 184 | 3000.42 | 3000.46 | 0.083839 | 0.041992 | 34.6478 | 34.6431 | 0.00028536 | -0.0047379 | 34.6478 | 34.6464 | 0.00026027 | -0.0013733
2019-006-0032.mrg | 2019-006-0032.sam | 186 | 3000.6 | 3000.54 | 0.10258 | -0.061279 | 34.6447 | 34.6431 | 0.00028818 | -0.0016365 | 34.6447 | 34.6464 | 0.00024911 | 0.0017281
2019-006-0032.mrg | 2019-006-0032.sam | 68 | 2501.13 | 2501.14 | 0.12808 | 0.0097656 | 34.6272 | 34.6222 | 0.00027372 | -0.0049553 | 34.6272 | 34.6254 | 0.00025605 | -0.0018463
2019-006-0032.mrg | 2019-006-0032.sam | 70 | 2000.77 | 2000.77 | 0.11661 | -0.0014648 | 34.5778 | 34.578 | 0.00029896 | 0.00021744 | 34.5778 | 34.581 | 0.00022098 | 0.0031929
2019-006-0032.mrg | 2019-006-0032.sam | 92 | 1499.85 | 1499.82 | 0.12145 | -0.022705 | 34.4924 | 34.4894 | 0.00030034 | -0.0030403 | 34.4924 | 34.4923 | 0.00020459 | -0.000103
2019-006-0032.mrg | 2019-006-0032.sam | 71 | 1249.91 | 1249.99 | 0.12763 | 0.083618 | 34.4394 | 34.4354 | 0.00025783 | -0.0040398 | 34.4394 | 34.438 | 0.00026543 | -0.0014305
2019-006-0032.mrg | 2019-006-0032.sam | 185 | 1000.53 | 1000.55 | 0.11167 | 0.026062 | 34.3464 | 34.3425 | 0.00027784 | -0.0039406 | 34.3464 | 34.3451 | 0.00024519 | -0.0013351
2019-006-0032.mrg | 2019-006-0032.sam | 108 | 4.958 | 4.9421 | 0.10812 | -0.015898 | 32.3792 | 32.3748 | 0.00028591 | -0.0043907 | 32.3792 | 32.3776 | 0.00020812 | -0.001648
2019-006-0040.mrg | 2019-006-0040.sam | 147 | 4.756 | 4.7831 | 0.10645 | 0.027129 | 32.3714 | 32.3696 | 0.00027983 | -0.0018349 | 32.3714 | 32.3708 | 0.00021394 | -0.00059891
2019-006-0042.mrg | 2019-006-0042.sam | 47 | 3358.74 | 3358.74 | 0.11119 | -0.0043945 | 34.6567 | 34.6527 | 0.00032084 | -0.0040207 | 34.6567 | 34.656 | 0.00025593 | -0.0006752
2019-006-0042.mrg | 2019-006-0042.sam | 46 | 2999.61 | 2999.58 | 0.11878 | -0.027588 | 34.648 | 34.6453 | 0.00024322 | -0.0026894 | 34.648 | 34.6486 | 0.00028413 | 0.00058746
2019-006-0042.mrg | 2019-006-0042.sam | 42 | 2499.8 | 2499.8 | 0.11385 | -0.0029297 | 34.625 | 34.6232 | 0.00027209 | -0.0018005 | 34.625 | 34.6262 | 0.00027956 | 0.0012131
2019-006-0042.mrg | 2019-006-0042.sam | 79 | 2250.6 | 2250.6 | 0.12128 | -0.0021973 | 34.6086 | 34.6053 | 0.00024067 | -0.0032844 | 34.6086 | 34.6082 | 0.00026843 | -0.00041962
2019-006-0042.mrg | 2019-006-0042.sam | 47 | 2000.31 | 2000.28 | 0.11078 | -0.025146 | 34.5817 | 34.5788 | 0.00024873 | -0.0029259 | 34.5817 | 34.5816 | 0.00025074 | -6.8665e-05
2019-006-0042.mrg | 2019-006-0042.sam | 69 | 1750.11 | 1750.13 | 0.099493 | 0.015869 | 34.5423 | 34.5396 | 0.00028167 | -0.0027428 | 34.5423 | 34.5422 | 0.00022994 | -0.00013733
2019-006-0042.mrg | 2019-006-0042.sam | 46 | 1500.43 | 1500.43 | 0.12767 | 0.00012207 | 34.4913 | 34.4898 | 0.00031395 | -0.0015068 | 34.4913 | 34.4924 | 0.00033467 | 0.001133
2019-006-0042.mrg | 2019-006-0042.sam | 36 | 1250.75 | 1250.73 | 0.13448 | -0.014282 | 34.4267 | 34.4244 | 0.00026095 | -0.0023079 | 34.4267 | 34.4269 | 0.00033912 | 0.00017166
2019-006-0042.mrg | 2019-006-0042.sam | 113 | 999.083 | 999.077 | 0.091517 | -0.0059815 | 34.3546 | 34.3534 | 0.00029423 | -0.001236 | 34.3546 | 34.3558 | 0.00036286 | 0.0012436
2019-006-0042.mrg | 2019-006-0042.sam | 33 | 801.601 | 801.614 | 0.12089 | 0.012512 | 34.2881 | 34.2881 | 0.00039595 | 1.5259e-05 | 34.2881 | 34.2904 | 0.00029701 | 0.0022888
2019-006-0042.mrg | 2019-006-0042.sam | 54 | 600.58 | 600.57 | 0.12557 | -0.0096436 | 34.152 | 34.1495 | 0.00029669 | -0.0025101 | 34.152 | 34.1519 | 0.00026468 | -6.1035e-05
2019-006-0042.mrg | 2019-006-0042.sam | 97 | 400.152 | 400.23 | 0.11584 | 0.077606 | 34.001 | 33.9992 | 0.00027122 | -0.0018082 | 34.001 | 34.0018 | 0.00033996 | 0.00078583
2019-006-0042.mrg | 2019-006-0042.sam | 50 | 301.25 | 301.242 | 0.12268 | -0.007782 | 33.9414 | 33.9395 | 0.00026796 | -0.0018692 | 33.9414 | 33.942 | 0.00023702 | 0.00059128
2019-006-0042.mrg | 2019-006-0042.sam | 76 | 251.104 | 251.135 | 0.11548 | 0.03096 | 33.903 | 33.9008 | 0.00028006 | -0.0022163 | 33.903 | 33.9032 | 0.00023285 | 0.00020981
2019-006-0042.mrg | 2019-006-0042.sam | 45 | 200.45 | 200.466 | 0.11375 | 0.016159 | 33.8932 | 33.8896 | 0.00024651 | -0.0036011 | 33.8932 | 33.8921 | 0.00024237 | -0.0010643
2019-006-0042.mrg | 2019-006-0042.sam | 66 | 174.967 | 174.963 | 0.12235 | -0.003952 | 33.8639 | 33.8632 | 0.00074556 | -0.00069046 | 33.8639 | 33.8659 | 0.00072397 | 0.0019989
2019-006-0042.mrg | 2019-006-0042.sam | 125 | 149.843 | 149.842 | 0.12915 | -0.00061035 | 33.7924 | 33.7778 | 0.0031998 | -0.014606 | 33.7924 | 33.7781 | 0.0020758 | -0.014324
2019-006-0042.mrg | 2019-006-0042.sam | 163 | 125.134 | 125.134 | 0.09213 | -0.00012207 | 33.4581 | 33.448 | 0.00050135 | -0.01009 | 33.4581 | 33.4507 | 0.00047427 | -0.0073929
2019-006-0042.mrg | 2019-006-0042.sam | 120 | 99.971 | 99.9276 | 0.088617 | -0.043419 | 32.765 | 32.7602 | 0.0018695 | -0.0048065 | 32.765 | 32.7619 | 0.0018811 | -0.0030518
2019-006-0042.mrg | 2019-006-0042.sam | 79 | 75.668 | 75.5711 | 0.10512 | -0.096939 | 32.3889 | 32.3842 | 0.00039235 | -0.0047226 | 32.3889 | 32.3871 | 0.00029715 | -0.0018234
2019-006-0042.mrg | 2019-006-0042.sam | 129 | 50.398 | 50.4283 | 0.08334 | 0.030258 | 32.3937 | 32.3909 | 0.001496 | -0.0027542 | 32.3937 | 32.3924 | 0.0012595 | -0.0013008
2019-006-0042.mrg | 2019-006-0042.sam | 73 | 25.246 | 25.2375 | 0.10778 | -0.0085068 | 32.3773 | 32.3732 | 0.00027495 | -0.0041237 | 32.3773 | 32.3752 | 0.00018616 | -0.0021133
2019-006-0042.mrg | 2019-006-0042.sam | 66 | 10.632 | 10.6401 | 0.11454 | 0.0080757 | 32.3771 | 32.3731 | 0.00030276 | -0.0040016 | 32.3771 | 32.3753 | 0.00028081 | -0.0018158
2019-006-0042.mrg | 2019-006-0042.sam | 157 | 5.296 | 5.2922 | 0.11447 | -0.0038023 | 32.375 | 32.3729 | 0.00027059 | -0.0020523 | 32.375 | 32.3753 | 0.0002357 | 0.00033951
2019-006-0044.mrg | 2019-006-0044.sam | 205 | 5.058 | 5.0809 | 0.11327 | 0.022863 | 32.385 | 32.3787 | 0.00034883 | -0.0063477 | 32.385 | 32.3798 | 0.00027433 | -0.0051727
2019-006-0045.mrg | 2019-006-0045.sam | 131 | 3499.43 | 3499.45 | 0.12298 | 0.022461 | 34.6656 | 34.6614 | 0.0003025 | -0.0042267 | 34.6656 | 34.6647 | 0.0002533 | -0.00092697
2019-006-0045.mrg | 2019-006-0045.sam | 135 | 3499.53 | 3499.55 | 0.11279 | 0.019287 | 34.6643 | 34.6613 | 0.00028716 | -0.0029602 | 34.6643 | 34.6647 | 0.00026249 | 0.00038147
2019-006-0045.mrg | 2019-006-0045.sam | 31 | 2999.34 | 2999.35 | 0.11315 | 0.0095215 | 34.6497 | 34.6489 | 0.00021892 | -0.0008049 | 34.6497 | 34.6523 | 0.00032487 | 0.0025787
2019-006-0045.mrg | 2019-006-0045.sam | 47 | 2500.63 | 2500.63 | 0.12239 | 0.0012207 | 34.6278 | 34.6259 | 0.00026318 | -0.0019226 | 34.6278 | 34.6289 | 0.00021954 | 0.0011482
2019-006-0045.mrg | 2019-006-0045.sam | 35 | 2000.29 | 2000.29 | 0.11442 | -0.0031738 | 34.5852 | 34.5824 | 0.00027145 | -0.0027924 | 34.5852 | 34.5853 | 0.00025078 | 6.8665e-05
2019-006-0045.mrg | 2019-006-0045.sam | 72 | 1500.97 | 1500.98 | 0.11987 | 0.0068359 | 34.4987 | 34.4971 | 0.00028973 | -0.0016174 | 34.4987 | 34.4996 | 0.00033063 | 0.00094604
2019-006-0045.mrg | 2019-006-0045.sam | 68 | 1249.95 | 1249.94 | 0.11914 | -0.011475 | 34.4248 | 34.4225 | 0.00026383 | -0.002346 | 34.4248 | 34.425 | 0.00023129 | 0.00021362
2019-006-0045.mrg | 2019-006-0045.sam | 151 | 1001.14 | 1001.17 | 0.11589 | 0.0354 | 34.3479 | 34.3468 | 0.00027695 | -0.001091 | 34.3479 | 34.3493 | 0.00023184 | 0.0014076
2019-006-0045.mrg | 2019-006-0045.sam | 139 | 5.796 | 5.8015 | 0.13166 | 0.0054679 | 32.3533 | 32.3482 | 0.004966 | -0.0051384 | 32.3533 | 32.3514 | 0.00027816 | -0.0019302
2019-006-0055.mrg | 2019-006-0055.sam | 120 | 5.496 | 5.5454 | 0.11558 | 0.0494 | 32.3519 | 32.3492 | 0.00026212 | -0.0026856 | 32.3519 | 32.3507 | 0.0002278 | -0.0012436
2019-006-0057.mrg | 2019-006-0057.sam | 86 | 3499.94 | 3499.91 | 0.11995 | -0.023682 | 34.6696 | 34.6641 | 0.00029578 | -0.0054665 | 34.6696 | 34.6679 | 0.00024427 | -0.0016823
2019-006-0057.mrg | 2019-006-0057.sam | 64 | 3000.98 | 3000.99 | 0.10398 | 0.0043945 | 34.6548 | 34.6497 | 0.00027862 | -0.0050812 | 34.6548 | 34.6533 | 0.00023442 | -0.0015488
2019-006-0057.mrg | 2019-006-0057.sam | 165 | 2500.59 | 2500.63 | 0.1161 | 0.038086 | 34.6278 | 34.6238 | 0.00029342 | -0.0039673 | 34.6278 | 34.6272 | 0.00024941 | -0.0006218
2019-006-0057.mrg | 2019-006-0057.sam | 41 | 2000.93 | 2000.92 | 0.13528 | -0.0074463 | 34.5856 | 34.5816 | 0.00038399 | -0.0040436 | 34.5856 | 34.5847 | 0.00025634 | -0.00086594
2019-006-0057.mrg | 2019-006-0057.sam | 42 | 1749.7 | 1749.7 | 0.1246 | -0.0031738 | 34.5491 | 34.5455 | 0.00028705 | -0.003582 | 34.5491 | 34.5487 | 0.00029795 | -0.00035095
2019-006-0057.mrg | 2019-006-0057.sam | 62 | 1500.54 | 1500.53 | 0.12813 | -0.012207 | 34.4988 | 34.4962 | 0.0002533 | -0.0025787 | 34.4988 | 34.4992 | 0.00026485 | 0.00037003
2019-006-0057.mrg | 2019-006-0057.sam | 166 | 5.473 | 5.4677 | 0.092168 | -0.0052953 | 32.3263 | 32.3212 | 0.001761 | -0.0050774 | 32.3263 | 32.3233 | 0.0044673 | -0.0029564
2019-006-0059.mrg | 2019-006-0059.sam | 229 | 5.36 | 5.3677 | 0.096918 | 0.0076766 | 32.3114 | 32.3073 | 0.00069931 | -0.0040817 | 32.3114 | 32.3081 | 0.0006651 | -0.003334
2019-006-0065.mrg | 2019-006-0065.sam | 103 | 3499.87 | 3499.86 | 0.087358 | -0.0073242 | 34.6704 | 34.666 | 0.00023188 | -0.0044098 | 34.6704 | 34.669 | 0.00023729 | -0.0013542
2019-006-0065.mrg | 2019-006-0065.sam | 127 | 3499.65 | 3499.6 | 0.11164 | -0.058105 | 34.67 | 34.666 | 0.00026062 | -0.0040131 | 34.67 | 34.6691 | 0.00025415 | -0.00091553
2019-006-0065.mrg | 2019-006-0065.sam | 18 | 3000.39 | 3000.37 | 0.14535 | -0.017334 | 34.6562 | 34.6519 | 0.00028201 | -0.0042953 | 34.6562 | 34.6548 | 0.0001974 | -0.0014343
2019-006-0065.mrg | 2019-006-0065.sam | 27 | 2501.19 | 2501.2 | 0.11719 | 0.016602 | 34.6317 | 34.6238 | 0.00027767 | -0.007946 | 34.6317 | 34.6264 | 0.00034602 | -0.0052643
2019-006-0065.mrg | 2019-006-0065.sam | 105 | 1999.5 | 1999.54 | 0.094701 | 0.039429 | 34.5848 | 34.5808 | 0.00024127 | -0.0040359 | 34.5848 | 34.5835 | 0.00026733 | -0.0013428
2019-006-0065.mrg | 2019-006-0065.sam | 51 | 1501.48 | 1501.47 | 0.11586 | -0.0046387 | 34.4982 | 34.4947 | 0.00029801 | -0.0034943 | 34.4982 | 34.4975 | 0.00036747 | -0.00069427
2019-006-0065.mrg | 2019-006-0065.sam | 30 | 1250.79 | 1250.79 | 0.12479 | -0.006958 | 34.4382 | 34.4348 | 0.00021871 | -0.003437 | 34.4382 | 34.437 | 0.00027954 | -0.0012321
2019-006-0065.mrg | 2019-006-0065.sam | 60 | 999.981 | 1000.01 | 0.12074 | 0.028442 | 34.3626 | 34.3586 | 0.00032184 | -0.0039902 | 34.3626 | 34.361 | 0.0003224 | -0.0015984
2019-006-0065.mrg | 2019-006-0065.sam | 141 | 5.221 | 5.2111 | 0.094376 | -0.0098867 | 32.3114 | 32.2989 | 0.00029619 | -0.01252 | 32.3114 | 32.3012 | 0.00031322 | -0.010227
2019-006-0066.mrg | 2019-006-0066.sam | 61 | 2000.26 | 2000.27 | 0.091333 | 0.015503 | 34.5793 | 34.5764 | 0.00024191 | -0.0028801 | 34.5793 | 34.5795 | 0.00021888 | 0.00015259
2019-006-0066.mrg | 2019-006-0066.sam | 44 | 1501.87 | 1501.83 | 0.10304 | -0.035278 | 34.4856 | 34.4819 | 0.00026158 | -0.0036545 | 34.4856 | 34.4845 | 0.00030127 | -0.0011291
2019-006-0066.mrg | 2019-006-0066.sam | 42 | 1001.36 | 1001.35 | 0.11904 | -0.011108 | 34.3538 | 34.352 | 0.00031468 | -0.0018463 | 34.3538 | 34.3543 | 0.00054896 | 0.00050354
2019-006-0066.mrg | 2019-006-0066.sam | 84 | 750.319 | 750.323 | 0.12141 | 0.0040283 | 34.2412 | 34.2386 | 0.00031223 | -0.0026436 | 34.2412 | 34.2412 | 0.00027669 | 2.2888e-05
2019-006-0066.mrg | 2019-006-0066.sam | 112 | 4.612 | 4.6249 | 0.096014 | 0.012937 | 32.2871 | 32.281 | 0.00038705 | -0.006134 | 32.2871 | 32.2832 | 0.00021424 | -0.0038719
2019-006-0069.mrg | 2019-006-0069.sam | 19 | 3499.88 | 3499.87 | 0.09827 | -0.0083008 | 34.6713 | 34.6655 | 0.00029465 | -0.005806 | 34.6713 | 34.6687 | 0.00029337 | -0.0026207
2019-006-0069.mrg | 2019-006-0069.sam | 205 | 2999.51 | 2999.49 | 0.10738 | -0.020752 | 34.6541 | 34.6498 | 0.00026968 | -0.0043144 | 34.6541 | 34.6527 | 0.00024881 | -0.0013924
2019-006-0069.mrg | 2019-006-0069.sam | 51 | 2499.89 | 2499.87 | 0.116 | -0.016113 | 34.6263 | 34.6237 | 0.00029355 | -0.0025673 | 34.6263 | 34.6264 | 0.00022656 | 0.00014496
2019-006-0069.mrg | 2019-006-0069.sam | 60 | 2000.29 | 2000.25 | 0.095015 | -0.04248 | 34.5816 | 34.5785 | 0.00024768 | -0.0030632 | 34.5816 | 34.5811 | 0.00022591 | -0.00051498
2019-006-0069.mrg | 2019-006-0069.sam | 25 | 1499.83 | 1499.8 | 0.098633 | -0.027832 | 34.5031 | 34.4972 | 0.00029403 | -0.0058594 | 34.5031 | 34.4995 | 0.00022027 | -0.0036201
2019-006-0069.mrg | 2019-006-0069.sam | 194 | 5.398 | 5.4018 | 0.11343 | 0.0037579 | 32.3271 | 32.3232 | 0.0007402 | -0.0039139 | 32.3271 | 32.3251 | 0.0016992 | -0.0020256
2019-006-0070.mrg | 2019-006-0070.sam | 103 | 5.416 | 5.3733 | 0.09134 | -0.042679 | 32.3461 | 32.343 | 0.00032632 | -0.0031052 | 32.3461 | 32.3441 | 0.00025929 | -0.0020256
2019-006-0072.mrg | 2019-006-0072.sam | 98 | 3500.28 | 3500.3 | 0.074647 | 0.017334 | 34.6681 | 34.6646 | 0.00028581 | -0.0035477 | 34.6681 | 34.6676 | 0.00023583 | -0.0005455
2019-006-0072.mrg | 2019-006-0072.sam | 65 | 3001.12 | 3001.15 | 0.1278 | 0.027588 | 34.6508 | 34.6479 | 0.00027004 | -0.0028877 | 34.6508 | 34.6507 | 0.00024853 | -8.3923e-05
2019-006-0072.mrg | 2019-006-0072.sam | 17 | 2500.13 | 2500.14 | 0.12166 | 0.0036621 | 34.622 | 34.6182 | 0.00034168 | -0.0038261 | 34.622 | 34.6208 | 0.00038088 | -0.0011673
2019-006-0072.mrg | 2019-006-0072.sam | 45 | 2000.4 | 2000.4 | 0.11404 | 0.0062256 | 34.5768 | 34.5698 | 0.00027075 | -0.0069847 | 34.5768 | 34.5722 | 0.00023122 | -0.0045891
2019-006-0072.mrg | 2019-006-0072.sam | 25 | 1500.65 | 1500.64 | 0.12918 | -0.0043945 | 34.4951 | 34.4888 | 0.00027273 | -0.00634 | 34.4951 | 34.491 | 0.00035254 | -0.0041122
2019-006-0072.mrg | 2019-006-0072.sam | 269 | 5.016 | 5.026 | 0.11091 | 0.010011 | 32.3881 | 32.3794 | 0.00026181 | -0.0086937 | 32.3881 | 32.3814 | 0.00020019 | -0.0066948
2019-006-0074.mrg | 2019-006-0074.sam | 46 | 2000.71 | 2000.73 | 0.12861 | 0.015259 | 34.5762 | 34.573 | 0.00030631 | -0.0032272 | 34.5762 | 34.5757 | 0.00025891 | -0.00053787
2019-006-0074.mrg | 2019-006-0074.sam | 45 | 1500.6 | 1500.6 | 0.12481 | -0.0019531 | 34.4915 | 34.4889 | 0.00029749 | -0.0026169 | 34.4915 | 34.4912 | 0.00026973 | -0.00030518
2019-006-0074.mrg | 2019-006-0074.sam | 65 | 1000.53 | 1000.52 | 0.11252 | -0.010315 | 34.3593 | 34.3569 | 0.00029477 | -0.0023918 | 34.3593 | 34.359 | 0.00025324 | -0.00034714
2019-006-0074.mrg | 2019-006-0074.sam | 70 | 751.047 | 751.038 | 0.1109 | -0.0093384 | 34.2511 | 34.2497 | 0.00026264 | -0.0014458 | 34.2511 | 34.2516 | 0.0002865 | 0.00050736
2019-006-0074.mrg | 2019-006-0074.sam | 226 | 5.753 | 5.7594 | 0.11425 | 0.0063677 | 32.3826 | 32.3796 | 0.00039235 | -0.0030136 | 32.3826 | 32.3816 | 0.00043425 | -0.0010452
2019-006-0075.mrg | 2019-006-0075.sam | 144 | 5.193 | 5.1658 | 0.094076 | -0.027173 | 32.3949 | 32.3928 | 0.00027206 | -0.0021133 | 32.3949 | 32.3934 | 0.00024389 | -0.0015259
2019-006-0077.mrg | 2019-006-0077.sam | 13 | 3999.02 | 3998.99 | 0.11705 | -0.031494 | 34.679 | 34.6752 | 0.00033017 | -0.0038376 | 34.679 | 34.6784 | 0.00027527 | -0.00060654
2019-006-0077.mrg | 2019-006-0077.sam | 30 | 3500.46 | 3500.49 | 0.13154 | 0.028564 | 34.671 | 34.6666 | 0.00024995 | -0.0044403 | 34.671 | 34.6697 | 0.00027476 | -0.001255
2019-006-0077.mrg | 2019-006-0077.sam | 90 | 3000.47 | 3000.49 | 0.10672 | 0.019287 | 34.6543 | 34.6508 | 0.00025894 | -0.0035286 | 34.6543 | 34.6536 | 0.00023263 | -0.00073624
2019-006-0077.mrg | 2019-006-0077.sam | 33 | 2499.58 | 2499.6 | 0.1073 | 0.020996 | 34.6276 | 34.6242 | 0.00034372 | -0.003376 | 34.6276 | 34.6268 | 0.00030993 | -0.00079727
2019-006-0077.mrg | 2019-006-0077.sam | 41 | 1999.58 | 1999.58 | 0.10591 | -0.0025635 | 34.5797 | 34.5764 | 0.00024728 | -0.0032844 | 34.5797 | 34.5791 | 0.00029423 | -0.00062942
2019-006-0077.mrg | 2019-006-0077.sam | 21 | 1499.76 | 1499.76 | 0.12885 | 0.0019531 | 34.4957 | 34.4934 | 0.0002439 | -0.0023041 | 34.4957 | 34.4957 | 0.00033148 | 7.6294e-06
2019-006-0077.mrg | 2019-006-0077.sam | 45 | 1249.47 | 1249.49 | 0.10598 | 0.025146 | 34.4383 | 34.436 | 0.00029814 | -0.0022621 | 34.4383 | 34.4383 | 0.00037011 | 4.5776e-05
2019-006-0077.mrg | 2019-006-0077.sam | 22 | 1000.67 | 1000.65 | 0.12143 | -0.01471 | 34.3655 | 34.3638 | 0.00037303 | -0.0016861 | 34.3655 | 34.3662 | 0.0003534 | 0.00072098
2019-006-0077.mrg | 2019-006-0077.sam | 185 | 5.355 | 5.3104 | 0.10322 | -0.044611 | 32.393 | 32.3883 | 0.00031495 | -0.0047302 | 32.393 | 32.3902 | 0.00031522 | -0.0028114
2019-006-0079.mrg | 2019-006-0079.sam | 275 | 2000.88 | 2000.87 | 0.12671 | -0.012939 | 34.5776 | 34.574 | 0.00026936 | -0.0036316 | 34.5776 | 34.5767 | 0.00026378 | -0.00086212
2019-006-0102.mrg | 2019-006-0102.sam | 127 | 1249.64 | 1249.7 | 0.08464 | 0.052246 | 34.4621 | 34.4584 | 0.00030696 | -0.0036774 | 34.4621 | 34.4617 | 0.0003082 | -0.00041199
2019-006-0102.mrg | 2019-006-0102.sam | 60 | 1250.19 | 1250.17 | 0.11951 | -0.011841 | 34.4663 | 34.4584 | 0.00027235 | -0.0079193 | 34.4663 | 34.4619 | 0.0002273 | -0.0044441
2019-006-0102.mrg | 2019-006-0102.sam | 176 | 999.374 | 999.332 | 0.093855 | -0.04248 | 34.3949 | 34.3903 | 0.00025127 | -0.0045624 | 34.3949 | 34.3933 | 0.00023115 | -0.0015564
2019-006-0102.mrg | 2019-006-0102.sam | 111 | 800.871 | 800.879 | 0.10698 | 0.0082398 | 34.3019 | 34.2985 | 0.00027876 | -0.0033569 | 34.3019 | 34.3014 | 0.00023516 | -0.00050736
2019-006-0102.mrg | 2019-006-0102.sam | 161 | 600.181 | 600.248 | 0.096628 | 0.067261 | 34.2029 | 34.1981 | 0.00026696 | -0.0047913 | 34.2029 | 34.2009 | 0.00025815 | -0.0020447
2019-006-0102.mrg | 2019-006-0102.sam | 165 | 5.301 | 5.2889 | 0.093541 | -0.012134 | 32.4479 | 32.4412 | 0.00039654 | -0.0067024 | 32.4479 | 32.4436 | 0.0003151 | -0.0043373
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 29
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2019-006-0006.mrg | 2019-006-0006.sam | 12 | 12 | 2019-006-0006.mrg & 2019-006-0006.sam
2 | 2019-006-0010.mrg | 2019-006-0010.sam | 14 | 14 | 2019-006-0010.mrg & 2019-006-0010.sam
3 | 2019-006-0011.mrg | 2019-006-0011.sam | 16 | 16 | 2019-006-0011.mrg & 2019-006-0011.sam
4 | 2019-006-0013.mrg | 2019-006-0013.sam | 18 | 22 | 2019-006-0013.mrg & 2019-006-0013.sam
5 | 2019-006-0019.mrg | 2019-006-0019.sam | 24 | 24 | 2019-006-0019.mrg & 2019-006-0019.sam
6 | 2019-006-0021.mrg | 2019-006-0021.sam | 26 | 26 | 2019-006-0021.mrg & 2019-006-0021.sam
7 | 2019-006-0022.mrg | 2019-006-0022.sam | 28 | 28 | 2019-006-0022.mrg & 2019-006-0022.sam
8 | 2019-006-0023.mrg | 2019-006-0023.sam | 30 | 35 | 2019-006-0023.mrg & 2019-006-0023.sam
9 | 2019-006-0026.mrg | 2019-006-0026.sam | 37 | 37 | 2019-006-0026.mrg & 2019-006-0026.sam
10 | 2019-006-0027.mrg | 2019-006-0027.sam | 39 | 39 | 2019-006-0027.mrg & 2019-006-0027.sam
11 | 2019-006-0028.mrg | 2019-006-0028.sam | 41 | 41 | 2019-006-0028.mrg & 2019-006-0028.sam
12 | 2019-006-0032.mrg | 2019-006-0032.sam | 43 | 50 | 2019-006-0032.mrg & 2019-006-0032.sam
13 | 2019-006-0040.mrg | 2019-006-0040.sam | 52 | 52 | 2019-006-0040.mrg & 2019-006-0040.sam
14 | 2019-006-0042.mrg | 2019-006-0042.sam | 54 | 77 | 2019-006-0042.mrg & 2019-006-0042.sam
15 | 2019-006-0044.mrg | 2019-006-0044.sam | 79 | 79 | 2019-006-0044.mrg & 2019-006-0044.sam
16 | 2019-006-0045.mrg | 2019-006-0045.sam | 81 | 89 | 2019-006-0045.mrg & 2019-006-0045.sam
17 | 2019-006-0055.mrg | 2019-006-0055.sam | 91 | 91 | 2019-006-0055.mrg & 2019-006-0055.sam
18 | 2019-006-0057.mrg | 2019-006-0057.sam | 93 | 99 | 2019-006-0057.mrg & 2019-006-0057.sam
19 | 2019-006-0059.mrg | 2019-006-0059.sam | 101 | 101 | 2019-006-0059.mrg & 2019-006-0059.sam
20 | 2019-006-0065.mrg | 2019-006-0065.sam | 103 | 111 | 2019-006-0065.mrg & 2019-006-0065.sam
21 | 2019-006-0066.mrg | 2019-006-0066.sam | 113 | 117 | 2019-006-0066.mrg & 2019-006-0066.sam
22 | 2019-006-0069.mrg | 2019-006-0069.sam | 119 | 124 | 2019-006-0069.mrg & 2019-006-0069.sam
23 | 2019-006-0070.mrg | 2019-006-0070.sam | 126 | 126 | 2019-006-0070.mrg & 2019-006-0070.sam
24 | 2019-006-0072.mrg | 2019-006-0072.sam | 128 | 133 | 2019-006-0072.mrg & 2019-006-0072.sam
25 | 2019-006-0074.mrg | 2019-006-0074.sam | 135 | 139 | 2019-006-0074.mrg & 2019-006-0074.sam
26 | 2019-006-0075.mrg | 2019-006-0075.sam | 141 | 141 | 2019-006-0075.mrg & 2019-006-0075.sam
27 | 2019-006-0077.mrg | 2019-006-0077.sam | 143 | 151 | 2019-006-0077.mrg & 2019-006-0077.sam
28 | 2019-006-0079.mrg | 2019-006-0079.sam | 153 | 153 | 2019-006-0079.mrg & 2019-006-0079.sam
29 | 2019-006-0102.mrg | 2019-006-0102.sam | 155 | 160 | 2019-006-0102.mrg & 2019-006-0102.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 29
2 | (Sal1 - Sal_Bot) | Pressure | 29
3 | (Sal0 - Sal_Bot) | File Pair # | 29
4 | (Sal1 - Sal_Bot) | File Pair # | 29

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.068088 |  | X | CALC Y
m= | -4.4539e-07 | 1.8657e-07 |  | Std Dev of Resid |  | 0.0015121 |  | 600.248 | -0.002772
b= | -0.0025046 | 0.00040019 |  | Resid Sum of Squares |  | 0.00017835 |  | 3998.99 | -0.0042857
# of Values Extracted = |  |  | 121
# of Values Used in Fit = |  |  | 80
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0297 | 31.6927 | 0.00094052 | -0.0062809 | -0.0025069 | 0.003774 |  | 1250.41 | -0.0036583
2 | NO | 5.1496 | 31.7169 | 0.00042182 | -0.0077171 | -0.0025069 | 0.0052102 |  | 1250 | -0.0035019
3 | NO | 5.0349 | 32.4938 | 0.00039708 | -0.0039482 | -0.0025069 | 0.0014414 |  | 1000.85 | -0.0016098
4 | YES | 1250.41 | 34.4777 | 0.00025257 | -0.0036583 | -0.0030615 | 0.00059676 |  | 800.079 | -0.0015488
4 | YES | 1250 | 34.4778 | 0.0002617 | -0.0035019 | -0.0030613 | 0.00044054 |  | 2000.48 | -0.0050163
4 | YES | 1000.85 | 34.4084 | 0.0002306 | -0.0016098 | -0.0029504 | -0.0013406 |  | 1250.5 | -0.0033188
4 | YES | 800.079 | 34.3106 | 0.0004017 | -0.0015488 | -0.002861 | -0.0013122 |  | 1000.02 | -0.0015907
4 | NO | 5.1918 | 32.3621 | 0.00056541 | -0.0033608 | -0.0025069 | 0.00085383 |  | 800.314 | -0.0055924
5 | NO | 5.3204 | 32.3382 | 0.00035 | -0.0036926 | -0.002507 | 0.0011856 |  | 3000.46 | -0.0047379
6 | NO | 5.8051 | 32.2771 | 0.0027959 | -0.0095749 | -0.0025072 | 0.0070677 |  | 3000.54 | -0.0016365
7 | NO | 5.1369 | 32.3009 | 0.0011912 | -0.0065689 | -0.0025069 | 0.004062 |  | 2501.14 | -0.0049553
8 | YES | 2000.48 | 34.5831 | 0.00026923 | -0.0050163 | -0.0033956 | 0.0016207 |  | 2000.77 | 0.00021744
8 | NO | 1500.42 | 34.5036 | 0.00029078 | -0.011723 | -0.0031729 | 0.0085497 |  | 1499.82 | -0.0030403
8 | YES | 1250.5 | 34.4523 | 0.00028453 | -0.0033188 | -0.0030616 | 0.00025722 |  | 1249.99 | -0.0040398
8 | YES | 1000.02 | 34.3654 | 0.00029824 | -0.0015907 | -0.00295 | -0.0013593 |  | 1000.55 | -0.0039406
8 | YES | 800.314 | 34.2702 | 0.00031399 | -0.0055924 | -0.0028611 | 0.0027313 |  | 3358.74 | -0.0040207
8 | NO | 5.8206 | 32.3382 | 0.0011345 | -0.0049667 | -0.0025072 | 0.0024595 |  | 2999.58 | -0.0026894
9 | NO | 6.0775 | 32.315 | 0.0010035 | -0.0031319 | -0.0025073 | 0.00062455 |  | 2499.8 | -0.0018005
10 | NO | 5.3679 | 32.336 | 0.00033891 | -0.0023346 | -0.002507 | -0.00017241 |  | 2250.6 | -0.0032844
11 | NO | 5.6232 | 32.3408 | 0.00029712 | -0.0030212 | -0.0025071 | 0.00051413 |  | 2000.28 | -0.0029259
12 | YES | 3000.46 | 34.6431 | 0.00028536 | -0.0047379 | -0.003841 | 0.00089686 |  | 1750.13 | -0.0027428
12 | YES | 3000.54 | 34.6431 | 0.00028818 | -0.0016365 | -0.003841 | -0.0022045 |  | 1500.43 | -0.0015068
12 | YES | 2501.14 | 34.6222 | 0.00027372 | -0.0049553 | -0.0036186 | 0.0013367 |  | 1250.73 | -0.0023079
12 | YES | 2000.77 | 34.578 | 0.00029896 | 0.00021744 | -0.0033957 | -0.0036132 |  | 999.077 | -0.001236
12 | YES | 1499.82 | 34.4894 | 0.00030034 | -0.0030403 | -0.0031726 | -0.00013231 |  | 801.614 | 1.5259e-05
12 | YES | 1249.99 | 34.4354 | 0.00025783 | -0.0040398 | -0.0030613 | 0.00097842 |  | 600.57 | -0.0025101
12 | YES | 1000.55 | 34.3425 | 0.00027784 | -0.0039406 | -0.0029502 | 0.00099033 |  | 3499.45 | -0.0042267
12 | NO | 4.9421 | 32.3748 | 0.00028591 | -0.0043907 | -0.0025068 | 0.0018839 |  | 3499.55 | -0.0029602
13 | NO | 4.7831 | 32.3696 | 0.00027983 | -0.0018349 | -0.0025067 | -0.00067187 |  | 2999.35 | -0.0008049
14 | YES | 3358.74 | 34.6527 | 0.00032084 | -0.0040207 | -0.0040006 | 2.0128e-05 |  | 2500.63 | -0.0019226
14 | YES | 2999.58 | 34.6453 | 0.00024322 | -0.0026894 | -0.0038406 | -0.0011512 |  | 2000.29 | -0.0027924
14 | YES | 2499.8 | 34.6232 | 0.00027209 | -0.0018005 | -0.003618 | -0.0018175 |  | 1500.98 | -0.0016174
14 | YES | 2250.6 | 34.6053 | 0.00024067 | -0.0032844 | -0.003507 | -0.00022256 |  | 1249.94 | -0.002346
14 | YES | 2000.28 | 34.5788 | 0.00024873 | -0.0029259 | -0.0033955 | -0.00046965 |  | 1001.17 | -0.001091
14 | YES | 1750.13 | 34.5396 | 0.00028167 | -0.0027428 | -0.0032841 | -0.00054133 |  | 3499.91 | -0.0054665
14 | YES | 1500.43 | 34.4898 | 0.00031395 | -0.0015068 | -0.0031729 | -0.0016661 |  | 3000.99 | -0.0050812
14 | YES | 1250.73 | 34.4244 | 0.00026095 | -0.0023079 | -0.0030617 | -0.00075378 |  | 2500.63 | -0.0039673
14 | YES | 999.077 | 34.3534 | 0.00029423 | -0.001236 | -0.0029496 | -0.0017136 |  | 2000.92 | -0.0040436
14 | YES | 801.614 | 34.2881 | 0.00039595 | 1.5259e-05 | -0.0028616 | -0.0028769 |  | 1749.7 | -0.003582
14 | YES | 600.57 | 34.1495 | 0.00029669 | -0.0025101 | -0.0027721 | -0.00026203 |  | 1500.53 | -0.0025787
14 | NO | 400.23 | 33.9992 | 0.00027122 | -0.0018082 | -0.0026829 | -0.0008747 |  | 3499.86 | -0.0044098
14 | NO | 301.242 | 33.9395 | 0.00026796 | -0.0018692 | -0.0026388 | -0.00076958 |  | 3499.6 | -0.0040131
14 | NO | 251.135 | 33.9008 | 0.00028006 | -0.0022163 | -0.0026165 | -0.00040012 |  | 3000.37 | -0.0042953
14 | NO | 200.466 | 33.8896 | 0.00024651 | -0.0036011 | -0.0025939 | 0.0010072 |  | 2501.2 | -0.007946
14 | NO | 174.963 | 33.8632 | 0.00074556 | -0.00069046 | -0.0025825 | -0.0018921 |  | 1999.54 | -0.0040359
14 | NO | 149.842 | 33.7778 | 0.0031998 | -0.014606 | -0.0025713 | 0.012035 |  | 1501.47 | -0.0034943
14 | NO | 125.134 | 33.448 | 0.00050135 | -0.01009 | -0.0025603 | 0.0075296 |  | 1250.79 | -0.003437
14 | NO | 99.9276 | 32.7602 | 0.0018695 | -0.0048065 | -0.0025491 | 0.0022574 |  | 1000.01 | -0.0039902
14 | NO | 75.5711 | 32.3842 | 0.00039235 | -0.0047226 | -0.0025383 | 0.0021843 |  | 2000.27 | -0.0028801
14 | NO | 50.4283 | 32.3909 | 0.001496 | -0.0027542 | -0.0025271 | 0.00022714 |  | 1501.83 | -0.0036545
14 | NO | 25.2375 | 32.3732 | 0.00027495 | -0.0041237 | -0.0025159 | 0.0016078 |  | 1001.35 | -0.0018463
14 | NO | 10.6401 | 32.3731 | 0.00030276 | -0.0040016 | -0.0025093 | 0.0014923 |  | 750.323 | -0.0026436
14 | NO | 5.2922 | 32.3729 | 0.00027059 | -0.0020523 | -0.002507 | -0.00045466 |  | 3499.87 | -0.005806
15 | NO | 5.0809 | 32.3787 | 0.00034883 | -0.0063477 | -0.0025069 | 0.0038408 |  | 2999.49 | -0.0043144
16 | YES | 3499.45 | 34.6614 | 0.0003025 | -0.0042267 | -0.0040632 | 0.00016345 |  | 2499.87 | -0.0025673
16 | YES | 3499.55 | 34.6613 | 0.00028716 | -0.0029602 | -0.0040633 | -0.0011031 |  | 2000.25 | -0.0030632
16 | YES | 2999.35 | 34.6489 | 0.00021892 | -0.0008049 | -0.0038405 | -0.0030356 |  | 1499.8 | -0.0058594
16 | YES | 2500.63 | 34.6259 | 0.00026318 | -0.0019226 | -0.0036184 | -0.0016958 |  | 3500.3 | -0.0035477
16 | YES | 2000.29 | 34.5824 | 0.00027145 | -0.0027924 | -0.0033955 | -0.00060316 |  | 3001.15 | -0.0028877
16 | YES | 1500.98 | 34.4971 | 0.00028973 | -0.0016174 | -0.0031731 | -0.0015557 |  | 2500.14 | -0.0038261
16 | YES | 1249.94 | 34.4225 | 0.00026383 | -0.002346 | -0.0030613 | -0.00071528 |  | 2000.4 | -0.0069847
16 | YES | 1001.17 | 34.3468 | 0.00027695 | -0.001091 | -0.0029505 | -0.0018595 |  | 1500.64 | -0.00634
16 | NO | 5.8015 | 32.3482 | 0.004966 | -0.0051384 | -0.0025072 | 0.0026312 |  | 2000.73 | -0.0032272
17 | NO | 5.5454 | 32.3492 | 0.00026212 | -0.0026856 | -0.0025071 | 0.00017847 |  | 1500.6 | -0.0026169
18 | YES | 3499.91 | 34.6641 | 0.00029578 | -0.0054665 | -0.0040634 | 0.001403 |  | 1000.52 | -0.0023918
18 | YES | 3000.99 | 34.6497 | 0.00027862 | -0.0050812 | -0.0038412 | 0.00124 |  | 751.038 | -0.0014458
18 | YES | 2500.63 | 34.6238 | 0.00029342 | -0.0039673 | -0.0036184 | 0.00034892 |  | 3998.99 | -0.0038376
18 | YES | 2000.92 | 34.5816 | 0.00038399 | -0.0040436 | -0.0033958 | 0.00064778 |  | 3500.49 | -0.0044403
18 | YES | 1749.7 | 34.5455 | 0.00028705 | -0.003582 | -0.0032839 | 0.00029809 |  | 3000.49 | -0.0035286
18 | YES | 1500.53 | 34.4962 | 0.0002533 | -0.0025787 | -0.0031729 | -0.00059419 |  | 2499.6 | -0.003376
18 | NO | 5.4677 | 32.3212 | 0.001761 | -0.0050774 | -0.002507 | 0.0025703 |  | 1999.58 | -0.0032844
19 | NO | 5.3677 | 32.3073 | 0.00069931 | -0.0040817 | -0.002507 | 0.0015747 |  | 1499.76 | -0.0023041
20 | YES | 3499.86 | 34.666 | 0.00023188 | -0.0044098 | -0.0040634 | 0.00034637 |  | 1249.49 | -0.0022621
20 | YES | 3499.6 | 34.666 | 0.00026062 | -0.0040131 | -0.0040633 | -5.024e-05 |  | 1000.65 | -0.0016861
20 | YES | 3000.37 | 34.6519 | 0.00028201 | -0.0042953 | -0.0038409 | 0.0004544 |  | 2000.87 | -0.0036316
20 | YES | 2501.2 | 34.6238 | 0.00027767 | -0.007946 | -0.0036186 | 0.0043274 |  | 1249.7 | -0.0036774
20 | YES | 1999.54 | 34.5808 | 0.00024127 | -0.0040359 | -0.0033952 | 0.00064076 |  | 1250.17 | -0.0079193
20 | YES | 1501.47 | 34.4947 | 0.00029801 | -0.0034943 | -0.0031734 | 0.00032091 |  | 999.332 | -0.0045624
20 | YES | 1250.79 | 34.4348 | 0.00021871 | -0.003437 | -0.0030617 | 0.00037534 |  | 800.879 | -0.0033569
20 | YES | 1000.01 | 34.3586 | 0.00032184 | -0.0039902 | -0.00295 | 0.0010402 |  | 600.248 | -0.0047913
20 | NO | 5.2111 | 32.2989 | 0.00029619 | -0.01252 | -0.0025069 | 0.010013
21 | YES | 2000.27 | 34.5764 | 0.00024191 | -0.0028801 | -0.0033955 | -0.00051541
21 | YES | 1501.83 | 34.4819 | 0.00026158 | -0.0036545 | -0.0031735 | 0.00048097
21 | YES | 1001.35 | 34.352 | 0.00031468 | -0.0018463 | -0.0029506 | -0.0011043
21 | YES | 750.323 | 34.2386 | 0.00031223 | -0.0026436 | -0.0028388 | -0.00019521
21 | NO | 4.6249 | 32.281 | 0.00038705 | -0.006134 | -0.0025067 | 0.0036274
22 | YES | 3499.87 | 34.6655 | 0.00029465 | -0.005806 | -0.0040634 | 0.0017425
22 | YES | 2999.49 | 34.6498 | 0.00026968 | -0.0043144 | -0.0038406 | 0.00047386
22 | YES | 2499.87 | 34.6237 | 0.00029355 | -0.0025673 | -0.003618 | -0.0010507
22 | YES | 2000.25 | 34.5785 | 0.00024768 | -0.0030632 | -0.0033955 | -0.0003323
22 | YES | 1499.8 | 34.4972 | 0.00029403 | -0.0058594 | -0.0031726 | 0.0026868
22 | NO | 5.4018 | 32.3232 | 0.0007402 | -0.0039139 | -0.002507 | 0.0014069
23 | NO | 5.3733 | 32.343 | 0.00032632 | -0.0031052 | -0.002507 | 0.00059816
24 | YES | 3500.3 | 34.6646 | 0.00028581 | -0.0035477 | -0.0040636 | -0.00051594
24 | YES | 3001.15 | 34.6479 | 0.00027004 | -0.0028877 | -0.0038413 | -0.00095357
24 | YES | 2500.14 | 34.6182 | 0.00034168 | -0.0038261 | -0.0036181 | 0.00020799
24 | YES | 2000.4 | 34.5698 | 0.00027075 | -0.0069847 | -0.0033956 | 0.0035891
24 | YES | 1500.64 | 34.4888 | 0.00027273 | -0.00634 | -0.003173 | 0.003167
24 | NO | 5.026 | 32.3794 | 0.00026181 | -0.0086937 | -0.0025068 | 0.0061869
25 | YES | 2000.73 | 34.573 | 0.00030631 | -0.0032272 | -0.0033957 | -0.00016849
25 | YES | 1500.6 | 34.4889 | 0.00029749 | -0.0026169 | -0.003173 | -0.00055608
25 | YES | 1000.52 | 34.3569 | 0.00029477 | -0.0023918 | -0.0029502 | -0.00055841
25 | YES | 751.038 | 34.2497 | 0.00026264 | -0.0014458 | -0.0028391 | -0.0013933
25 | NO | 5.7594 | 32.3796 | 0.00039235 | -0.0030136 | -0.0025072 | 0.00050643
26 | NO | 5.1658 | 32.3928 | 0.00027206 | -0.0021133 | -0.0025069 | -0.00039357
27 | YES | 3998.99 | 34.6752 | 0.00033017 | -0.0038376 | -0.0042857 | -0.00044813
27 | YES | 3500.49 | 34.6666 | 0.00024995 | -0.0044403 | -0.0040637 | 0.00037661
27 | YES | 3000.49 | 34.6508 | 0.00025894 | -0.0035286 | -0.003841 | -0.00031241
27 | YES | 2499.6 | 34.6242 | 0.00034372 | -0.003376 | -0.0036179 | -0.0002419
27 | YES | 1999.58 | 34.5764 | 0.00024728 | -0.0032844 | -0.0033952 | -0.00011076
27 | YES | 1499.76 | 34.4934 | 0.0002439 | -0.0023041 | -0.0031726 | -0.00086851
27 | YES | 1249.49 | 34.436 | 0.00029814 | -0.0022621 | -0.0030611 | -0.000799
27 | YES | 1000.65 | 34.3638 | 0.00037303 | -0.0016861 | -0.0029503 | -0.0012642
27 | NO | 5.3104 | 32.3883 | 0.00031495 | -0.0047302 | -0.002507 | 0.0022232
28 | YES | 2000.87 | 34.574 | 0.00026936 | -0.0036316 | -0.0033958 | 0.00023581
29 | YES | 1249.7 | 34.4584 | 0.00030696 | -0.0036774 | -0.0030612 | 0.00061615
29 | YES | 1250.17 | 34.4584 | 0.00027235 | -0.0079193 | -0.0030614 | 0.0048579
29 | YES | 999.332 | 34.3903 | 0.00025127 | -0.0045624 | -0.0029497 | 0.0016127
29 | YES | 800.879 | 34.2985 | 0.00027876 | -0.0033569 | -0.0028613 | 0.00049562
29 | YES | 600.248 | 34.1981 | 0.00026696 | -0.0047913 | -0.002772 | 0.0020193
29 | NO | 5.2889 | 32.4412 | 0.00039654 | -0.0067024 | -0.002507 | 0.0041955 |  |  |  | -0.0033705 | 0.0015564
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0071332 |  | X | CALC Y
m= | -1.3823e-07 | 1.8466e-07 |  | Std Dev of Resid |  | 0.0014966 |  | 600.248 | -0.00040479
b= | -0.00032182 | 0.00039608 |  | Resid Sum of Squares |  | 0.00017471 |  | 3998.99 | -0.00087461
# of Values Extracted = |  |  | 121
# of Values Used in Fit = |  |  | 80
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0297 | 31.6948 | 0.0013737 | -0.0041828 | -0.00032251 | 0.0038603 |  | 1250.41 | -0.001194
2 | NO | 5.1496 | 31.7188 | 0.00065426 | -0.0058098 | -0.00032253 | 0.0054873 |  | 1250 | -0.0010033
3 | NO | 5.0349 | 32.4947 | 0.00039457 | -0.0030251 | -0.00032251 | 0.0027025 |  | 1000.85 | 0.00057221
4 | YES | 1250.41 | 34.4802 | 0.00025068 | -0.001194 | -0.00049466 | 0.00069934 |  | 800.079 | 0.00040054
4 | YES | 1250 | 34.4803 | 0.00025609 | -0.0010033 | -0.00049461 | 0.00050866 |  | 2000.48 | -0.0016289
4 | YES | 1000.85 | 34.4106 | 0.00020643 | 0.00057221 | -0.00046017 | -0.0010324 |  | 1250.5 | -0.00069809
4 | YES | 800.079 | 34.3125 | 0.00034369 | 0.00040054 | -0.00043241 | -0.00083296 |  | 1000.02 | 0.00094223
4 | NO | 5.1918 | 32.3635 | 0.00078339 | -0.0019646 | -0.00032253 | 0.001642 |  | 800.314 | -0.003315
5 | NO | 5.3204 | 32.3401 | 0.00025866 | -0.0018425 | -0.00032255 | 0.0015199 |  | 3000.46 | -0.0013733
6 | NO | 5.8051 | 32.2808 | 0.0028966 | -0.0058899 | -0.00032262 | 0.0055673 |  | 3000.54 | 0.0017281
7 | NO | 5.1369 | 32.304 | 0.00061638 | -0.0035172 | -0.00032253 | 0.0031946 |  | 2501.14 | -0.0018463
8 | YES | 2000.48 | 34.5865 | 0.00027599 | -0.0016289 | -0.00059835 | 0.0010305 | 34.5881 | 2000.77 | 0.0031929
8 | NO | 1500.42 | 34.5063 | 0.00020375 | -0.0090103 | -0.00052922 | 0.0084811 | 34.5153 | 1499.82 | -0.000103
8 | YES | 1250.5 | 34.4549 | 0.00024428 | -0.00069809 | -0.00049468 | 0.00020341 | 34.4556 | 1249.99 | -0.0014305
8 | YES | 1000.02 | 34.3679 | 0.00027498 | 0.00094223 | -0.00046005 | -0.0014023 | 34.367 | 1000.55 | -0.0013351
8 | YES | 800.314 | 34.2725 | 0.00025056 | -0.003315 | -0.00043245 | 0.0028825 | 34.2758 | 3358.74 | -0.0006752
8 | NO | 5.8206 | 32.3411 | 0.001154 | -0.0021172 | -0.00032262 | 0.0017945 | 32.3432 | 2999.58 | 0.00058746
9 | NO | 6.0775 | 32.3162 | 0.0014399 | -0.0018692 | -0.00032266 | 0.0015465 | 32.3181 | 2499.8 | 0.0012131
10 | NO | 5.3679 | 32.3374 | 0.0003025 | -0.00093842 | -0.00032256 | 0.00061586 | 32.3383 | 2250.6 | -0.00041962
11 | NO | 5.6232 | 32.3421 | 0.00023535 | -0.0017471 | -0.00032259 | 0.0014245 | 32.3438 | 2000.28 | -6.8665e-05
12 | YES | 3000.46 | 34.6464 | 0.00026027 | -0.0013733 | -0.00073658 | 0.00063671 | 34.6478 | 1750.13 | -0.00013733
12 | YES | 3000.54 | 34.6464 | 0.00024911 | 0.0017281 | -0.00073659 | -0.0024646 | 34.6447 | 1500.43 | 0.001133
12 | YES | 2501.14 | 34.6254 | 0.00025605 | -0.0018463 | -0.00066755 | 0.0011788 | 34.6272 | 1250.73 | 0.00017166
12 | YES | 2000.77 | 34.581 | 0.00022098 | 0.0031929 | -0.00059839 | -0.0037913 | 34.5778 | 999.077 | 0.0012436
12 | YES | 1499.82 | 34.4923 | 0.00020459 | -0.000103 | -0.00052914 | -0.00042614 | 34.4924 | 801.614 | 0.0022888
12 | YES | 1249.99 | 34.438 | 0.00026543 | -0.0014305 | -0.00049461 | 0.0009359 | 34.4394 | 600.57 | -6.1035e-05
12 | YES | 1000.55 | 34.3451 | 0.00024519 | -0.0013351 | -0.00046013 | 0.00087501 | 34.3464 | 3499.45 | -0.00092697
12 | NO | 4.9421 | 32.3776 | 0.00020812 | -0.001648 | -0.0003225 | 0.0013254 | 32.3792 | 3499.55 | 0.00038147
13 | NO | 4.7831 | 32.3708 | 0.00021394 | -0.00059891 | -0.00032248 | 0.00027643 |  | 2999.35 | 0.0025787
14 | YES | 3358.74 | 34.656 | 0.00025593 | -0.0006752 | -0.0007861 | -0.0001109 |  | 2500.63 | 0.0011482
14 | YES | 2999.58 | 34.6486 | 0.00028413 | 0.00058746 | -0.00073646 | -0.0013239 |  | 2000.29 | 6.8665e-05
14 | YES | 2499.8 | 34.6262 | 0.00027956 | 0.0012131 | -0.00066737 | -0.0018804 |  | 1500.98 | 0.00094604
14 | YES | 2250.6 | 34.6082 | 0.00026843 | -0.00041962 | -0.00063292 | -0.00021331 |  | 1249.94 | 0.00021362
14 | YES | 2000.28 | 34.5816 | 0.00025074 | -6.8665e-05 | -0.00059832 | -0.00052966 |  | 1001.17 | 0.0014076
14 | YES | 1750.13 | 34.5422 | 0.00022994 | -0.00013733 | -0.00056374 | -0.00042641 |  | 3499.91 | -0.0016823
14 | YES | 1500.43 | 34.4924 | 0.00033467 | 0.001133 | -0.00052922 | -0.0016622 |  | 3000.99 | -0.0015488
14 | YES | 1250.73 | 34.4269 | 0.00033912 | 0.00017166 | -0.00049471 | -0.00066637 |  | 2500.63 | -0.0006218
14 | YES | 999.077 | 34.3558 | 0.00036286 | 0.0012436 | -0.00045992 | -0.0017035 |  | 2000.92 | -0.00086594
14 | YES | 801.614 | 34.2904 | 0.00029701 | 0.0022888 | -0.00043263 | -0.0027214 |  | 1749.7 | -0.00035095
14 | YES | 600.57 | 34.1519 | 0.00026468 | -6.1035e-05 | -0.00040483 | -0.0003438 |  | 1500.53 | 0.00037003
14 | NO | 400.23 | 34.0018 | 0.00033996 | 0.00078583 | -0.00037714 | -0.001163 |  | 3499.86 | -0.0013542
14 | NO | 301.242 | 33.942 | 0.00023702 | 0.00059128 | -0.00036346 | -0.00095474 |  | 3499.6 | -0.00091553
14 | NO | 251.135 | 33.9032 | 0.00023285 | 0.00020981 | -0.00035653 | -0.00056634 |  | 3000.37 | -0.0014343
14 | NO | 200.466 | 33.8921 | 0.00024237 | -0.0010643 | -0.00034953 | 0.00071477 |  | 2501.2 | -0.0052643
14 | NO | 174.963 | 33.8659 | 0.00072397 | 0.0019989 | -0.000346 | -0.0023449 |  | 1999.54 | -0.0013428
14 | NO | 149.842 | 33.7781 | 0.0020758 | -0.014324 | -0.00034253 | 0.013982 |  | 1501.47 | -0.00069427
14 | NO | 125.134 | 33.4507 | 0.00047427 | -0.0073929 | -0.00033911 | 0.0070538 |  | 1250.79 | -0.0012321
14 | NO | 99.9276 | 32.7619 | 0.0018811 | -0.0030518 | -0.00033563 | 0.0027161 |  | 1000.01 | -0.0015984
14 | NO | 75.5711 | 32.3871 | 0.00029715 | -0.0018234 | -0.00033226 | 0.0014912 |  | 2000.27 | 0.00015259
14 | NO | 50.4283 | 32.3924 | 0.0012595 | -0.0013008 | -0.00032879 | 0.00097202 |  | 1501.83 | -0.0011291
14 | NO | 25.2375 | 32.3752 | 0.00018616 | -0.0021133 | -0.00032531 | 0.001788 |  | 1001.35 | 0.00050354
14 | NO | 10.6401 | 32.3753 | 0.00028081 | -0.0018158 | -0.00032329 | 0.0014925 |  | 750.323 | 2.2888e-05
14 | NO | 5.2922 | 32.3753 | 0.0002357 | 0.00033951 | -0.00032255 | -0.00066206 |  | 3499.87 | -0.0026207
15 | NO | 5.0809 | 32.3798 | 0.00027433 | -0.0051727 | -0.00032252 | 0.0048502 |  | 2999.49 | -0.0013924
16 | YES | 3499.45 | 34.6647 | 0.0002533 | -0.00092697 | -0.00080555 | 0.00012142 |  | 2499.87 | 0.00014496
16 | YES | 3499.55 | 34.6647 | 0.00026249 | 0.00038147 | -0.00080557 | -0.001187 |  | 2000.25 | -0.00051498
16 | YES | 2999.35 | 34.6523 | 0.00032487 | 0.0025787 | -0.00073642 | -0.0033152 |  | 1499.8 | -0.0036201
16 | YES | 2500.63 | 34.6289 | 0.00021954 | 0.0011482 | -0.00066748 | -0.0018157 |  | 3500.3 | -0.0005455
16 | YES | 2000.29 | 34.5853 | 0.00025078 | 6.8665e-05 | -0.00059832 | -0.00066699 |  | 3001.15 | -8.3923e-05
16 | YES | 1500.98 | 34.4996 | 0.00033063 | 0.00094604 | -0.0005293 | -0.0014753 |  | 2500.14 | -0.0011673
16 | YES | 1249.94 | 34.425 | 0.00023129 | 0.00021362 | -0.0004946 | -0.00070822 |  | 2000.4 | -0.0045891
16 | YES | 1001.17 | 34.3493 | 0.00023184 | 0.0014076 | -0.00046021 | -0.0018678 |  | 1500.64 | -0.0041122
16 | NO | 5.8015 | 32.3514 | 0.00027816 | -0.0019302 | -0.00032262 | 0.0016076 |  | 2000.73 | -0.00053787
17 | NO | 5.5454 | 32.3507 | 0.0002278 | -0.0012436 | -0.00032258 | 0.00092101 |  | 1500.6 | -0.00030518
18 | YES | 3499.91 | 34.6679 | 0.00024427 | -0.0016823 | -0.00080562 | 0.00087666 |  | 1000.52 | -0.00034714
18 | YES | 3000.99 | 34.6533 | 0.00023442 | -0.0015488 | -0.00073665 | 0.00081212 |  | 751.038 | 0.00050736
18 | YES | 2500.63 | 34.6272 | 0.00024941 | -0.0006218 | -0.00066748 | -4.5688e-05 |  | 3998.99 | -0.00060654
18 | YES | 2000.92 | 34.5847 | 0.00025634 | -0.00086594 | -0.00059841 | 0.00026753 |  | 3500.49 | -0.001255
18 | YES | 1749.7 | 34.5487 | 0.00029795 | -0.00035095 | -0.00056368 | -0.00021273 |  | 3000.49 | -0.00073624
18 | YES | 1500.53 | 34.4992 | 0.00026485 | 0.00037003 | -0.00052924 | -0.00089926 |  | 2499.6 | -0.00079727
18 | NO | 5.4677 | 32.3233 | 0.0044673 | -0.0029564 | -0.00032257 | 0.0026338 |  | 1999.58 | -0.00062942
19 | NO | 5.3677 | 32.3081 | 0.0006651 | -0.003334 | -0.00032256 | 0.0030115 |  | 1499.76 | 7.6294e-06
20 | YES | 3499.86 | 34.669 | 0.00023729 | -0.0013542 | -0.00080561 | 0.00054861 |  | 1249.49 | 4.5776e-05
20 | YES | 3499.6 | 34.6691 | 0.00025415 | -0.00091553 | -0.00080557 | 0.00010995 |  | 1000.65 | 0.00072098
20 | YES | 3000.37 | 34.6548 | 0.0001974 | -0.0014343 | -0.00073656 | 0.00069777 |  | 2000.87 | -0.00086212
20 | YES | 2501.2 | 34.6264 | 0.00034602 | -0.0052643 | -0.00066756 | 0.0045967 |  | 1249.7 | -0.00041199
20 | YES | 1999.54 | 34.5835 | 0.00026733 | -0.0013428 | -0.00059822 | 0.00074455 |  | 1250.17 | -0.0044441
20 | YES | 1501.47 | 34.4975 | 0.00036747 | -0.00069427 | -0.00052937 | 0.00016491 |  | 999.332 | -0.0015564
20 | YES | 1250.79 | 34.437 | 0.00027954 | -0.0012321 | -0.00049472 | 0.00073743 |  | 800.879 | -0.00050736
20 | YES | 1000.01 | 34.361 | 0.0003224 | -0.0015984 | -0.00046005 | 0.0011383 |  | 600.248 | -0.0020447
20 | NO | 5.2111 | 32.3012 | 0.00031322 | -0.010227 | -0.00032254 | 0.0099047
21 | YES | 2000.27 | 34.5795 | 0.00021888 | 0.00015259 | -0.00059832 | -0.00075091
21 | YES | 1501.83 | 34.4845 | 0.00030127 | -0.0011291 | -0.00052942 | 0.00059973
21 | YES | 1001.35 | 34.3543 | 0.00054896 | 0.00050354 | -0.00046024 | -0.00096378
21 | YES | 750.323 | 34.2412 | 0.00027669 | 2.2888e-05 | -0.00042554 | -0.00044842
21 | NO | 4.6249 | 32.2832 | 0.00021424 | -0.0038719 | -0.00032246 | 0.0035495
22 | YES | 3499.87 | 34.6687 | 0.00029337 | -0.0026207 | -0.00080561 | 0.0018151
22 | YES | 2999.49 | 34.6527 | 0.00024881 | -0.0013924 | -0.00073644 | 0.00065592
22 | YES | 2499.87 | 34.6264 | 0.00022656 | 0.00014496 | -0.00066738 | -0.00081234
22 | YES | 2000.25 | 34.5811 | 0.00022591 | -0.00051498 | -0.00059832 | -8.3332e-05
22 | YES | 1499.8 | 34.4995 | 0.00022027 | -0.0036201 | -0.00052914 | 0.003091
22 | NO | 5.4018 | 32.3251 | 0.0016992 | -0.0020256 | -0.00032256 | 0.001703
23 | NO | 5.3733 | 32.3441 | 0.00025929 | -0.0020256 | -0.00032256 | 0.001703
24 | YES | 3500.3 | 34.6676 | 0.00023583 | -0.0005455 | -0.00080567 | -0.00026017
24 | YES | 3001.15 | 34.6507 | 0.00024853 | -8.3923e-05 | -0.00073667 | -0.00065275
24 | YES | 2500.14 | 34.6208 | 0.00038088 | -0.0011673 | -0.00066742 | 0.00049988
24 | YES | 2000.4 | 34.5722 | 0.00023122 | -0.0045891 | -0.00059834 | 0.0039907
24 | YES | 1500.64 | 34.491 | 0.00035254 | -0.0041122 | -0.00052925 | 0.003583
24 | NO | 5.026 | 32.3814 | 0.00020019 | -0.0066948 | -0.00032251 | 0.0063723
25 | YES | 2000.73 | 34.5757 | 0.00025891 | -0.00053787 | -0.00059838 | -6.051e-05
25 | YES | 1500.6 | 34.4912 | 0.00026973 | -0.00030518 | -0.00052925 | -0.00022407
25 | YES | 1000.52 | 34.359 | 0.00025324 | -0.00034714 | -0.00046012 | -0.00011298
25 | YES | 751.038 | 34.2516 | 0.0002865 | 0.00050736 | -0.00042563 | -0.00093299
25 | NO | 5.7594 | 32.3816 | 0.00043425 | -0.0010452 | -0.00032261 | 0.00072262
26 | NO | 5.1658 | 32.3934 | 0.00024389 | -0.0015259 | -0.00032253 | 0.0012033
27 | YES | 3998.99 | 34.6784 | 0.00027527 | -0.00060654 | -0.00087461 | -0.00026807
27 | YES | 3500.49 | 34.6697 | 0.00027476 | -0.001255 | -0.0008057 | 0.00044934
27 | YES | 3000.49 | 34.6536 | 0.00023263 | -0.00073624 | -0.00073658 | -3.4408e-07
27 | YES | 2499.6 | 34.6268 | 0.00030993 | -0.00079727 | -0.00066734 | 0.00012993
27 | YES | 1999.58 | 34.5791 | 0.00029423 | -0.00062942 | -0.00059822 | 3.1202e-05
27 | YES | 1499.76 | 34.4957 | 0.00033148 | 7.6294e-06 | -0.00052913 | -0.00053676
27 | YES | 1249.49 | 34.4383 | 0.00037011 | 4.5776e-05 | -0.00049454 | -0.00054031
27 | YES | 1000.65 | 34.3662 | 0.0003534 | 0.00072098 | -0.00046014 | -0.0011811
27 | NO | 5.3104 | 32.3902 | 0.00031522 | -0.0028114 | -0.00032255 | 0.0024889
28 | YES | 2000.87 | 34.5767 | 0.00026378 | -0.00086212 | -0.0005984 | 0.00026372
29 | YES | 1249.7 | 34.4617 | 0.0003082 | -0.00041199 | -0.00049457 | -8.2579e-05
29 | YES | 1250.17 | 34.4619 | 0.0002273 | -0.0044441 | -0.00049463 | 0.0039495
29 | YES | 999.332 | 34.3933 | 0.00023115 | -0.0015564 | -0.00045996 | 0.0010964
29 | YES | 800.879 | 34.3014 | 0.00023516 | -0.00050736 | -0.00043252 | 7.4831e-05
29 | YES | 600.248 | 34.2009 | 0.00025815 | -0.0020447 | -0.00040479 | 0.0016399
29 | NO | 5.2889 | 32.4436 | 0.0003151 | -0.0043373 | -0.00032255 | 0.0040148 |  |  |  | -0.00059056 | 0.0014924
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.063766 |  | X | CALC Y
m= | -5.9085e-05 | 2.5635e-05 |  | Std Dev of Resid |  | 0.0015156 |  | 4 | -0.0025101
b= | -0.0022738 | 0.00050511 |  | Resid Sum of Squares |  | 0.00017918 |  | 29 | -0.0039872
# of Values Extracted = |  |  | 121
# of Values Used in Fit = |  |  | 80
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0297 | 31.6927 | 0.00094052 | -0.0062809 | -0.0023328 | 0.0039481 |  | 4 | -0.0036583
2 | NO | 5.1496 | 31.7169 | 0.00042182 | -0.0077171 | -0.0023919 | 0.0053252 |  | 4 | -0.0035019
3 | NO | 5.0349 | 32.4938 | 0.00039708 | -0.0039482 | -0.002451 | 0.0014972 |  | 4 | -0.0016098
4 | YES | 1250.41 | 34.4777 | 0.00025257 | -0.0036583 | -0.0025101 | 0.0011482 |  | 4 | -0.0015488
4 | YES | 1250 | 34.4778 | 0.0002617 | -0.0035019 | -0.0025101 | 0.00099179 |  | 8 | -0.0050163
4 | YES | 1000.85 | 34.4084 | 0.0002306 | -0.0016098 | -0.0025101 | -0.0009003 |  | 8 | -0.0033188
4 | YES | 800.079 | 34.3106 | 0.0004017 | -0.0015488 | -0.0025101 | -0.00096133 |  | 8 | -0.0015907
4 | NO | 5.1918 | 32.3621 | 0.00056541 | -0.0033608 | -0.0025101 | 0.00085065 |  | 8 | -0.0055924
5 | NO | 5.3204 | 32.3382 | 0.00035 | -0.0036926 | -0.0025692 | 0.0011234 |  | 12 | -0.0047379
6 | NO | 5.8051 | 32.2771 | 0.0027959 | -0.0095749 | -0.0026283 | 0.0069466 |  | 12 | -0.0016365
7 | NO | 5.1369 | 32.3009 | 0.0011912 | -0.0065689 | -0.0026874 | 0.0038816 |  | 12 | -0.0049553
8 | YES | 2000.48 | 34.5831 | 0.00026923 | -0.0050163 | -0.0027464 | 0.0022699 |  | 12 | 0.00021744
8 | NO | 1500.42 | 34.5036 | 0.00029078 | -0.011723 | -0.0027464 | 0.0089762 |  | 12 | -0.0030403
8 | YES | 1250.5 | 34.4523 | 0.00028453 | -0.0033188 | -0.0027464 | 0.00057235 |  | 12 | -0.0040398
8 | YES | 1000.02 | 34.3654 | 0.00029824 | -0.0015907 | -0.0027464 | -0.0011557 |  | 12 | -0.0039406
8 | YES | 800.314 | 34.2702 | 0.00031399 | -0.0055924 | -0.0027464 | 0.0028459 |  | 14 | -0.0040207
8 | NO | 5.8206 | 32.3382 | 0.0011345 | -0.0049667 | -0.0027464 | 0.0022203 |  | 14 | -0.0026894
9 | NO | 6.0775 | 32.315 | 0.0010035 | -0.0031319 | -0.0028055 | 0.00032634 |  | 14 | -0.0018005
10 | NO | 5.3679 | 32.336 | 0.00033891 | -0.0023346 | -0.0028646 | -0.00053002 |  | 14 | -0.0032844
11 | NO | 5.6232 | 32.3408 | 0.00029712 | -0.0030212 | -0.0029237 | 9.7545e-05 |  | 14 | -0.0029259
12 | YES | 3000.46 | 34.6431 | 0.00028536 | -0.0047379 | -0.0029828 | 0.0017551 |  | 14 | -0.0027428
12 | YES | 3000.54 | 34.6431 | 0.00028818 | -0.0016365 | -0.0029828 | -0.0013463 |  | 14 | -0.0015068
12 | YES | 2501.14 | 34.6222 | 0.00027372 | -0.0049553 | -0.0029828 | 0.0019725 |  | 14 | -0.0023079
12 | YES | 2000.77 | 34.578 | 0.00029896 | 0.00021744 | -0.0029828 | -0.0032002 |  | 14 | -0.001236
12 | YES | 1499.82 | 34.4894 | 0.00030034 | -0.0030403 | -0.0029828 | 5.753e-05 |  | 14 | 1.5259e-05
12 | YES | 1249.99 | 34.4354 | 0.00025783 | -0.0040398 | -0.0029828 | 0.001057 |  | 14 | -0.0025101
12 | YES | 1000.55 | 34.3425 | 0.00027784 | -0.0039406 | -0.0029828 | 0.0009578 |  | 16 | -0.0042267
12 | NO | 4.9421 | 32.3748 | 0.00028591 | -0.0043907 | -0.0029828 | 0.0014079 |  | 16 | -0.0029602
13 | NO | 4.7831 | 32.3696 | 0.00027983 | -0.0018349 | -0.0030419 | -0.001207 |  | 16 | -0.0008049
14 | YES | 3358.74 | 34.6527 | 0.00032084 | -0.0040207 | -0.0031009 | 0.00091974 |  | 16 | -0.0019226
14 | YES | 2999.58 | 34.6453 | 0.00024322 | -0.0026894 | -0.0031009 | -0.00041159 |  | 16 | -0.0027924
14 | YES | 2499.8 | 34.6232 | 0.00027209 | -0.0018005 | -0.0031009 | -0.0013004 |  | 16 | -0.0016174
14 | YES | 2250.6 | 34.6053 | 0.00024067 | -0.0032844 | -0.0031009 | 0.0001835 |  | 16 | -0.002346
14 | YES | 2000.28 | 34.5788 | 0.00024873 | -0.0029259 | -0.0031009 | -0.00017508 |  | 16 | -0.001091
14 | YES | 1750.13 | 34.5396 | 0.00028167 | -0.0027428 | -0.0031009 | -0.00035818 |  | 18 | -0.0054665
14 | YES | 1500.43 | 34.4898 | 0.00031395 | -0.0015068 | -0.0031009 | -0.0015941 |  | 18 | -0.0050812
14 | YES | 1250.73 | 34.4244 | 0.00026095 | -0.0023079 | -0.0031009 | -0.00079306 |  | 18 | -0.0039673
14 | YES | 999.077 | 34.3534 | 0.00029423 | -0.001236 | -0.0031009 | -0.001865 |  | 18 | -0.0040436
14 | YES | 801.614 | 34.2881 | 0.00039595 | 1.5259e-05 | -0.0031009 | -0.0031162 |  | 18 | -0.003582
14 | YES | 600.57 | 34.1495 | 0.00029669 | -0.0025101 | -0.0031009 | -0.00059088 |  | 18 | -0.0025787
14 | NO | 400.23 | 33.9992 | 0.00027122 | -0.0018082 | -0.0031009 | -0.0012928 |  | 20 | -0.0044098
14 | NO | 301.242 | 33.9395 | 0.00026796 | -0.0018692 | -0.0031009 | -0.0012317 |  | 20 | -0.0040131
14 | NO | 251.135 | 33.9008 | 0.00028006 | -0.0022163 | -0.0031009 | -0.00088461 |  | 20 | -0.0042953
14 | NO | 200.466 | 33.8896 | 0.00024651 | -0.0036011 | -0.0031009 | 0.00050012 |  | 20 | -0.007946
14 | NO | 174.963 | 33.8632 | 0.00074556 | -0.00069046 | -0.0031009 | -0.0024105 |  | 20 | -0.0040359
14 | NO | 149.842 | 33.7778 | 0.0031998 | -0.014606 | -0.0031009 | 0.011506 |  | 20 | -0.0034943
14 | NO | 125.134 | 33.448 | 0.00050135 | -0.01009 | -0.0031009 | 0.006989 |  | 20 | -0.003437
14 | NO | 99.9276 | 32.7602 | 0.0018695 | -0.0048065 | -0.0031009 | 0.0017056 |  | 20 | -0.0039902
14 | NO | 75.5711 | 32.3842 | 0.00039235 | -0.0047226 | -0.0031009 | 0.0016217 |  | 21 | -0.0028801
14 | NO | 50.4283 | 32.3909 | 0.001496 | -0.0027542 | -0.0031009 | -0.00034674 |  | 21 | -0.0036545
14 | NO | 25.2375 | 32.3732 | 0.00027495 | -0.0041237 | -0.0031009 | 0.0010227 |  | 21 | -0.0018463
14 | NO | 10.6401 | 32.3731 | 0.00030276 | -0.0040016 | -0.0031009 | 0.00090067 |  | 21 | -0.0026436
14 | NO | 5.2922 | 32.3729 | 0.00027059 | -0.0020523 | -0.0031009 | -0.0010486 |  | 22 | -0.005806
15 | NO | 5.0809 | 32.3787 | 0.00034883 | -0.0063477 | -0.00316 | 0.0031876 |  | 22 | -0.0043144
16 | YES | 3499.45 | 34.6614 | 0.0003025 | -0.0042267 | -0.0032191 | 0.0010076 |  | 22 | -0.0025673
16 | YES | 3499.55 | 34.6613 | 0.00028716 | -0.0029602 | -0.0032191 | -0.00025891 |  | 22 | -0.0030632
16 | YES | 2999.35 | 34.6489 | 0.00021892 | -0.0008049 | -0.0032191 | -0.0024142 |  | 22 | -0.0058594
16 | YES | 2500.63 | 34.6259 | 0.00026318 | -0.0019226 | -0.0032191 | -0.0012965 |  | 24 | -0.0035477
16 | YES | 2000.29 | 34.5824 | 0.00027145 | -0.0027924 | -0.0032191 | -0.00042676 |  | 24 | -0.0028877
16 | YES | 1500.98 | 34.4971 | 0.00028973 | -0.0016174 | -0.0032191 | -0.0016017 |  | 24 | -0.0038261
16 | YES | 1249.94 | 34.4225 | 0.00026383 | -0.002346 | -0.0032191 | -0.00087308 |  | 24 | -0.0069847
16 | YES | 1001.17 | 34.3468 | 0.00027695 | -0.001091 | -0.0032191 | -0.0021281 |  | 24 | -0.00634
16 | NO | 5.8015 | 32.3482 | 0.004966 | -0.0051384 | -0.0032191 | 0.0019193 |  | 25 | -0.0032272
17 | NO | 5.5454 | 32.3492 | 0.00026212 | -0.0026856 | -0.0032782 | -0.00059265 |  | 25 | -0.0026169
18 | YES | 3499.91 | 34.6641 | 0.00029578 | -0.0054665 | -0.0033373 | 0.0021292 |  | 25 | -0.0023918
18 | YES | 3000.99 | 34.6497 | 0.00027862 | -0.0050812 | -0.0033373 | 0.0017439 |  | 25 | -0.0014458
18 | YES | 2500.63 | 34.6238 | 0.00029342 | -0.0039673 | -0.0033373 | 0.00063 |  | 27 | -0.0038376
18 | YES | 2000.92 | 34.5816 | 0.00038399 | -0.0040436 | -0.0033373 | 0.00070629 |  | 27 | -0.0044403
18 | YES | 1749.7 | 34.5455 | 0.00028705 | -0.003582 | -0.0033373 | 0.00024471 |  | 27 | -0.0035286
18 | YES | 1500.53 | 34.4962 | 0.0002533 | -0.0025787 | -0.0033373 | -0.00075855 |  | 27 | -0.003376
18 | NO | 5.4677 | 32.3212 | 0.001761 | -0.0050774 | -0.0033373 | 0.0017401 |  | 27 | -0.0032844
19 | NO | 5.3677 | 32.3073 | 0.00069931 | -0.0040817 | -0.0033964 | 0.00068536 |  | 27 | -0.0023041
20 | YES | 3499.86 | 34.666 | 0.00023188 | -0.0044098 | -0.0034555 | 0.00095433 |  | 27 | -0.0022621
20 | YES | 3499.6 | 34.666 | 0.00026062 | -0.0040131 | -0.0034555 | 0.0005576 |  | 27 | -0.0016861
20 | YES | 3000.37 | 34.6519 | 0.00028201 | -0.0042953 | -0.0034555 | 0.00083989 |  | 28 | -0.0036316
20 | YES | 2501.2 | 34.6238 | 0.00027767 | -0.007946 | -0.0034555 | 0.0044906 |  | 29 | -0.0036774
20 | YES | 1999.54 | 34.5808 | 0.00024127 | -0.0040359 | -0.0034555 | 0.00058049 |  | 29 | -0.0079193
20 | YES | 1501.47 | 34.4947 | 0.00029801 | -0.0034943 | -0.0034555 | 3.8802e-05 |  | 29 | -0.0045624
20 | YES | 1250.79 | 34.4348 | 0.00021871 | -0.003437 | -0.0034555 | -1.8418e-05 |  | 29 | -0.0033569
20 | YES | 1000.01 | 34.3586 | 0.00032184 | -0.0039902 | -0.0034555 | 0.00053471 |  | 29 | -0.0047913
20 | NO | 5.2111 | 32.2989 | 0.00029619 | -0.01252 | -0.0034555 | 0.0090643
21 | YES | 2000.27 | 34.5764 | 0.00024191 | -0.0028801 | -0.0035145 | -0.00063444
21 | YES | 1501.83 | 34.4819 | 0.00026158 | -0.0036545 | -0.0035145 | 0.00013994
21 | YES | 1001.35 | 34.352 | 0.00031468 | -0.0018463 | -0.0035145 | -0.0016682
21 | YES | 750.323 | 34.2386 | 0.00031223 | -0.0026436 | -0.0035145 | -0.00087095
21 | NO | 4.6249 | 32.281 | 0.00038705 | -0.006134 | -0.0035145 | 0.0026195
22 | YES | 3499.87 | 34.6655 | 0.00029465 | -0.005806 | -0.0035736 | 0.0022323
22 | YES | 2999.49 | 34.6498 | 0.00026968 | -0.0043144 | -0.0035736 | 0.00074079
22 | YES | 2499.87 | 34.6237 | 0.00029355 | -0.0025673 | -0.0035736 | -0.0010063
22 | YES | 2000.25 | 34.5785 | 0.00024768 | -0.0030632 | -0.0035736 | -0.00051043
22 | YES | 1499.8 | 34.4972 | 0.00029403 | -0.0058594 | -0.0035736 | 0.0022858
22 | NO | 5.4018 | 32.3232 | 0.0007402 | -0.0039139 | -0.0035736 | 0.00034025
23 | NO | 5.3733 | 32.343 | 0.00032632 | -0.0031052 | -0.0036327 | -0.00052755
24 | YES | 3500.3 | 34.6646 | 0.00028581 | -0.0035477 | -0.0036918 | -0.00014413
24 | YES | 3001.15 | 34.6479 | 0.00027004 | -0.0028877 | -0.0036918 | -0.00080407
24 | YES | 2500.14 | 34.6182 | 0.00034168 | -0.0038261 | -0.0036918 | 0.00013434
24 | YES | 2000.4 | 34.5698 | 0.00027075 | -0.0069847 | -0.0036918 | 0.0032929
24 | YES | 1500.64 | 34.4888 | 0.00027273 | -0.00634 | -0.0036918 | 0.0026482
24 | NO | 5.026 | 32.3794 | 0.00026181 | -0.0086937 | -0.0036918 | 0.0050019
25 | YES | 2000.73 | 34.573 | 0.00030631 | -0.0032272 | -0.0037509 | -0.00052365
25 | YES | 1500.6 | 34.4889 | 0.00029749 | -0.0026169 | -0.0037509 | -0.001134
25 | YES | 1000.52 | 34.3569 | 0.00029477 | -0.0023918 | -0.0037509 | -0.0013591
25 | YES | 751.038 | 34.2497 | 0.00026264 | -0.0014458 | -0.0037509 | -0.0023051
25 | NO | 5.7594 | 32.3796 | 0.00039235 | -0.0030136 | -0.0037509 | -0.00073727
26 | NO | 5.1658 | 32.3928 | 0.00027206 | -0.0021133 | -0.00381 | -0.0016966
27 | YES | 3998.99 | 34.6752 | 0.00033017 | -0.0038376 | -0.0038691 | -3.1461e-05
27 | YES | 3500.49 | 34.6666 | 0.00024995 | -0.0044403 | -0.0038691 | 0.00057126
27 | YES | 3000.49 | 34.6508 | 0.00025894 | -0.0035286 | -0.0038691 | -0.00034046
27 | YES | 2499.6 | 34.6242 | 0.00034372 | -0.003376 | -0.0038691 | -0.00049304
27 | YES | 1999.58 | 34.5764 | 0.00024728 | -0.0032844 | -0.0038691 | -0.0005846
27 | YES | 1499.76 | 34.4934 | 0.0002439 | -0.0023041 | -0.0038691 | -0.001565
27 | YES | 1249.49 | 34.436 | 0.00029814 | -0.0022621 | -0.0038691 | -0.0016069
27 | YES | 1000.65 | 34.3638 | 0.00037303 | -0.0016861 | -0.0038691 | -0.002183
27 | NO | 5.3104 | 32.3883 | 0.00031495 | -0.0047302 | -0.0038691 | 0.00086117
28 | YES | 2000.87 | 34.574 | 0.00026936 | -0.0036316 | -0.0039281 | -0.00029655
29 | YES | 1249.7 | 34.4584 | 0.00030696 | -0.0036774 | -0.0039872 | -0.00030985
29 | YES | 1250.17 | 34.4584 | 0.00027235 | -0.0079193 | -0.0039872 | 0.0039321
29 | YES | 999.332 | 34.3903 | 0.00025127 | -0.0045624 | -0.0039872 | 0.00057516
29 | YES | 800.879 | 34.2985 | 0.00027876 | -0.0033569 | -0.0039872 | -0.00063029
29 | YES | 600.248 | 34.1981 | 0.00026696 | -0.0047913 | -0.0039872 | 0.00080404
29 | NO | 5.2889 | 32.4412 | 0.00039654 | -0.0067024 | -0.0039872 | 0.0027152
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.065227 |  | X | CALC Y
m= | -5.73e-05 | 2.4561e-05 |  | Std Dev of Resid |  | 0.0014522 |  | 4 | 0.00024387
b= | 0.00047307 | 0.00048396 |  | Resid Sum of Squares |  | 0.00016448 |  | 29 | -0.0011886
# of Values Extracted = |  |  | 121
# of Values Used in Fit = |  |  | 80
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0297 | 31.6948 | 0.0013737 | -0.0041828 | 0.00041577 | 0.0045986 |  | 4 | -0.001194
2 | NO | 5.1496 | 31.7188 | 0.00065426 | -0.0058098 | 0.00035847 | 0.0061683 |  | 4 | -0.0010033
3 | NO | 5.0349 | 32.4947 | 0.00039457 | -0.0030251 | 0.00030117 | 0.0033262 |  | 4 | 0.00057221
4 | YES | 1250.41 | 34.4802 | 0.00025068 | -0.001194 | 0.00024387 | 0.0014379 |  | 4 | 0.00040054
4 | YES | 1250 | 34.4803 | 0.00025609 | -0.0010033 | 0.00024387 | 0.0012471 |  | 8 | -0.0016289
4 | YES | 1000.85 | 34.4106 | 0.00020643 | 0.00057221 | 0.00024387 | -0.00032833 |  | 8 | -0.00069809
4 | YES | 800.079 | 34.3125 | 0.00034369 | 0.00040054 | 0.00024387 | -0.00015667 |  | 8 | 0.00094223
4 | NO | 5.1918 | 32.3635 | 0.00078339 | -0.0019646 | 0.00024387 | 0.0022084 |  | 8 | -0.003315
5 | NO | 5.3204 | 32.3401 | 0.00025866 | -0.0018425 | 0.00018657 | 0.0020291 |  | 12 | -0.0013733
6 | NO | 5.8051 | 32.2808 | 0.0028966 | -0.0058899 | 0.00012927 | 0.0060192 |  | 12 | 0.0017281
7 | NO | 5.1369 | 32.304 | 0.00061638 | -0.0035172 | 7.1971e-05 | 0.0035891 |  | 12 | -0.0018463
8 | YES | 2000.48 | 34.5865 | 0.00027599 | -0.0016289 | 1.4671e-05 | 0.0016436 |  | 12 | 0.0031929
8 | NO | 1500.42 | 34.5063 | 0.00020375 | -0.0090103 | 1.4671e-05 | 0.009025 |  | 12 | -0.000103
8 | YES | 1250.5 | 34.4549 | 0.00024428 | -0.00069809 | 1.4671e-05 | 0.00071276 |  | 12 | -0.0014305
8 | YES | 1000.02 | 34.3679 | 0.00027498 | 0.00094223 | 1.4671e-05 | -0.00092756 |  | 12 | -0.0013351
8 | YES | 800.314 | 34.2725 | 0.00025056 | -0.003315 | 1.4671e-05 | 0.0033296 |  | 14 | -0.0006752
8 | NO | 5.8206 | 32.3411 | 0.001154 | -0.0021172 | 1.4671e-05 | 0.0021318 |  | 14 | 0.00058746
9 | NO | 6.0775 | 32.3162 | 0.0014399 | -0.0018692 | -4.2629e-05 | 0.0018266 |  | 14 | 0.0012131
10 | NO | 5.3679 | 32.3374 | 0.0003025 | -0.00093842 | -9.9929e-05 | 0.00083849 |  | 14 | -0.00041962
11 | NO | 5.6232 | 32.3421 | 0.00023535 | -0.0017471 | -0.00015723 | 0.0015899 |  | 14 | -6.8665e-05
12 | YES | 3000.46 | 34.6464 | 0.00026027 | -0.0013733 | -0.00021453 | 0.0011588 |  | 14 | -0.00013733
12 | YES | 3000.54 | 34.6464 | 0.00024911 | 0.0017281 | -0.00021453 | -0.0019426 |  | 14 | 0.001133
12 | YES | 2501.14 | 34.6254 | 0.00025605 | -0.0018463 | -0.00021453 | 0.0016318 |  | 14 | 0.00017166
12 | YES | 2000.77 | 34.581 | 0.00022098 | 0.0031929 | -0.00021453 | -0.0034074 |  | 14 | 0.0012436
12 | YES | 1499.82 | 34.4923 | 0.00020459 | -0.000103 | -0.00021453 | -0.00011153 |  | 14 | 0.0022888
12 | YES | 1249.99 | 34.438 | 0.00026543 | -0.0014305 | -0.00021453 | 0.001216 |  | 14 | -6.1035e-05
12 | YES | 1000.55 | 34.3451 | 0.00024519 | -0.0013351 | -0.00021453 | 0.0011206 |  | 16 | -0.00092697
12 | NO | 4.9421 | 32.3776 | 0.00020812 | -0.001648 | -0.00021453 | 0.0014334 |  | 16 | 0.00038147
13 | NO | 4.7831 | 32.3708 | 0.00021394 | -0.00059891 | -0.00027183 | 0.00032708 |  | 16 | 0.0025787
14 | YES | 3358.74 | 34.656 | 0.00025593 | -0.0006752 | -0.00032913 | 0.00034607 |  | 16 | 0.0011482
14 | YES | 2999.58 | 34.6486 | 0.00028413 | 0.00058746 | -0.00032913 | -0.00091659 |  | 16 | 6.8665e-05
14 | YES | 2499.8 | 34.6262 | 0.00027956 | 0.0012131 | -0.00032913 | -0.0015422 |  | 16 | 0.00094604
14 | YES | 2250.6 | 34.6082 | 0.00026843 | -0.00041962 | -0.00032913 | 9.0487e-05 |  | 16 | 0.00021362
14 | YES | 2000.28 | 34.5816 | 0.00025074 | -6.8665e-05 | -0.00032913 | -0.00026047 |  | 16 | 0.0014076
14 | YES | 1750.13 | 34.5422 | 0.00022994 | -0.00013733 | -0.00032913 | -0.0001918 |  | 18 | -0.0016823
14 | YES | 1500.43 | 34.4924 | 0.00033467 | 0.001133 | -0.00032913 | -0.0014621 |  | 18 | -0.0015488
14 | YES | 1250.73 | 34.4269 | 0.00033912 | 0.00017166 | -0.00032913 | -0.00050079 |  | 18 | -0.0006218
14 | YES | 999.077 | 34.3558 | 0.00036286 | 0.0012436 | -0.00032913 | -0.0015727 |  | 18 | -0.00086594
14 | YES | 801.614 | 34.2904 | 0.00029701 | 0.0022888 | -0.00032913 | -0.002618 |  | 18 | -0.00035095
14 | YES | 600.57 | 34.1519 | 0.00026468 | -6.1035e-05 | -0.00032913 | -0.0002681 |  | 18 | 0.00037003
14 | NO | 400.23 | 34.0018 | 0.00033996 | 0.00078583 | -0.00032913 | -0.001115 |  | 20 | -0.0013542
14 | NO | 301.242 | 33.942 | 0.00023702 | 0.00059128 | -0.00032913 | -0.00092041 |  | 20 | -0.00091553
14 | NO | 251.135 | 33.9032 | 0.00023285 | 0.00020981 | -0.00032913 | -0.00053894 |  | 20 | -0.0014343
14 | NO | 200.466 | 33.8921 | 0.00024237 | -0.0010643 | -0.00032913 | 0.00073517 |  | 20 | -0.0052643
14 | NO | 174.963 | 33.8659 | 0.00072397 | 0.0019989 | -0.00032913 | -0.002328 |  | 20 | -0.0013428
14 | NO | 149.842 | 33.7781 | 0.0020758 | -0.014324 | -0.00032913 | 0.013995 |  | 20 | -0.00069427
14 | NO | 125.134 | 33.4507 | 0.00047427 | -0.0073929 | -0.00032913 | 0.0070637 |  | 20 | -0.0012321
14 | NO | 99.9276 | 32.7619 | 0.0018811 | -0.0030518 | -0.00032913 | 0.0027226 |  | 20 | -0.0015984
14 | NO | 75.5711 | 32.3871 | 0.00029715 | -0.0018234 | -0.00032913 | 0.0014943 |  | 21 | 0.00015259
14 | NO | 50.4283 | 32.3924 | 0.0012595 | -0.0013008 | -0.00032913 | 0.00097168 |  | 21 | -0.0011291
14 | NO | 25.2375 | 32.3752 | 0.00018616 | -0.0021133 | -0.00032913 | 0.0017842 |  | 21 | 0.00050354
14 | NO | 10.6401 | 32.3753 | 0.00028081 | -0.0018158 | -0.00032913 | 0.0014867 |  | 21 | 2.2888e-05
14 | NO | 5.2922 | 32.3753 | 0.0002357 | 0.00033951 | -0.00032913 | -0.00066864 |  | 22 | -0.0026207
15 | NO | 5.0809 | 32.3798 | 0.00027433 | -0.0051727 | -0.00038643 | 0.0047863 |  | 22 | -0.0013924
16 | YES | 3499.45 | 34.6647 | 0.0002533 | -0.00092697 | -0.00044373 | 0.00048324 |  | 22 | 0.00014496
16 | YES | 3499.55 | 34.6647 | 0.00026249 | 0.00038147 | -0.00044373 | -0.0008252 |  | 22 | -0.00051498
16 | YES | 2999.35 | 34.6523 | 0.00032487 | 0.0025787 | -0.00044373 | -0.0030225 |  | 22 | -0.0036201
16 | YES | 2500.63 | 34.6289 | 0.00021954 | 0.0011482 | -0.00044373 | -0.001592 |  | 24 | -0.0005455
16 | YES | 2000.29 | 34.5853 | 0.00025078 | 6.8665e-05 | -0.00044373 | -0.0005124 |  | 24 | -8.3923e-05
16 | YES | 1500.98 | 34.4996 | 0.00033063 | 0.00094604 | -0.00044373 | -0.0013898 |  | 24 | -0.0011673
16 | YES | 1249.94 | 34.425 | 0.00023129 | 0.00021362 | -0.00044373 | -0.00065735 |  | 24 | -0.0045891
16 | YES | 1001.17 | 34.3493 | 0.00023184 | 0.0014076 | -0.00044373 | -0.0018514 |  | 24 | -0.0041122
16 | NO | 5.8015 | 32.3514 | 0.00027816 | -0.0019302 | -0.00044373 | 0.0014865 |  | 25 | -0.00053787
17 | NO | 5.5454 | 32.3507 | 0.0002278 | -0.0012436 | -0.00050103 | 0.00074256 |  | 25 | -0.00030518
18 | YES | 3499.91 | 34.6679 | 0.00024427 | -0.0016823 | -0.00055833 | 0.0011239 |  | 25 | -0.00034714
18 | YES | 3000.99 | 34.6533 | 0.00023442 | -0.0015488 | -0.00055833 | 0.00099044 |  | 25 | 0.00050736
18 | YES | 2500.63 | 34.6272 | 0.00024941 | -0.0006218 | -0.00055833 | 6.3464e-05 |  | 27 | -0.00060654
18 | YES | 2000.92 | 34.5847 | 0.00025634 | -0.00086594 | -0.00055833 | 0.0003076 |  | 27 | -0.001255
18 | YES | 1749.7 | 34.5487 | 0.00029795 | -0.00035095 | -0.00055833 | -0.00020738 |  | 27 | -0.00073624
18 | YES | 1500.53 | 34.4992 | 0.00026485 | 0.00037003 | -0.00055833 | -0.00092836 |  | 27 | -0.00079727
18 | NO | 5.4677 | 32.3233 | 0.0044673 | -0.0029564 | -0.00055833 | 0.0023981 |  | 27 | -0.00062942
19 | NO | 5.3677 | 32.3081 | 0.0006651 | -0.003334 | -0.00061563 | 0.0027184 |  | 27 | 7.6294e-06
20 | YES | 3499.86 | 34.669 | 0.00023729 | -0.0013542 | -0.00067293 | 0.00068129 |  | 27 | 4.5776e-05
20 | YES | 3499.6 | 34.6691 | 0.00025415 | -0.00091553 | -0.00067293 | 0.00024259 |  | 27 | 0.00072098
20 | YES | 3000.37 | 34.6548 | 0.0001974 | -0.0014343 | -0.00067293 | 0.0007614 |  | 28 | -0.00086212
20 | YES | 2501.2 | 34.6264 | 0.00034602 | -0.0052643 | -0.00067293 | 0.0045913 |  | 29 | -0.00041199
20 | YES | 1999.54 | 34.5835 | 0.00026733 | -0.0013428 | -0.00067293 | 0.00066984 |  | 29 | -0.0044441
20 | YES | 1501.47 | 34.4975 | 0.00036747 | -0.00069427 | -0.00067293 | 2.1343e-05 |  | 29 | -0.0015564
20 | YES | 1250.79 | 34.437 | 0.00027954 | -0.0012321 | -0.00067293 | 0.00055922 |  | 29 | -0.00050736
20 | YES | 1000.01 | 34.361 | 0.0003224 | -0.0015984 | -0.00067293 | 0.00092543 |  | 29 | -0.0020447
20 | NO | 5.2111 | 32.3012 | 0.00031322 | -0.010227 | -0.00067293 | 0.0095543
21 | YES | 2000.27 | 34.5795 | 0.00021888 | 0.00015259 | -0.00073023 | -0.00088282
21 | YES | 1501.83 | 34.4845 | 0.00030127 | -0.0011291 | -0.00073023 | 0.00039892
21 | YES | 1001.35 | 34.3543 | 0.00054896 | 0.00050354 | -0.00073023 | -0.0012338
21 | YES | 750.323 | 34.2412 | 0.00027669 | 2.2888e-05 | -0.00073023 | -0.00075312
21 | NO | 4.6249 | 32.2832 | 0.00021424 | -0.0038719 | -0.00073023 | 0.0031417
22 | YES | 3499.87 | 34.6687 | 0.00029337 | -0.0026207 | -0.00078753 | 0.0018332
22 | YES | 2999.49 | 34.6527 | 0.00024881 | -0.0013924 | -0.00078753 | 0.00060483
22 | YES | 2499.87 | 34.6264 | 0.00022656 | 0.00014496 | -0.00078753 | -0.00093249
22 | YES | 2000.25 | 34.5811 | 0.00022591 | -0.00051498 | -0.00078753 | -0.00027255
22 | YES | 1499.8 | 34.4995 | 0.00022027 | -0.0036201 | -0.00078753 | 0.0028326
22 | NO | 5.4018 | 32.3251 | 0.0016992 | -0.0020256 | -0.00078753 | 0.0012381
23 | NO | 5.3733 | 32.3441 | 0.00025929 | -0.0020256 | -0.00084483 | 0.0011808
24 | YES | 3500.3 | 34.6676 | 0.00023583 | -0.0005455 | -0.00090213 | -0.00035663
24 | YES | 3001.15 | 34.6507 | 0.00024853 | -8.3923e-05 | -0.00090213 | -0.00081821
24 | YES | 2500.14 | 34.6208 | 0.00038088 | -0.0011673 | -0.00090213 | 0.00026517
24 | YES | 2000.4 | 34.5722 | 0.00023122 | -0.0045891 | -0.00090213 | 0.0036869
24 | YES | 1500.64 | 34.491 | 0.00035254 | -0.0041122 | -0.00090213 | 0.0032101
24 | NO | 5.026 | 32.3814 | 0.00020019 | -0.0066948 | -0.00090213 | 0.0057927
25 | YES | 2000.73 | 34.5757 | 0.00025891 | -0.00053787 | -0.00095943 | -0.00042156
25 | YES | 1500.6 | 34.4912 | 0.00026973 | -0.00030518 | -0.00095943 | -0.00065426
25 | YES | 1000.52 | 34.359 | 0.00025324 | -0.00034714 | -0.00095943 | -0.0006123
25 | YES | 751.038 | 34.2516 | 0.0002865 | 0.00050736 | -0.00095943 | -0.0014668
25 | NO | 5.7594 | 32.3816 | 0.00043425 | -0.0010452 | -0.00095943 | 8.5797e-05
26 | NO | 5.1658 | 32.3934 | 0.00024389 | -0.0015259 | -0.0010167 | 0.00050915
27 | YES | 3998.99 | 34.6784 | 0.00027527 | -0.00060654 | -0.001074 | -0.0004675
27 | YES | 3500.49 | 34.6697 | 0.00027476 | -0.001255 | -0.001074 | 0.00018101
27 | YES | 3000.49 | 34.6536 | 0.00023263 | -0.00073624 | -0.001074 | -0.0003378
27 | YES | 2499.6 | 34.6268 | 0.00030993 | -0.00079727 | -0.001074 | -0.00027676
27 | YES | 1999.58 | 34.5791 | 0.00029423 | -0.00062942 | -0.001074 | -0.00044461
27 | YES | 1499.76 | 34.4957 | 0.00033148 | 7.6294e-06 | -0.001074 | -0.0010817
27 | YES | 1249.49 | 34.4383 | 0.00037011 | 4.5776e-05 | -0.001074 | -0.0011198
27 | YES | 1000.65 | 34.3662 | 0.0003534 | 0.00072098 | -0.001074 | -0.001795
27 | NO | 5.3104 | 32.3902 | 0.00031522 | -0.0028114 | -0.001074 | 0.0017374
28 | YES | 2000.87 | 34.5767 | 0.00026378 | -0.00086212 | -0.0011313 | -0.00026921
29 | YES | 1249.7 | 34.4617 | 0.0003082 | -0.00041199 | -0.0011886 | -0.00077665
29 | YES | 1250.17 | 34.4619 | 0.0002273 | -0.0044441 | -0.0011886 | 0.0032555
29 | YES | 999.332 | 34.3933 | 0.00023115 | -0.0015564 | -0.0011886 | 0.00036777
29 | YES | 800.879 | 34.3014 | 0.00023516 | -0.00050736 | -0.0011886 | -0.00068128
29 | YES | 600.248 | 34.2009 | 0.00025815 | -0.0020447 | -0.0011886 | 0.00085605
29 | NO | 5.2889 | 32.4436 | 0.0003151 | -0.0043373 | -0.0011886 | 0.0031487
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE