## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 19 March 2025 | ### Updated channel names & formats in TOB files. G.G. |
| ### 20 January 2021 | ### Transmissivity:Green was recalibrated to correct an error in the original conversion. S.H. |
| 10 July 2020 | Reversed order of transmissometers in CTD & CHE files. |

## PROCESSING NOTES

Cruise: 2019-014

Agency: Ecosystem Sciences Division

Location: North-East Pacific

Project: Seamounts

Chief Scientist: Norgard T.

Platform: John P. Tully

Date: 16 July 2019 –  29 July 2019

Processed by: Germaine Gatien

Date of Processing: 20 November 2019 – 10 January 2020

Number of original HEX files: 	27 		Number of CTD files: 27

Number of rosette files:	 25			Number of bottle casts processed: 25

Number of original TSG files: 1			Number of processed TSG files: 13

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (#1883DG & 1185DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (#3640) and an altimeter.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully Laptop.

A Guildline model 8400B Autosal serial # 68572 was used to analyze salinity samples.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in excellent order with comments about problems encountered.

The format for the CTD file names was incorrect. 3 digits should be used for event #s and hyphens should separate the different sections of the file names.

There were 2 transmissometers mounted during this cruise. The configuration file used at sea had the parameters reversed from the positions in which they were actually mounted. This was corrected in processing.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

Because this cruise was sampling seamounts, bottom depths obtained from a sounder are not very useful; steep slopes mean the altimetry is a better measure of water depth when the CTD was near the bottom of casts. The bottom depth header entry as entered at sea was replaced by an estimate based on maximum depth sampled plus altimetry at the bottom of the cast. Plots were examined to ensure that the altimetry header entries were reasonable. Even in the presence of very noisy altimetry data the algorithm worked very well at determining reasonable values for altimetry at the bottom.

There were no downcast data for event #22. A CTD file was prepared using upcast data but fewer significant figures are reported for temperature and salinity channels to reflect the lower quality in upcast data due to the CTD travelling through the wake of the rosette.

The primary temperature and salinity channels were chosen for most casts, but for event #18 there was a problem with the primary conductivity so the secondary channels were prepared for archival for the CTD file. The primary data were used for the CHE file.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.30 mL/L from 0 to 300db

±0.16 mL/L from 300db to 800db

±0.06 mL/L from 800db to 1250db

±0.03 mL/L below 1250db

The TSG data were delivered in a single file; this was split into separate files for each date.

Comparisons were made between the IOS TSG data and co-incident CTD casts. There were no loop samples. TSG Intake Temperature was recalibrated by adding 0.057C°. TSG salinity was lower than CTD salinity by a fairly consistent amount for the first 6 files but then started drifting lower. Recalibration was applied by adding 0.075psu, 0.095psu or 0.130psu depending on date.

Temperature and salinity for both TSGs are shown with 3 decimal places to reflect uncertainty in the calibrations.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The formats in the file names were incorrect, but were fixed prior to conversion.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Since their last calibration most of these sensors have been used on 8 previous casts (7 for the DO sensor).

- There were 2 transmissometers mounted on the CTD and it was reported after the cruise that they were mounted in reverse order to that indicated in the configuration file.

- The calibration control files were checked. The only problem was the reversal of the 2 transmissometers as also occurred during 2019-005 and 2019-006; they were mounted in the opposite order to that in the original con file. After that correction the file was saved as 2019-014-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2019-014-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and only 1 cast had significant spiking. CTDEDIT was used to clean a spiky section in the secondary salinity around 800db in cast #29 and the output file was copied to *.BOT.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The file was sorted again on event #/sample #.

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-014-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF 2019-014 OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-014oxy.csv. That file was converted into individual *.OXY files.

At this point note was made of all comments in the DO file that had comments starting with “ALL:”. The rosette sheets were checked to note what other samples were taken from the same bottle so the flags will be applied to all samples from the other csv files that are affected by the DO analysts observation. (A few of these were missed and added at the MRGREO stage.) In some cases the comment only concerned order of bottle firing so no flags were added to other variables.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2019-014 CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2019-014chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2019-014SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 4 to 14 days after collection. The files were simplified and saved as 2019-014sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2019-014*.xlsx. This includes a precision study.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions.

- Event #1 had a ROS file but no sampling, so will not be processed further. (Stn name TEST)

- Event #29, sample 139. An oxygen sample was planned but no sample was found; presume not taken. Flag 9 and comment added.

- Sample numbers got assigned to the wrong Niskins in ADDSAMP – Fixed ADDSAMP and recreated CST, SAM and SAMAVG files and repeated merges.

- Same as for event 41 Fixed ADDSAMP etc.

- Event 47, Niskin #1 fired during downcast –remove from SAMAVG and re-merge.

The data were exported again and all looked fine.

##### Compare

Salinity

Compare was run with pressure as reference channel. The samples all came from below 800db.

There were only 3 significant outliers and they all involved bottle values that had been flagged 4 due to missing nylon liners. The CTD looked low relative to these bottles, likely due to evaporation of samples. Given all bottles came from depths where we do not expect problems from flushing and known problems with the sampling, these differences are large enough to justify rejecting the results. The analyst decided that those values should be replaced with pad values.

- Event 2, sample 1 at 3280db (high by 0.008psu)and sample 5 at 1499db (high by 0.06psu).

- Event 33, sample 161 at 1000db (high by 0.02).

When those values were rejected the primary CTD salinity was lower than bottles by 0.0029psu with a standard deviation of ~0.0009psu. The secondary CTD salinity was higher than bottles by 0.0010psu with a standard deviation of 0.0011psu.

There were many bottles fired at the bottom of casts. Bottom bottles tend to stand out in comparisons because any errors due to incomplete flushing of Niskin bottles have the opposite sign to bottles fired during the upcast. The primary CTD salinity was low by an average of 0.0030psu when bottom bottles were excluded and by 0.0029 if only bottom bottles were included. Similarly the secondary salinity was high by 0.0008psu when excluding bottom bottles and by an average of 0.0014psu based on bottom bottles only. So there is no significant error due to incomplete flushing.

A fit against event # shows no noticeable drift with time.

For full details for the COMPARE run see file 2019-014-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There were 2 major outliers (out of line by >0.5mL/L):

- Event #2 – sample 2 – significant outlier in comparison and bottle value is out of line in profile, looks too high. Given it is out of line in profile and in the fit, a 5 flag was assigned by the analyst.

- Event #2 – sample 6 – significant outlier in comparison and bottle value is out of line in profile – high for just below the OMZ. Given it is out of line in profile and in the fit, a 5 flag was assigned by the analyst.

The following were Chauvenet outliers and out of line by >0.15mL/L

Event #37 – sample 205 – both reps look high but the first is reasonably close to CTD in the fit. The analyst chose to use the value 0.485 and assigned a 2 flag with an explanation for that.

Event #57 – sample 333 – the first rep looks much better in the fit than the second. The analyst chose to use the value 6.282 and assigned a 2 flag with an explanation for that.

A fit that includes most of the near-surface samples was:

CTD DO Corrected = CTD DO * 1.0307 + 0.019 Fit #1

There is more noise at the higher end of the DO range. This is likely due to noisier CTD DO data and the effect of incomplete flushing of Niskin bottles. There is a DO gradient reversal near the surface. The Niskin bottle contents likely contain some water from lower in the water column because they don’t flush completely. Generally in offshore waters flushing is better than in protected areas due to high sea state. However, even in those conditions there can be significant errors in the presence of high vertical DO gradients. That is most often the case in the top 150db for this cruise, an area where DO values are mostly between 4 and 8mL/L. A fit that excluded the top 100db plus outliers based on residuals was:.

CTD DO Corrected = CTD DO * 1.035 + 0.013 Fit #2

This is very close to the fit found for 2019-006:

CTD DO Corrected = CTD DO * 1.035 + 0.014

This suggests that using fit #2 is likely a better choice than Fit #1.

A hysteresis check was done by comparing differences above and below 1000db. There were no obvious problems with deeper values in line with shallower ones in the same DO range.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

All Extracted CHL values were low. The slope of the FL versus CHL plot was ~2.8 as is typical of these fluorometers when CHL is very low.

For details see file 2019-014-fl-chl-comp.xls.

After the analysts made corrections to values, flag and comments in salinity and dissolved oxygen analysis spreadsheets, the merge process was repeated to the MRGCLN2 stage.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-014-ctd.xmlcon. The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. PAR was converted but was actually not mounted on the CTD so should be removed later. The T and C pairs were reasonably close during downcasts with upcasts quite noisy, as usual. Fluorescence and Dissolved Oxygen profiles looked normal. The “Green” transmissometer was generally higher than the “Red” one; the profiles had similar shapes. The altimetry looked noisy at the bottom so careful checking will be necessary.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many SBE DO sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find differences between the pairs of temperature, conductivity and salinity channels. Lines highlighted in grey are from other cruises using the same equipment before and after this one.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2019-005-0051 | 1000 | -0.0010 | +0.00015 | +0.0028 | High, V.Noisy |
| “ | 1900 | -0.0006 | +0.00015 | +0.0025 | “ |
| 2019-005-0159 | 1000 | -0.0009 | +0.00020 | +0.0034 | F.High, Noisy |
| “ | 1900 | -0.0007 | +0.00018 | +0.0030 | “ |
| 2019-006-0032 | 1000 | -0.0007 | +0.00025 | +0.0040 | High, XNoisy |
| “ | 1900 | -0.0004 | +0.00025 | +0.0035 | “ |
| “ | 2900 | -0.0003 | +0.00026 | +0.0035 | “ |
| 2019-006-0057 | 1000 | -0.0007 | +0.00027 | +0.0039 | High, Noisy |
| “ | 1900 | -0.0004 | +0.00026 | +0.0037 | “ |
| “ | 2900 | -0.0005 | +0.00026 | +0.0037 | “ |
| 2019-006-0077 | 1000 | -0.0007 | +0.00021 | +0.0032 | High, XNoisy |
| “ | 1900 | -0.0004 | +0.00021 | +0.0030 | “ |
| “ | 2900 | -0.0006 | +0.00021 | +0.0030 | “ |
| “ | 3900 | -0.0005 | +0.00022 | +0.0032 | “ |
| 2019-014-0002 | 1000 | -0.0008 | +0.00028 | +0.0040 | High, Noisy |
|  | 1900 | -0.0003 | +0.00028 | +0.0035 | “ |
|  | 2500 | -0.0002 | +0.00030 | +0.0048 | “ |
| 2019-014-0033 | 1000 | -0.0009 | +0.00034 | +0.0048 | High, Noisy |
|  | 1900 | -0.0005 | +0.00033 | +0.0044 | “ |
|  | 2500 | -0.0002 | +0.00032 | +0.0043 | “ |
| 2019-014-0049 | 1000 | -0.0009 | +0.00034 | +0.0050 | High, V.Noisy |
|  | 1900 | -0.0003 | +0.00033 | +0.0045 | “ |
|  | 2500 | -0.0004 | +0.00033 | +0.0045 | “ |
| 2019-069-0030 | 1000 | -0.0004 | +0.00041 | +0.0054 | High, Noisy |
| “ | 1900 | -0.0003 | +0.00040 | +0.0052 | “ |
| 2019-069-0051 | 1000 | -0.0009 | +0.00041 | +0.0059 | High, F Noisy |
| “ | 1900 | -0.0004 | +0.00039 | +0.0053 | “ |

The temperature differences are similar to the cruises that preceded and followed it. The conductivity and salinity differences are intermediate between those 2 cruises and suggest a slight drift to larger differences. The data from cruise 2019-069 are preliminary.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run. An error in station name for event #43 was found and corrected in the headers of the IOS, CLN, SAM, SAMAVG, MRG and MRGCLN2 files.

The header check was run. The maximum fluorescence value occurred during event #22 and a plot suggests there might have been a few values that were off-scale; this is not clear, and, at most, affects some records in a 1.5m section at about 32db. No action was taken to remove those points as they are likely ok. The value maximum value recorded was 14.788ug/L. No other problems were noted. There was no CHL sampling for that cast and no downcast data acquired.

Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header. Many of the check values were very high. A second calculation was made using log book entries for water depth and that often reduced the check values but they were often still higher than usual. It is assumed that the problem is because many of the casts were on the steep slopes of seamounts so that the time of the sounder reading has a big impact.

So the bottom depths obtained from a sounder are not very useful. The altimetry is a better measure of water depth when the CTD was near the bottom of casts. The bottom depth header entry as entered at sea was replaced by an estimate based on maximum depth sampled plus altimetry at the bottom of the cast. Plots were examined to ensure that the altimetry header entries were reasonable. Even in the presence of very noisy altimetry data the algorithm worked very well at determining reasonable values for altimetry at the bottom.

File 2019-014-header-merge.csv was prepared to add the estimated water depths to the headers using program Merge CSV File to Headers of the CLN files.

The same step was applied to the MRGCLN2 files with output MRGMRH. There were no cases of bottle files that had not sampled near the bottom.

A note of explanation will be added to the headers.

The following information is available concerning the pressure accuracy:

- During the previous cruise using this CTD there was some suggestion that the pressure could be reading too high by <1m, but there was other evidence that cast doubt on that observation.

- A surface check shows a mean pressure of 2.7db which is fairly low for offshore work on the Tully. So it is unlikely that the pressure is reading significantly low.

- At the end of cast # 22 the CTD was left running for almost 7 minutes near the surface with pumps on. Pressure dropped very briefly to about 0.2db and then very gradually increased to about 0.4db. There was a single spike in transmissivity and primary conductivity to near 0 when pressure first dropped to ~0.2db so the CTD may have briefly left the water, but both channels returned immediately to in-water values and did not spike low again. The secondary conductivity had many spikes to near 0 values. The top of the rosette would have been out of the water and the secondary sensors were likely closer to the surface than the primary. These observations suggest that the surface pressure is within ±0.3db.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts to assess whether the same settings used for 2019-005 and 2019-006 to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space) were appropriate for this cruise and they were. The best setting for both sensors was -0.7records. However, while the secondary data were improved, they did not look particularly good with that setting.

SHIFT was run twice on all SBE911 casts using -0.7 records for the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

DELETE was run on all casts but cast #22 was put through REVERSE and then put through DELETE with extension DELREV on the output.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature and conductivity sensors have been used on 8 cruises since they were last serviced. Two of the cruises did not have useful calibration sampling. During 2018-030 the primary salinity was found to read low by 0.002 and the secondary by 0.001. During 2018-032 there were problems with the salinometer so the comparison was not trusted. During 2019-005 the primary salinity was low by 0.0023psu and the secondary had no significant difference. During 2019-006 the primary was  0.0006psu  while the secondary was low by 0.0034psu.

During 2018-032 DO was recalibrated using a slope of 1.0324 and offset of 0.0096. During 2019-003 and 2019-001 the primary salinity was low by ~0.002 and the secondary was low by <0.001psu. During 2019-005 the primary was low by ~0.0023 and the secondary was very close to the bottles. During 2019-006 the primary was low by 0.0006 and the secondary was low by 0.0034psu.

The DO sensor #3234 was used for 6 of those cruises noted above. The fits had slopes varying from 1.025 to 1.032 and offsets from 0.0096 to 0.0275. During 2019-005 and 2019-006 the fits had slopes of 1.03 and 1.035 respectively with offsets of 0.023 and 0.014mL/L.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data fell within those ranges.

Repeat Casts –There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

As was found during 2019-006 the primary channels looked like the best choice to archive for most casts because there was fewer unstable features in T-S plots and because the salinity comparison with bottles showed less pressure and time dependence. During 2019-006 there were a few casts with poor primary salinity data and there was 1 cast with poor data for this cruise as well. So the secondary channels were chosen for cast #18.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. No further editing was necessary.

##### Recalibration

No pressure correction will be applied. The differences from the comparison with bottles will be used to recalibrate salinity (+0.0029/-0.0010) since flushing errors are likely very small. Downcast differences are slightly larger than 0.0039 – closer to 0.0045, but the differences underway may include some small alignment error.

File 2019-014-recal1.ccf was prepared to recalibrate by adding 0.0029psu to channel Salinity:T0:C0 and subtracting 0.0010psu from channel Salinity :T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.035 + 0.013

This correction was first applied to the SAM and MRGMRH files.

COMPARE was rerun for salinity to ensure that corrections were appropriate and they were. The primary salinity was found to be high by an average of 0.00003psu and the secondary high by 0.00004psu. See file 2019-014-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded using roughly the same points as in the original fit, the CTD DO was low by an average of 0.0019mL/L. This shows that the recalibration was done correctly. See file 2019-014-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When large outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.007mL/L (standard deviation of 0.04mL/L). This is as good a fit as we are likely to achieve.

A display of differences versus pressure were examined; based on that Oxygen:Dissolved:SBE data in the CTD files for this cruise are considered, very roughly, to be:

±0.30 mL/L from 0 to 300db

±0.16 mL/L from 300db to 800db

±0.06 mL/L from 800db to 1250db

±0.03 mL/L below 1250db

No further recalibration was applied. See 2019-014-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #18 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, PAR and Flag.

For casts #18  REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0, PAR and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the comments about processing including a note about bottom depths being based on maximum sampling depth plus altimetry at the bottom of casts.

The Standards Check routine was run and the only comment concerned non-standard formats for temperature and salinity for cast #22 – this was deliberate choice so no change was made.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 5m were almost all between 103% and 105% with the exception of the test cast which was >140% and 2 casts that were ~ 95%. In the offshore area values from 101-105% are typical at 5m while the test cast was in Saanich Inlet where surface saturation is frequently very high. This check is very rough but suggests that CTD DO values are reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The primary channels were noisy in the downcast data so the secondary were chosen for archival, but there does not appear to be a problem in the bottle file, so the primary are suitable for all casts.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, PAR and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-014-bottles-final.xls. The entries were compared with the rosette log sheets. Some salinity bottle values had been missed in event #2, so processing of that file was repeated and the missing data were then added to the spreadsheet.  A few flags had not been adjusted to reflect comments marked ALL, so those were fixed. No other problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise.

The IOS SBE TSG45 data were delivered in 5 files but 4 contained only a few records and were from the last day of the cruise when the ship was not moving.

The useful file was TSG-_20190717-181931.raw. It has extension RAW but is in csv format, so it was opened in EXCEL and saved as a CSV file, 2019-014-tsg.csv.

The file is very large making it awkward to handle. It will be divided into separate files for each day.

The spreadsheet was adjusted as follows:

- The first 19 records were removed because they were heavily corrupted.

- 2 lines of headers were added – channel names and units.

- The file was then saved in CSV format – the name is not critical at this point.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column L. Then a concentration value was calculated in column F using offset -0.04 and scale 13.1 as determined in the most recent recalibration of the fluorometer.

- Copy/Special Paste was used to save the values and then the voltage channel was removed.

- A file break column was filled with 2019-014-0717-181931 and then as each new day started it was changed to enable the file to be divided into daily files. (format 2019-014-MMDD-HHMMSS with 000000 for the time for all files except the first.)

- Time and Date formats are a problem – when converting from RAW choose TEXT but once opened in EXCEL set Time Format  to HH:MM:SS and save the file again. (Any time the file is opened in EXCEL the time format may have to be set again before saving.) Using “custom” formats works best to ensure the hours are in 2-digit format.

The file was then converted to IOS format with some header info added.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Day of Year and set Time Zero to the end of the previous year. A record number was also added.

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end.

a.) Plots

A track plot was produced and added to the end of this document.

Time-series plots were examined on screen. Some editing will be required for the following files:

- July 17 – during the first half hour the flow was variable and heating in the loop high.

- July 20 – one very large spike in lab temperature and salinity – this was due to a pad value listed as NaN rather than -99. A text editor was used to change to -99.0000.

- July 21 - 1 small salinity spike around 0430

- July 29 – the flow was off at the end of the file so temperature, fluorescence and salinity need to be padded.

- There were a few cases of other odd pad values that had to be fixed to get reasonable plots.

Overall the data look good with no salinity spiking other the few cases mentioned above. The temperature differences are noisy but there is a lot of variability. During stops the differences look normal.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2019-014-ctd-tsg-comp.xlsx. There were 26 CTD casts that overlapped with the TSG records, but of those had no CTD data at the right depth. There were no CTD data from July 17, 18, 25, 28 and 29.

8 TSG files were opened in EXCEL. The median and standard deviations over 30s were calculated for intake and lab temperature, fluorescence and salinity.

TSG Data were selected for the times of CTD casts and then added to file 2019-014-ctd-tsg-loop-rosette-comp.xlsx.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no differences greater than 0.00078º while the median differences were 0.00003º and 0.00005º. So the matches look good.

c.) Comparisons

- Comparison of T, S and fluorescence from TSG and CTD data

|  | Tint-Tctd | Tlab-Tctd | Stsg-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| average | 0.1321 | 0.4070 | -0.0981 | 1.279 |
| median | 0.0566 | 0.3218 | -0.0998 | 1.197 |
| stdev | 0.3022 | 0.3474 | 0.0399 | 0.450 |
| min | -0.0701 | 0.2362 | -0.1419 | 0.314 |
| max | 1.3855 | 1.9365 | 0.0267 | 2.162 |

The intake temperature is higher than the CTD temperature by a median of 0.057C°. During 2019-006 when this equipment was last used, the median difference was 0.015 C°, but 2019-014 sampled waters with greater variability so small differences in sampling depth and time of sampling likely explain the higher value. This is still good agreement so the intake thermistor calibration appears to be good.

The heating in the loop (Lab temperature – Intake Temperature) has a median value of ~0.27C°. This is slightly lower than the results for 2019-006 which probably reflects slightly higher intake temperatures. The results suggest that whatever problem arose earlier in the year during 2019-001 and 2019-005, when heating in the loop appeared to be much larger than usual, has been resolved.

The TSG salinity is lower than the CTD salinity by a median of 0.10psu. The differences were fairly steady up to event #27 with the TSG reading low by between ~0.07 and 0.08psu after which the difference grew steadily until the TSG was reading low by ~0.14psu. There are no CTD data for the final 2 days. During 2019-006 it was low by a median of 0.057psu. The flow rate looks very steady and there is no indication of bubbles. TSG Fluorescence is low in the offshore areas but there is a slight gradual increase during the period when the salinity differences increased. Possibly there was some biological contamination affecting the TSG salinity readings.

The ratio of TSG fluorescence to CTD fluorescence has a median value of 1.2 with a range from 0.3 to 2.2. The ratio of TSG becomes somewhat higher at about the same time as the TSG salinity appears to get lower than CTD salinity. All fluorescence values were <1ug/L.

- There were no loop samples and no near-surface rosette salinity samples.

d.) Calibration History

- During 2018-026 this system was used for the first time. The intake temperature sensor read higher than the CTD temperature by a median of 0.03C°. The lab temperature was higher than the intake temperature by more than expected (~0.5) based on previous experience with heating in the loop on the Tully. The salinity traces were quite smooth with only a few very small spikes. Salinity read lower than the CTD by a median of 0.009psu. The differences were not sufficiently consistent to allow recalibration. The data were archived but with a warning that there was insufficient calibration sampling to assess quality.

- The results from 2018-025 suggest that there were problems with flow in the loop.

- During 2018-034 Salinity was lower than the CTD by a median of 0.030psu. Lab temperatures were higher than CTD temperatures by a median of 0.66 C°. The intake temperatures were high by a median of 0.072 C° using all data or 0.058 C° in a “quiet section”. Fluorescence from the TSG was lower than the CTD fluorescence except in areas where CHL was low. The comparisons have a lot of scatter and recalibration was not attempted. The data were archived but with a comment about the intake temperature being further from the CTD data than expected and heating in the loop being higher than when the SBE21 was in use. Because of the limitations only 3 decimal places will be shown for temperature and salinity.

- During 2018-040 the TSG fluorescence was removed as it looked unreliable, starting with values close to the CTD but increasing until it was 60 times the CTD values. Salinity was lower than that from the CTD by a median of 0.032 and lower than loop samples by 0.034psu. The intake temperature was higher than the CTD temperature by a median of 0.0032C°.  The lab temperature was higher than the intake temperature by ~0.28 C°.

- During 2019-001 the TSG fluorescence values started unbelievably high, but the values dropped through the cruise getting closer to CTD and loop values. The TSG intake temperature was higher than the CTD by about 0.05C°. The lab temperature was higher than the intake temperature by ~0.5C°. Salinity was lower than CTD and loop by between 0.003 and 0.009psu.

- During 2019-005 the TSG fluorescence values were about 50% of those from the CTD and ~40% of loop extracted chlorophyll samples. The pattern of variability is similar to that from the CTD. The TSG intake temperature was higher than the CTD by about 0.19C°. The lab temperature was higher than the intake temperature by ~0.5C°. Salinity was higher than CTD and loop by 0.4 and 0.34psu, respectively. The salinity is usually lower than the CTD and loop and the two TSG temperature channels are way out of line with expected values. TSG data were not archived.

- During 2019-006 the TSG fluorescence had a large temporal drift that was not related to chlorophyll level. The data were not archived. The TSG intake temperature was higher than the CTD by 0.013C°. The lab temperature was higher than the CTD by ~0.32° and higher than the intake by ~0.28°. TSG salinity was reading lower than the CTD by about 0.057psu and lower than loop samples by 0.061psu.

e.) Conclusions re IOS TSG

1. The TSG clock worked well.

2. There were 25 points of comparison with CTD data. There were no CTD casts on July 17, 18, 25, 28 and 29.

3. After an initial period large variability, the flow rate was steady with values ~2L/min. This is much higher than for other recent uses of this equipment.

4. The intake temperature is reading higher than the CTD by a median difference of 0.057C° which is a little higher than for 2019-006 but is quite reasonable given greater variability.

5. TSG Salinity is reading lower than the CTD salinity by a median of 0.10psu. The differences grew with time during the cruise with values between ~0.07 and 0.08psu until event #27 after which it rose steadily to 0.14psu. We expect salinity to be slightly low due to small bubbles though there is no evidence in the traces that bubbles were a serious issue; in fact, the standard deviation in the TSG salinity is very low. The TSG fluorescence also appears to be drifting upward relative to the CTD values during the same period, so there may be a biological contamination issue.

6. The heating in the loop (Lab temperature – Intake Temperature) has a median value of ~0.27C°. This is slightly lower than the results for 2019-006 which probably reflects slightly higher intake temperatures.

7. There are no loop samples so there is no way to judge underway quality or calibration.

8. The intake temperature will be recalibrated by subtracting 0.057C°. The salinity will be recalibrated by adding 0.075psu for July 17 to 22, 0.095psu for July 23 to 25 and 0.130psu for July 26 to 29 as a reasonable estimate.

9. Only 3 decimal places should be used for the temperature and salinity channels.

10. While Fluorescence values appear to be drifting upwards slightly this is much less significant than during the previous cruise.

g.) Editing

All *.REO files were copied to EDT and then the following files were put through CTDEDIT:

- July 17 – Temperature, salinity, conductivity, fluorescence removed for records #1-1024 because flow was off or not well established; salinity was cleaned very lightly.

- July 21 – Salinity was cleaned very lightly

- July 29 – Temperature, salinity, conductivity, fluorescence values were padded from record #9899 to the end because flow was off.

The file for July 20 was edited using a text editor to replace records with NaN with -99.0000.

h.) Recalibration

CALIBRATE was applied to decrease intake temperature by 0.057C° and increase salinity by 0.075psu, 0.095psu or 0.130psu depending on date using files 2019-014-recal1.ccf, 2019-014-recal2.ccf and 2019-014-recal3.ccf.

Sigma-T was derived.

Further problems with odd pad values (such as -98.925) were found in the salinity pad values for 1 cast. This was fixed in the DQT2 file.

i.) Preparing Final Files

REMOVE was used to remove channel Pressure, Temperature:Difference and record #.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments. The temperature and salinity format was chosen as F8.3.

The files were saved with format 2019-014-MMDD-HHMMSS.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

Particulars

General – PAR was included in the configuration file but no PAR sensor was mounted during this cruise.

1. Test cast – all bottles fired but none sampled. CTD file prepared but not CHE.

5. Niskin #2 fired at bottom.

9. Acquisition did not start until the CTD was at about 10m on the downcast.

11. Bottle 3 sample 60 skipped as cast only wen to 2883m

22. Upcast only archived.

27. Bottle 5 not fired (intentionally). Sample 115 not taken.

35. Stopped ~108m of downcast to reposition ship and then continued. Niskin 21 at surface not fired.

37. Stopped ~753m of downcast to reposition ship and then continued. Bottle #20 fired out of order because #2 fired accidentally at 2364m– not sampled. Sample 197 was taken with bottle 20 at 2000m.

40. Bottle 18 fired out of order – bottom deeper than anticipated.

43. Stopped ~62m and ~880m of downcast for ship to adjust position

47. Bottle 1 fired at 722m on downcast. Firing of other bottles started at #2.

49. Extra nutrient samples taken for cool storage study.

51. Sample numbers repeated from previous cast. #283 to 287 were replaced with 9283 to 9287.

59. Niskin bottle did not fire.

2019-014

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 21Nov2018 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 16Dec2018 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3640 |  | Jan 2019 |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 2Aug17 | Factory |  |  |
| Conductivity | 0620 | 2Aug17 | Factory |  |  |
| Temperature SBE38 | ? | ? | Factory |  |  |
| WETStar Fluorometer | 953 | 5May2017 | Factory |  |  |