## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6 | Run at 2020/01/30 14:55:35 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 0.2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2019-023-0002.mrg | 2019-023-0002.sam | 241 | 5.609 | 5.6088 | 0.058754 | -0.00018263 | 29.8302 | 29.8296 | 0.013931 | -0.00059318 | 29.8302 | 29.8357 | 0.014531 | 0.0055065
2019-023-0006.mrg | 2019-023-0006.sam | 350 | 5.1 | 5.1959 | 0.083074 | 0.095851 | 31.2545 | 31.2248 | 0.0014056 | -0.029665 | 31.2545 | 31.2335 | 0.0012586 | -0.021006
2019-023-0027.mrg | 2019-023-0027.sam | 204 | 5.055 | 4.9747 | 0.077361 | -0.080343 | 32.2483 | 32.2356 | 0.001293 | -0.012688 | 32.2483 | 32.2437 | 0.0015287 | -0.0046425
2019-023-0034.mrg | 2019-023-0034.sam | 194 | 1000.65 | 1000.62 | 0.0886 | -0.032837 | 34.3893 | 34.3796 | 0.00026958 | -0.0097046 | 34.3893 | 34.3893 | 0.00025026 | -3.0518e-05
2019-023-0034.mrg | 2019-023-0034.sam | 89 | 750.341 | 750.366 | 0.11582 | 0.024597 | 34.2452 | 34.2346 | 0.00026957 | -0.010563 | 34.2452 | 34.2439 | 0.00025939 | -0.0013275
2019-023-0034.mrg | 2019-023-0034.sam | 214 | 500.868 | 500.893 | 0.078446 | 0.025146 | 34.0715 | 34.0617 | 0.00025059 | -0.0097847 | 34.0715 | 34.0709 | 0.00022875 | -0.00064468
2019-023-0044.mrg | 2019-023-0044.sam | 138 | 1251.17 | 1251.21 | 0.099302 | 0.03479 | 34.4621 | 34.4513 | 0.00036996 | -0.01083 | 34.4621 | 34.461 | 0.00040462 | -0.0011025
2019-023-0044.mrg | 2019-023-0044.sam | 49 | 1000.21 | 1000.21 | 0.11352 | 0.00067139 | 34.3542 | 34.3439 | 0.00064954 | -0.010269 | 34.3542 | 34.3527 | 0.00065006 | -0.0014725
2019-023-0044.mrg | 2019-023-0044.sam | 170 | 750.552 | 750.528 | 0.11679 | -0.024414 | 34.2069 | 34.1981 | 0.00090442 | -0.0087891 | 34.2069 | 34.2074 | 0.00084636 | 0.00049591
2019-023-0044.mrg | 2019-023-0044.sam | 78 | 500.403 | 500.444 | 0.10221 | 0.041321 | 34.0812 | 34.0723 | 0.00033208 | -0.0089149 | 34.0812 | 34.0814 | 0.00027799 | 0.00017929
2019-023-0044.mrg | 2019-023-0044.sam | 71 | 5.212 | 5.1938 | 0.12339 | -0.018225 | 32.2876 | 32.278 | 0.00029264 | -0.0095978 | 32.2876 | 32.2858 | 0.00024054 | -0.0017967
2019-023-0051.mrg | 2019-023-0051.sam | 115 | 5.405 | 5.4165 | 0.10565 | 0.011521 | 32.2205 | 32.2096 | 0.00027556 | -0.010944 | 32.2205 | 32.2178 | 0.00021944 | -0.0027351
2019-023-0069.mrg | 2019-023-0069.sam | 241 | 4.971 | 4.9509 | 0.08653 | -0.020129 | 32.204 | 32.1954 | 0.00030775 | -0.0086174 | 32.204 | 32.2033 | 0.00026654 | -0.00069427
2019-023-0071.mrg | 2019-023-0071.sam | 88 | 750.499 | 750.524 | 0.12072 | 0.02533 | 34.2232 | 34.2129 | 0.00087447 | -0.010342 | 34.2232 | 34.2227 | 0.00090323 | -0.00047684
2019-023-0071.mrg | 2019-023-0071.sam | 70 | 501.054 | 501.004 | 0.091235 | -0.050018 | 34.0754 | 34.0669 | 0.00030947 | -0.0085144 | 34.0754 | 34.0761 | 0.00029344 | 0.0006752
2019-023-0076.mrg | 2019-023-0076.sam | 137 | 1500.05 | 1500.04 | 0.11881 | -0.012451 | 34.5187 | 34.5095 | 0.00053409 | -0.0092239 | 34.5187 | 34.5193 | 0.00035236 | 0.00055313
2019-023-0076.mrg | 2019-023-0076.sam | 143 | 1249.88 | 1249.87 | 0.118 | -0.0081787 | 34.463 | 34.4549 | 0.00028318 | -0.0081177 | 34.463 | 34.4639 | 0.00023672 | 0.00089264
2019-023-0076.mrg | 2019-023-0076.sam | 162 | 1000.44 | 1000.48 | 0.1181 | 0.040039 | 34.3824 | 34.3745 | 0.0002512 | -0.0079231 | 34.3824 | 34.3832 | 0.00024855 | 0.0007515
2019-023-0076.mrg | 2019-023-0076.sam | 85 | 750.801 | 750.717 | 0.089359 | -0.083984 | 34.2462 | 34.2396 | 0.00031431 | -0.006649 | 34.2462 | 34.2482 | 0.0003456 | 0.0019722
2019-023-0076.mrg | 2019-023-0076.sam | 82 | 500.146 | 500.089 | 0.11117 | -0.057098 | 34.0899 | 34.0838 | 0.00031774 | -0.0060806 | 34.0899 | 34.0926 | 0.0002595 | 0.0026627
2019-023-0076.mrg | 2019-023-0076.sam | 105 | 400.083 | 400.07 | 0.11323 | -0.013458 | 34.0272 | 34.0201 | 0.00044288 | -0.007061 | 34.0272 | 34.0285 | 0.00029626 | 0.0012817
2019-023-0076.mrg | 2019-023-0076.sam | 126 | 299.94 | 299.857 | 0.095541 | -0.083374 | 33.9813 | 33.9736 | 0.00034282 | -0.0077286 | 33.9813 | 33.982 | 0.00023857 | 0.00074768
2019-023-0076.mrg | 2019-023-0076.sam | 45 | 249.85 | 249.854 | 0.10983 | 0.0044403 | 33.957 | 33.9504 | 0.00035442 | -0.0065651 | 33.957 | 33.9592 | 0.00021977 | 0.002224
2019-023-0076.mrg | 2019-023-0076.sam | 106 | 200.454 | 200.426 | 0.11229 | -0.028305 | 33.9292 | 33.9224 | 0.0004017 | -0.0068054 | 33.9292 | 33.931 | 0.00089757 | 0.0017662
2019-023-0076.mrg | 2019-023-0076.sam | 84 | 175.48 | 175.452 | 0.11954 | -0.027847 | 33.8899 | 33.8828 | 0.00054129 | -0.0070801 | 33.8899 | 33.8919 | 0.0005365 | 0.0020142
2019-023-0076.mrg | 2019-023-0076.sam | 149 | 150.049 | 150.053 | 0.11789 | 0.0041962 | 33.7328 | 33.7056 | 0.001202 | -0.027187 | 33.7328 | 33.7141 | 0.00073009 | -0.018738
2019-023-0076.mrg | 2019-023-0076.sam | 85 | 125.16 | 125.184 | 0.1082 | 0.023682 | 33.3454 | 33.3311 | 0.0050448 | -0.014332 | 33.3454 | 33.3408 | 0.0048721 | -0.0045929
2019-023-0076.mrg | 2019-023-0076.sam | 167 | 100.136 | 100.22 | 0.093301 | 0.08387 | 32.7256 | 32.6854 | 0.00089063 | -0.040199 | 32.7256 | 32.6939 | 0.00055802 | -0.031666
2019-023-0076.mrg | 2019-023-0076.sam | 180 | 75.089 | 75.0778 | 0.14175 | -0.011162 | 32.4304 | 32.4244 | 0.0008663 | -0.0060158 | 32.4304 | 32.4331 | 0.0005256 | 0.0027008
2019-023-0076.mrg | 2019-023-0076.sam | 212 | 50.249 | 50.2436 | 0.10987 | -0.0053978 | 32.3821 | 32.372 | 0.0010826 | -0.010052 | 32.3821 | 32.3808 | 0.0012142 | -0.0013084
2019-023-0076.mrg | 2019-023-0076.sam | 147 | 30.198 | 30.2145 | 0.12354 | 0.016462 | 32.3022 | 32.2824 | 0.0050227 | -0.019836 | 32.3022 | 32.2903 | 0.0023368 | -0.011879
2019-023-0076.mrg | 2019-023-0076.sam | 241 | 19.951 | 19.9512 | 0.036406 | 0.00024414 | 32.2699 | 32.2687 | 0.0011 | -0.0012321 | 32.2699 | 32.2755 | 0.00067805 | 0.0055618
2019-023-0076.mrg | 2019-023-0076.sam | 229 | 9.969 | 9.9804 | 0.11841 | 0.011424 | 32.1998 | 32.1901 | 0.0011591 | -0.0096779 | 32.1998 | 32.1968 | 0.0015605 | -0.0029793
2019-023-0076.mrg | 2019-023-0076.sam | 368 | 5.032 | 5.0607 | 0.084707 | 0.028726 | 32.1953 | 32.187 | 0.00030421 | -0.0082626 | 32.1953 | 32.1943 | 0.00037329 | -0.0010261
2019-023-0079.mrg | 2019-023-0079.sam | 170 | 6.068 | 5.9902 | 0.089819 | -0.0778 | 32.1895 | 32.1804 | 0.00044399 | -0.0091438 | 32.1895 | 32.186 | 0.00048196 | -0.0034637
2019-023-0084.mrg | 2019-023-0084.sam | 237 | 750.547 | 750.543 | 0.086349 | -0.00354 | 34.246 | 34.2379 | 0.00027364 | -0.0081444 | 34.246 | 34.247 | 0.00032325 | 0.00097656
2019-023-0084.mrg | 2019-023-0084.sam | 181 | 4.921 | 4.8932 | 0.11973 | -0.027796 | 32.2615 | 32.2537 | 0.00044244 | -0.007782 | 32.2615 | 32.2611 | 0.00056509 | -0.00035477
2019-023-0094.mrg | 2019-023-0094.sam | 241 | 749.936 | 749.936 | 0.093686 | -0.00018311 | 34.2544 | 34.243 | 0.0003167 | -0.01144 | 34.2544 | 34.2523 | 0.00051134 | -0.0020905
2019-023-0094.mrg | 2019-023-0094.sam | 205 | 4.799 | 4.8232 | 0.11712 | 0.024147 | 32.1654 | 32.1579 | 0.00031998 | -0.0075378 | 32.1654 | 32.1649 | 0.00028663 | -0.00048065
2019-023-0099.mrg | 2019-023-0099.sam | 116 | 2214.2 | 2214.25 | 0.089974 | 0.045166 | 34.6118 | 34.6054 | 0.00026601 | -0.0064201 | 34.6118 | 34.6147 | 0.00025213 | 0.0029259
2019-023-0099.mrg | 2019-023-0099.sam | 64 | 1500.83 | 1500.79 | 0.11202 | -0.04126 | 34.5083 | 34.5008 | 0.00029056 | -0.0074806 | 34.5083 | 34.5095 | 0.00028943 | 0.0012436
2019-023-0099.mrg | 2019-023-0099.sam | 15 | 1248.58 | 1248.57 | 0.12947 | -0.0040283 | 34.441 | 34.4335 | 0.00021097 | -0.0075226 | 34.441 | 34.4422 | 0.00036176 | 0.0012054
2019-023-0099.mrg | 2019-023-0099.sam | 36 | 1000.57 | 1000.56 | 0.10968 | -0.0040283 | 34.3558 | 34.3508 | 0.00025555 | -0.0049934 | 34.3558 | 34.3595 | 0.00024773 | 0.0036507
2019-023-0099.mrg | 2019-023-0099.sam | 275 | 5.137 | 5.0812 | 0.073267 | -0.055811 | 32.2075 | 32.1993 | 0.0020667 | -0.0082436 | 32.2075 | 32.2066 | 0.0004392 | -0.00085068
2019-023-0101.mrg | 2019-023-0101.sam | 17 | 1251.85 | 1251.85 | 0.1272 | -0.0037842 | 34.4417 | 34.4332 | 0.00022122 | -0.0085487 | 34.4417 | 34.4426 | 0.00020315 | 0.00085831
2019-023-0101.mrg | 2019-023-0101.sam | 93 | 1001.32 | 1001.22 | 0.097045 | -0.098328 | 34.3645 | 34.3568 | 0.00028007 | -0.0076828 | 34.3645 | 34.3659 | 0.00024413 | 0.0013695
2019-023-0105.mrg | 2019-023-0105.sam | 63 | 150.745 | 150.754 | 0.12021 | 0.0094452 | 33.8524 | 33.845 | 0.0015316 | -0.0073967 | 33.8524 | 33.8547 | 0.0016378 | 0.0022888
2019-023-0121.mrg | 2019-023-0121.sam | 93 | 1500.38 | 1500.4 | 0.10216 | 0.023804 | 34.5028 | 34.4949 | 0.00024983 | -0.0078812 | 34.5028 | 34.5039 | 0.00024486 | 0.0010681
2019-023-0121.mrg | 2019-023-0121.sam | 11 | 1250.39 | 1250.36 | 0.13473 | -0.022949 | 34.4423 | 34.4357 | 0.0003227 | -0.0066261 | 34.4423 | 34.4447 | 0.00040855 | 0.0024376
2019-023-0121.mrg | 2019-023-0121.sam | 82 | 999.962 | 999.87 | 0.081421 | -0.091675 | 34.3644 | 34.3551 | 0.00026929 | -0.0093269 | 34.3644 | 34.3639 | 0.00026983 | -0.00049973
2019-023-0123.mrg | 2019-023-0123.sam | 76 | 1499.58 | 1499.59 | 0.1089 | 0.01416 | 34.5043 | 34.4963 | 0.00024231 | -0.0079498 | 34.5043 | 34.5057 | 0.00022011 | 0.0013542
2019-023-0123.mrg | 2019-023-0123.sam | 39 | 1251.05 | 1251.05 | 0.11557 | 0.0031738 | 34.4407 | 34.4326 | 0.00031126 | -0.0080833 | 34.4407 | 34.4415 | 0.0002856 | 0.00079346
2019-023-0123.mrg | 2019-023-0123.sam | 116 | 1000.7 | 1000.75 | 0.10635 | 0.052795 | 34.3638 | 34.3562 | 0.00039073 | -0.0076332 | 34.3638 | 34.3652 | 0.00036381 | 0.0013809
2019-023-0123.mrg | 2019-023-0123.sam | 174 | 5.101 | 5.0508 | 0.079317 | -0.050218 | 32.0263 | 32.0272 | 0.0028602 | 0.00094986 | 32.0263 | 32.0353 | 0.002518 | 0.0090065
2019-023-0170.mrg | 2019-023-0170.sam | 241 | 4.949 | 4.9489 | 0.033244 | -5.3883e-05 | 30.9871 | 31.1365 | 0.0012558 | 0.14937 | 30.9871 | 31.1465 | 0.00088073 | 0.15943
2019-023-0184.mrg | 2019-023-0184.sam | 295 | 4.939 | 4.9585 | 0.039045 | 0.019468 | 26.3684 | 25.8467 | 0.0018722 | -0.52165 | 26.3684 | 25.8535 | 0.0048599 | -0.51492
2019-023-0195.mrg | 2019-023-0195.sam | 241 | 5.217 | 5.2171 | 0.031582 | 0.00010776 | 27.5159 | 26.3558 | 0.0022341 | -1.1601 | 27.5159 | 26.364 | 0.0032158 | -1.1519
2019-023-0203.mrg | 2019-023-0203.sam | 220 | 4.888 | 4.9316 | 0.034402 | 0.043613 | 28.5856 | 28.4227 | 0.028644 | -0.16286 | 28.5856 | 28.4434 | 0.030712 | -0.14222
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 21
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2019-023-0002.mrg | 2019-023-0002.sam | 12 | 12 | 2019-023-0002.mrg & 2019-023-0002.sam
2 | 2019-023-0006.mrg | 2019-023-0006.sam | 14 | 14 | 2019-023-0006.mrg & 2019-023-0006.sam
3 | 2019-023-0027.mrg | 2019-023-0027.sam | 16 | 16 | 2019-023-0027.mrg & 2019-023-0027.sam
4 | 2019-023-0034.mrg | 2019-023-0034.sam | 18 | 20 | 2019-023-0034.mrg & 2019-023-0034.sam
5 | 2019-023-0044.mrg | 2019-023-0044.sam | 22 | 26 | 2019-023-0044.mrg & 2019-023-0044.sam
6 | 2019-023-0051.mrg | 2019-023-0051.sam | 28 | 28 | 2019-023-0051.mrg & 2019-023-0051.sam
7 | 2019-023-0069.mrg | 2019-023-0069.sam | 30 | 30 | 2019-023-0069.mrg & 2019-023-0069.sam
8 | 2019-023-0071.mrg | 2019-023-0071.sam | 32 | 33 | 2019-023-0071.mrg & 2019-023-0071.sam
9 | 2019-023-0076.mrg | 2019-023-0076.sam | 35 | 53 | 2019-023-0076.mrg & 2019-023-0076.sam
10 | 2019-023-0079.mrg | 2019-023-0079.sam | 55 | 55 | 2019-023-0079.mrg & 2019-023-0079.sam
11 | 2019-023-0084.mrg | 2019-023-0084.sam | 57 | 58 | 2019-023-0084.mrg & 2019-023-0084.sam
12 | 2019-023-0094.mrg | 2019-023-0094.sam | 60 | 61 | 2019-023-0094.mrg & 2019-023-0094.sam
13 | 2019-023-0099.mrg | 2019-023-0099.sam | 63 | 67 | 2019-023-0099.mrg & 2019-023-0099.sam
14 | 2019-023-0101.mrg | 2019-023-0101.sam | 69 | 70 | 2019-023-0101.mrg & 2019-023-0101.sam
15 | 2019-023-0105.mrg | 2019-023-0105.sam | 72 | 72 | 2019-023-0105.mrg & 2019-023-0105.sam
16 | 2019-023-0121.mrg | 2019-023-0121.sam | 74 | 76 | 2019-023-0121.mrg & 2019-023-0121.sam
17 | 2019-023-0123.mrg | 2019-023-0123.sam | 78 | 81 | 2019-023-0123.mrg & 2019-023-0123.sam
18 | 2019-023-0170.mrg | 2019-023-0170.sam | 83 | 83 | 2019-023-0170.mrg & 2019-023-0170.sam
19 | 2019-023-0184.mrg | 2019-023-0184.sam | 85 | 85 | 2019-023-0184.mrg & 2019-023-0184.sam
20 | 2019-023-0195.mrg | 2019-023-0195.sam | 87 | 87 | 2019-023-0195.mrg & 2019-023-0195.sam
21 | 2019-023-0203.mrg | 2019-023-0203.sam | 89 | 89 | 2019-023-0203.mrg & 2019-023-0203.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 21
2 | (Sal1 - Sal_Bot) | Pressure | 21
3 | (Sal0 - Sal_Bot) | File Pair # | 21
4 | (Sal0 - Sal_Bot) | File Pair # | 21

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 8.0554e-05 |  | X | CALC Y
m= | -2.9279e-08 | 5.8589e-07 |  | Std Dev of Resid |  | 0.0015248 |  | 175.452 | -0.0081804
b= | -0.0081753 | 0.0006039 |  | Resid Sum of Squares |  | 7.2079e-05 |  | 2214.25 | -0.0082401
# of Values Extracted = |  |  | 58
# of Values Used in Fit = |  |  | 33
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.6088 | 29.8296 | 0.013931 | -0.00059318 | -0.0081755 | -0.0075823 |  | 1000.62 | -0.0097046
2 | NO | 5.1959 | 31.2248 | 0.0014056 | -0.029665 | -0.0081755 | 0.02149 |  | 750.366 | -0.010563
3 | NO | 4.9747 | 32.2356 | 0.001293 | -0.012688 | -0.0081754 | 0.0045123 |  | 500.893 | -0.0097847
4 | YES | 1000.62 | 34.3796 | 0.00026958 | -0.0097046 | -0.0082046 | 0.0015 |  | 1251.21 | -0.01083
4 | YES | 750.366 | 34.2346 | 0.00026957 | -0.010563 | -0.0081973 | 0.0023656 |  | 1000.21 | -0.010269
4 | YES | 500.893 | 34.0617 | 0.00025059 | -0.0097847 | -0.00819 | 0.0015947 |  | 750.528 | -0.0087891
5 | YES | 1251.21 | 34.4513 | 0.00036996 | -0.01083 | -0.0082119 | 0.002618 |  | 500.444 | -0.0089149
5 | YES | 1000.21 | 34.3439 | 0.00064954 | -0.010269 | -0.0082046 | 0.0020646 |  | 750.524 | -0.010342
5 | YES | 750.528 | 34.1981 | 0.00090442 | -0.0087891 | -0.0081973 | 0.00059178 |  | 501.004 | -0.0085144
5 | YES | 500.444 | 34.0723 | 0.00033208 | -0.0089149 | -0.00819 | 0.000725 |  | 1500.04 | -0.0092239
5 | NO | 5.1938 | 32.278 | 0.00029264 | -0.0095978 | -0.0081755 | 0.0014223 |  | 1249.87 | -0.0081177
6 | NO | 5.4165 | 32.2096 | 0.00027556 | -0.010944 | -0.0081755 | 0.0027689 |  | 1000.48 | -0.0079231
7 | NO | 4.9509 | 32.1954 | 0.00030775 | -0.0086174 | -0.0081754 | 0.00044195 |  | 750.717 | -0.006649
8 | YES | 750.524 | 34.2129 | 0.00087447 | -0.010342 | -0.0081973 | 0.0021443 |  | 500.089 | -0.0060806
8 | YES | 501.004 | 34.0669 | 0.00030947 | -0.0085144 | -0.00819 | 0.00032443 |  | 400.07 | -0.007061
9 | YES | 1500.04 | 34.5095 | 0.00053409 | -0.0092239 | -0.0082192 | 0.0010047 |  | 299.857 | -0.0077286
9 | YES | 1249.87 | 34.4549 | 0.00028318 | -0.0081177 | -0.0082119 | -9.4216e-05 |  | 249.854 | -0.0065651
9 | YES | 1000.48 | 34.3745 | 0.0002512 | -0.0079231 | -0.0082046 | -0.00028146 |  | 200.426 | -0.0068054
9 | YES | 750.717 | 34.2396 | 0.00031431 | -0.006649 | -0.0081973 | -0.0015483 |  | 175.452 | -0.0070801
9 | YES | 500.089 | 34.0838 | 0.00031774 | -0.0060806 | -0.0081899 | -0.0021093 |  | 750.543 | -0.0081444
9 | YES | 400.07 | 34.0201 | 0.00044288 | -0.007061 | -0.008187 | -0.001126 |  | 749.936 | -0.01144
9 | YES | 299.857 | 33.9736 | 0.00034282 | -0.0077286 | -0.0081841 | -0.0004555 |  | 2214.25 | -0.0064201
9 | YES | 249.854 | 33.9504 | 0.00035442 | -0.0065651 | -0.0081826 | -0.0016175 |  | 1500.79 | -0.0074806
9 | YES | 200.426 | 33.9224 | 0.0004017 | -0.0068054 | -0.0081812 | -0.0013757 |  | 1248.57 | -0.0075226
9 | YES | 175.452 | 33.8828 | 0.00054129 | -0.0070801 | -0.0081804 | -0.0011004 |  | 1000.56 | -0.0049934
9 | NO | 150.053 | 33.7056 | 0.001202 | -0.027187 | -0.0081797 | 0.019008 |  | 1251.85 | -0.0085487
9 | NO | 125.184 | 33.3311 | 0.0050448 | -0.014332 | -0.008179 | 0.0061528 |  | 1001.22 | -0.0076828
9 | NO | 100.22 | 32.6854 | 0.00089063 | -0.040199 | -0.0081782 | 0.032021 |  | 1500.4 | -0.0078812
9 | NO | 75.0778 | 32.4244 | 0.0008663 | -0.0060158 | -0.0081775 | -0.0021617 |  | 1250.36 | -0.0066261
9 | NO | 50.2436 | 32.372 | 0.0010826 | -0.010052 | -0.0081768 | 0.0018749 |  | 999.87 | -0.0093269
9 | NO | 30.2145 | 32.2824 | 0.0050227 | -0.019836 | -0.0081762 | 0.01166 |  | 1499.59 | -0.0079498
9 | NO | 19.9512 | 32.2687 | 0.0011 | -0.0012321 | -0.0081759 | -0.0069437 |  | 1251.05 | -0.0080833
9 | NO | 9.9804 | 32.1901 | 0.0011591 | -0.0096779 | -0.0081756 | 0.0015023 |  | 1000.75 | -0.0076332 | -0.0082024 | 0.0015009
9 | NO | 5.0607 | 32.187 | 0.00030421 | -0.0082626 | -0.0081754 | 8.7181e-05
10 | NO | 5.9902 | 32.1804 | 0.00044399 | -0.0091438 | -0.0081755 | 0.00096835
11 | YES | 750.543 | 34.2379 | 0.00027364 | -0.0081444 | -0.0081973 | -5.2896e-05
11 | NO | 4.8932 | 32.2537 | 0.00044244 | -0.007782 | -0.0081754 | -0.00039346
12 | YES | 749.936 | 34.243 | 0.0003167 | -0.01144 | -0.0081973 | 0.003243
12 | NO | 4.8232 | 32.1579 | 0.00031998 | -0.0075378 | -0.0081754 | -0.0006376
13 | YES | 2214.25 | 34.6054 | 0.00026601 | -0.0064201 | -0.0082401 | -0.00182
13 | YES | 1500.79 | 34.5008 | 0.00029056 | -0.0074806 | -0.0082192 | -0.00073862
13 | YES | 1248.57 | 34.4335 | 0.00021097 | -0.0075226 | -0.0082119 | -0.00068928
13 | YES | 1000.56 | 34.3508 | 0.00025555 | -0.0049934 | -0.0082046 | -0.0032112
13 | NO | 5.0812 | 32.1993 | 0.0020667 | -0.0082436 | -0.0081754 | 6.8111e-05
14 | YES | 1251.85 | 34.4332 | 0.00022122 | -0.0085487 | -0.008212 | 0.00033679
14 | YES | 1001.22 | 34.3568 | 0.00028007 | -0.0076828 | -0.0082046 | -0.00052182
15 | NO | 150.754 | 33.845 | 0.0015316 | -0.0073967 | -0.0081797 | -0.00078301
16 | YES | 1500.4 | 34.4949 | 0.00024983 | -0.0078812 | -0.0082192 | -0.00033807
16 | YES | 1250.36 | 34.4357 | 0.0003227 | -0.0066261 | -0.0082119 | -0.0015858
16 | YES | 999.87 | 34.3551 | 0.00026929 | -0.0093269 | -0.0082046 | 0.0011224
17 | YES | 1499.59 | 34.4963 | 0.00024231 | -0.0079498 | -0.0082192 | -0.00026938
17 | YES | 1251.05 | 34.4326 | 0.00031126 | -0.0080833 | -0.0082119 | -0.00012859
17 | YES | 1000.75 | 34.3562 | 0.00039073 | -0.0076332 | -0.0082046 | -0.00057139
17 | NO | 5.0508 | 32.0272 | 0.0028602 | 0.00094986 | -0.0081754 | -0.0091253
18 | NO | 4.9489 | 31.1365 | 0.0012558 | 0.14937 | -0.0081754 | -0.15755
19 | NO | 4.9585 | 25.8467 | 0.0018722 | -0.52165 | -0.0081754 | 0.51348
20 | NO | 5.2171 | 26.3558 | 0.0022341 | -1.1601 | -0.0081755 | 1.1519
21 | NO | 4.9316 | 28.4227 | 0.028644 | -0.16286 | -0.0081754 | 0.15468
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0073564 |  | X | CALC Y
m= | 2.4338e-07 | 5.0778e-07 |  | Std Dev of Resid |  | 0.0013215 |  | 175.452 | 0.00066088
b= | 0.00061818 | 0.00052339 |  | Resid Sum of Squares |  | 5.414e-05 |  | 2214.25 | 0.0011571
# of Values Extracted = |  |  | 58
# of Values Used in Fit = |  |  | 33
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.6088 | 29.8357 | 0.014531 | 0.0055065 | 0.00061954 | -0.004887 |  | 1000.62 | -3.0518e-05
2 | NO | 5.1959 | 31.2335 | 0.0012586 | -0.021006 | 0.00061944 | 0.021625 |  | 750.366 | -0.0013275
3 | NO | 4.9747 | 32.2437 | 0.0015287 | -0.0046425 | 0.00061939 | 0.0052619 |  | 500.893 | -0.00064468
4 | YES | 1000.62 | 34.3893 | 0.00025026 | -3.0518e-05 | 0.00086171 | 0.00089223 |  | 1251.21 | -0.0011025
4 | YES | 750.366 | 34.2439 | 0.00025939 | -0.0013275 | 0.00080081 | 0.0021283 |  | 1000.21 | -0.0014725
4 | YES | 500.893 | 34.0709 | 0.00022875 | -0.00064468 | 0.00074009 | 0.0013848 |  | 750.528 | 0.00049591
5 | YES | 1251.21 | 34.461 | 0.00040462 | -0.0011025 | 0.0009227 | 0.0020252 |  | 500.444 | 0.00017929
5 | YES | 1000.21 | 34.3527 | 0.00065006 | -0.0014725 | 0.00086161 | 0.0023341 |  | 750.524 | -0.00047684
5 | YES | 750.528 | 34.2074 | 0.00084636 | 0.00049591 | 0.00080085 | 0.00030493 |  | 501.004 | 0.0006752
5 | YES | 500.444 | 34.0814 | 0.00027799 | 0.00017929 | 0.00073998 | 0.00056069 |  | 1500.04 | 0.00055313
5 | NO | 5.1938 | 32.2858 | 0.00024054 | -0.0017967 | 0.00061944 | 0.0024162 |  | 1249.87 | 0.00089264
6 | NO | 5.4165 | 32.2178 | 0.00021944 | -0.0027351 | 0.0006195 | 0.0033546 |  | 1000.48 | 0.0007515
7 | NO | 4.9509 | 32.2033 | 0.00026654 | -0.00069427 | 0.00061938 | 0.0013137 |  | 750.717 | 0.0019722
8 | YES | 750.524 | 34.2227 | 0.00090323 | -0.00047684 | 0.00080084 | 0.0012777 |  | 500.089 | 0.0026627
8 | YES | 501.004 | 34.0761 | 0.00029344 | 0.0006752 | 0.00074012 | 6.4914e-05 |  | 400.07 | 0.0012817
9 | YES | 1500.04 | 34.5193 | 0.00035236 | 0.00055313 | 0.00098326 | 0.00043013 |  | 299.857 | 0.00074768
9 | YES | 1249.87 | 34.4639 | 0.00023672 | 0.00089264 | 0.00092238 | 2.9738e-05 |  | 249.854 | 0.002224
9 | YES | 1000.48 | 34.3832 | 0.00024855 | 0.0007515 | 0.00086168 | 0.00011018 |  | 200.426 | 0.0017662
9 | YES | 750.717 | 34.2482 | 0.0003456 | 0.0019722 | 0.00080089 | -0.0011713 |  | 175.452 | 0.0020142
9 | YES | 500.089 | 34.0926 | 0.0002595 | 0.0026627 | 0.00073989 | -0.0019228 |  | 750.543 | 0.00097656
9 | YES | 400.07 | 34.0285 | 0.00029626 | 0.0012817 | 0.00071555 | -0.00056619 |  | 749.936 | -0.0020905
9 | YES | 299.857 | 33.982 | 0.00023857 | 0.00074768 | 0.00069116 | -5.6522e-05 |  | 2214.25 | 0.0029259
9 | YES | 249.854 | 33.9592 | 0.00021977 | 0.002224 | 0.00067899 | -0.001545 |  | 1500.79 | 0.0012436
9 | YES | 200.426 | 33.931 | 0.00089757 | 0.0017662 | 0.00066696 | -0.0010992 |  | 1248.57 | 0.0012054
9 | YES | 175.452 | 33.8919 | 0.0005365 | 0.0020142 | 0.00066088 | -0.0013533 |  | 1000.56 | 0.0036507
9 | NO | 150.053 | 33.7141 | 0.00073009 | -0.018738 | 0.0006547 | 0.019392 |  | 1251.85 | 0.00085831
9 | NO | 125.184 | 33.3408 | 0.0048721 | -0.0045929 | 0.00064865 | 0.0052415 |  | 1001.22 | 0.0013695
9 | NO | 100.22 | 32.6939 | 0.00055802 | -0.031666 | 0.00064257 | 0.032308 |  | 1500.4 | 0.0010681
9 | NO | 75.0778 | 32.4331 | 0.0005256 | 0.0027008 | 0.00063645 | -0.0020644 |  | 1250.36 | 0.0024376
9 | NO | 50.2436 | 32.3808 | 0.0012142 | -0.0013084 | 0.00063041 | 0.0019388 |  | 999.87 | -0.00049973
9 | NO | 30.2145 | 32.2903 | 0.0023368 | -0.011879 | 0.00062553 | 0.012505 |  | 1499.59 | 0.0013542
9 | NO | 19.9512 | 32.2755 | 0.00067805 | 0.0055618 | 0.00062303 | -0.0049388 |  | 1251.05 | 0.00079346
9 | NO | 9.9804 | 32.1968 | 0.0015605 | -0.0029793 | 0.00062061 | 0.0035999 |  | 1000.75 | 0.0013809 | 0.00084351 | 0.0013055
9 | NO | 5.0607 | 32.1943 | 0.00037329 | -0.0010261 | 0.00061941 | 0.0016456
10 | NO | 5.9902 | 32.186 | 0.00048196 | -0.0034637 | 0.00061964 | 0.0040834
11 | YES | 750.543 | 34.247 | 0.00032325 | 0.00097656 | 0.00080085 | -0.00017571
11 | NO | 4.8932 | 32.2611 | 0.00056509 | -0.00035477 | 0.00061937 | 0.00097414
12 | YES | 749.936 | 34.2523 | 0.00051134 | -0.0020905 | 0.0008007 | 0.0028912
12 | NO | 4.8232 | 32.1649 | 0.00028663 | -0.00048065 | 0.00061935 | 0.0011
13 | YES | 2214.25 | 34.6147 | 0.00025213 | 0.0029259 | 0.0011571 | -0.0017688
13 | YES | 1500.79 | 34.5095 | 0.00028943 | 0.0012436 | 0.00098345 | -0.00026014
13 | YES | 1248.57 | 34.4422 | 0.00036176 | 0.0012054 | 0.00092206 | -0.00028338
13 | YES | 1000.56 | 34.3595 | 0.00024773 | 0.0036507 | 0.0008617 | -0.002789
13 | NO | 5.0812 | 32.2066 | 0.0004392 | -0.00085068 | 0.00061942 | 0.0014701
14 | YES | 1251.85 | 34.4426 | 0.00020315 | 0.00085831 | 0.00092286 | 6.4552e-05
14 | YES | 1001.22 | 34.3659 | 0.00024413 | 0.0013695 | 0.00086186 | -0.00050762
15 | NO | 150.754 | 33.8547 | 0.0016378 | 0.0022888 | 0.00065487 | -0.0016339
16 | YES | 1500.4 | 34.5039 | 0.00024486 | 0.0010681 | 0.00098335 | -8.4768e-05
16 | YES | 1250.36 | 34.4447 | 0.00040855 | 0.0024376 | 0.0009225 | -0.0015151
16 | YES | 999.87 | 34.3639 | 0.00026983 | -0.00049973 | 0.00086153 | 0.0013613
17 | YES | 1499.59 | 34.5057 | 0.00022011 | 0.0013542 | 0.00098315 | -0.00037107
17 | YES | 1251.05 | 34.4415 | 0.0002856 | 0.00079346 | 0.00092266 | 0.00012921
17 | YES | 1000.75 | 34.3652 | 0.00036381 | 0.0013809 | 0.00086175 | -0.00051917
17 | NO | 5.0508 | 32.0353 | 0.002518 | 0.0090065 | 0.00061941 | -0.0083871
18 | NO | 4.9489 | 31.1465 | 0.00088073 | 0.15943 | 0.00061938 | -0.15881
19 | NO | 4.9585 | 25.8535 | 0.0048599 | -0.51492 | 0.00061939 | 0.51554
20 | NO | 5.2171 | 26.364 | 0.0032158 | -1.1519 | 0.00061945 | 1.1525
21 | NO | 4.9316 | 28.4434 | 0.030712 | -0.14222 | 0.00061938 | 0.14284
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.17525 |  | X | CALC Y
m= | 0.00015142 | 5.8998e-05 |  | Std Dev of Resid |  | 0.0013848 |  | 4 | -0.0091568
b= | -0.0097625 | 0.00065392 |  | Resid Sum of Squares |  | 5.9452e-05 |  | 17 | -0.0071884
# of Values Extracted = |  |  | 58
# of Values Used in Fit = |  |  | 33
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.6088 | 29.8296 | 0.013931 | -0.00059318 | -0.0096111 | -0.0090179 |  | 4 | -0.0097046
2 | NO | 5.1959 | 31.2248 | 0.0014056 | -0.029665 | -0.0094597 | 0.020205 |  | 4 | -0.010563
3 | NO | 4.9747 | 32.2356 | 0.001293 | -0.012688 | -0.0093082 | 0.0033795 |  | 4 | -0.0097847
4 | YES | 1000.62 | 34.3796 | 0.00026958 | -0.0097046 | -0.0091568 | 0.00054778 |  | 5 | -0.01083
4 | YES | 750.366 | 34.2346 | 0.00026957 | -0.010563 | -0.0091568 | 0.0014061 |  | 5 | -0.010269
4 | YES | 500.893 | 34.0617 | 0.00025059 | -0.0097847 | -0.0091568 | 0.00062789 |  | 5 | -0.0087891
5 | YES | 1251.21 | 34.4513 | 0.00036996 | -0.01083 | -0.0090054 | 0.0018245 |  | 5 | -0.0089149
5 | YES | 1000.21 | 34.3439 | 0.00064954 | -0.010269 | -0.0090054 | 0.0012638 |  | 8 | -0.010342
5 | YES | 750.528 | 34.1981 | 0.00090442 | -0.0087891 | -0.0090054 | -0.00021633 |  | 8 | -0.0085144
5 | YES | 500.444 | 34.0723 | 0.00033208 | -0.0089149 | -0.0090054 | -9.0444e-05 |  | 9 | -0.0092239
5 | NO | 5.1938 | 32.278 | 0.00029264 | -0.0095978 | -0.0090054 | 0.00059239 |  | 9 | -0.0081177
6 | NO | 5.4165 | 32.2096 | 0.00027556 | -0.010944 | -0.008854 | 0.0020904 |  | 9 | -0.0079231
7 | NO | 4.9509 | 32.1954 | 0.00030775 | -0.0086174 | -0.0087026 | -8.5153e-05 |  | 9 | -0.006649
8 | YES | 750.524 | 34.2129 | 0.00087447 | -0.010342 | -0.0085511 | 0.0017905 |  | 9 | -0.0060806
8 | YES | 501.004 | 34.0669 | 0.00030947 | -0.0085144 | -0.0085511 | -3.6733e-05 |  | 9 | -0.007061
9 | YES | 1500.04 | 34.5095 | 0.00053409 | -0.0092239 | -0.0083997 | 0.00082423 |  | 9 | -0.0077286
9 | YES | 1249.87 | 34.4549 | 0.00028318 | -0.0081177 | -0.0083997 | -0.00028203 |  | 9 | -0.0065651
9 | YES | 1000.48 | 34.3745 | 0.0002512 | -0.0079231 | -0.0083997 | -0.00047658 |  | 9 | -0.0068054
9 | YES | 750.717 | 34.2396 | 0.00031431 | -0.006649 | -0.0083997 | -0.0017507 |  | 9 | -0.0070801
9 | YES | 500.089 | 34.0838 | 0.00031774 | -0.0060806 | -0.0083997 | -0.0023191 |  | 11 | -0.0081444
9 | YES | 400.07 | 34.0201 | 0.00044288 | -0.007061 | -0.0083997 | -0.0013387 |  | 12 | -0.01144
9 | YES | 299.857 | 33.9736 | 0.00034282 | -0.0077286 | -0.0083997 | -0.00067113 |  | 13 | -0.0064201
9 | YES | 249.854 | 33.9504 | 0.00035442 | -0.0065651 | -0.0083997 | -0.0018346 |  | 13 | -0.0074806
9 | YES | 200.426 | 33.9224 | 0.0004017 | -0.0068054 | -0.0083997 | -0.0015943 |  | 13 | -0.0075226
9 | YES | 175.452 | 33.8828 | 0.00054129 | -0.0070801 | -0.0083997 | -0.0013196 |  | 13 | -0.0049934
9 | NO | 150.053 | 33.7056 | 0.001202 | -0.027187 | -0.0083997 | 0.018788 |  | 14 | -0.0085487
9 | NO | 125.184 | 33.3311 | 0.0050448 | -0.014332 | -0.0083997 | 0.0059321 |  | 14 | -0.0076828
9 | NO | 100.22 | 32.6854 | 0.00089063 | -0.040199 | -0.0083997 | 0.0318 |  | 16 | -0.0078812
9 | NO | 75.0778 | 32.4244 | 0.0008663 | -0.0060158 | -0.0083997 | -0.0023839 |  | 16 | -0.0066261
9 | NO | 50.2436 | 32.372 | 0.0010826 | -0.010052 | -0.0083997 | 0.001652 |  | 16 | -0.0093269
9 | NO | 30.2145 | 32.2824 | 0.0050227 | -0.019836 | -0.0083997 | 0.011437 |  | 17 | -0.0079498
9 | NO | 19.9512 | 32.2687 | 0.0011 | -0.0012321 | -0.0083997 | -0.0071676 |  | 17 | -0.0080833
9 | NO | 9.9804 | 32.1901 | 0.0011591 | -0.0096779 | -0.0083997 | 0.0012782 |  | 17 | -0.0076332
9 | NO | 5.0607 | 32.187 | 0.00030421 | -0.0082626 | -0.0083997 | -0.00013708
10 | NO | 5.9902 | 32.1804 | 0.00044399 | -0.0091438 | -0.0082483 | 0.00089554
11 | YES | 750.543 | 34.2379 | 0.00027364 | -0.0081444 | -0.0080969 | 4.7507e-05
11 | NO | 4.8932 | 32.2537 | 0.00044244 | -0.007782 | -0.0080969 | -0.00031489
12 | YES | 749.936 | 34.243 | 0.0003167 | -0.01144 | -0.0079455 | 0.0034948
12 | NO | 4.8232 | 32.1579 | 0.00031998 | -0.0075378 | -0.0079455 | -0.00040761
13 | YES | 2214.25 | 34.6054 | 0.00026601 | -0.0064201 | -0.007794 | -0.0013739
13 | YES | 1500.79 | 34.5008 | 0.00029056 | -0.0074806 | -0.007794 | -0.00031341
13 | YES | 1248.57 | 34.4335 | 0.00021097 | -0.0075226 | -0.007794 | -0.00027145
13 | YES | 1000.56 | 34.3508 | 0.00025555 | -0.0049934 | -0.007794 | -0.0028006
13 | NO | 5.0812 | 32.1993 | 0.0020667 | -0.0082436 | -0.007794 | 0.00044953
14 | YES | 1251.85 | 34.4332 | 0.00022122 | -0.0085487 | -0.0076426 | 0.00090613
14 | YES | 1001.22 | 34.3568 | 0.00028007 | -0.0076828 | -0.0076426 | 4.0187e-05
15 | NO | 150.754 | 33.845 | 0.0015316 | -0.0073967 | -0.0074912 | -9.4493e-05
16 | YES | 1500.4 | 34.4949 | 0.00024983 | -0.0078812 | -0.0073398 | 0.00054139
16 | YES | 1250.36 | 34.4357 | 0.0003227 | -0.0066261 | -0.0073398 | -0.00071364
16 | YES | 999.87 | 34.3551 | 0.00026929 | -0.0093269 | -0.0073398 | 0.0019872
17 | YES | 1499.59 | 34.4963 | 0.00024231 | -0.0079498 | -0.0071884 | 0.00076148
17 | YES | 1251.05 | 34.4326 | 0.00031126 | -0.0080833 | -0.0071884 | 0.00089499
17 | YES | 1000.75 | 34.3562 | 0.00039073 | -0.0076332 | -0.0071884 | 0.00044486
17 | NO | 5.0508 | 32.0272 | 0.0028602 | 0.00094986 | -0.0071884 | -0.0081382
18 | NO | 4.9489 | 31.1365 | 0.0012558 | 0.14937 | -0.0070369 | -0.15641
19 | NO | 4.9585 | 25.8467 | 0.0018722 | -0.52165 | -0.0068855 | 0.51477
20 | NO | 5.2171 | 26.3558 | 0.0022341 | -1.1601 | -0.0067341 | 1.1534
21 | NO | 4.9316 | 28.4227 | 0.028644 | -0.16286 | -0.0065827 | 0.15628
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Sal0
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.01005 |  | X | CALC Y
m= | -0.0007618 | 0.0013579 |  | Std Dev of Resid |  | 0.0015172 |  | 33.8828 | -0.0079016
b= | 0.01791 | 0.046548 |  | Resid Sum of Squares |  | 7.136e-05 |  | 34.6054 | -0.008452
# of Values Extracted = |  |  | 58
# of Values Used in Fit = |  |  | 33
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Sal0
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.6088 | 29.8296 | 0.013931 | -0.00059318 | -0.0048138 | -0.0042206 |  | 34.3796 | -0.0097046
2 | NO | 5.1959 | 31.2248 | 0.0014056 | -0.029665 | -0.0058767 | 0.023788 |  | 34.2346 | -0.010563
3 | NO | 4.9747 | 32.2356 | 0.001293 | -0.012688 | -0.0066467 | 0.006041 |  | 34.0617 | -0.0097847
4 | YES | 1000.62 | 34.3796 | 0.00026958 | -0.0097046 | -0.00828 | 0.0014246 |  | 34.4513 | -0.01083
4 | YES | 750.366 | 34.2346 | 0.00026957 | -0.010563 | -0.0081696 | 0.0023933 |  | 34.3439 | -0.010269
4 | YES | 500.893 | 34.0617 | 0.00025059 | -0.0097847 | -0.0080378 | 0.0017469 |  | 34.1981 | -0.0087891
5 | YES | 1251.21 | 34.4513 | 0.00036996 | -0.01083 | -0.0083346 | 0.0024953 |  | 34.0723 | -0.0089149
5 | YES | 1000.21 | 34.3439 | 0.00064954 | -0.010269 | -0.0082528 | 0.0020164 |  | 34.2129 | -0.010342
5 | YES | 750.528 | 34.1981 | 0.00090442 | -0.0087891 | -0.0081417 | 0.00064731 |  | 34.0669 | -0.0085144
5 | YES | 500.444 | 34.0723 | 0.00033208 | -0.0089149 | -0.0080459 | 0.00086904 |  | 34.5095 | -0.0092239
5 | NO | 5.1938 | 32.278 | 0.00029264 | -0.0095978 | -0.006679 | 0.0029188 |  | 34.4549 | -0.0081177
6 | NO | 5.4165 | 32.2096 | 0.00027556 | -0.010944 | -0.0066269 | 0.0043175 |  | 34.3745 | -0.0079231
7 | NO | 4.9509 | 32.1954 | 0.00030775 | -0.0086174 | -0.0066161 | 0.0020013 |  | 34.2396 | -0.006649
8 | YES | 750.524 | 34.2129 | 0.00087447 | -0.010342 | -0.008153 | 0.0021886 |  | 34.0838 | -0.0060806
8 | YES | 501.004 | 34.0669 | 0.00030947 | -0.0085144 | -0.0080418 | 0.0004726 |  | 34.0201 | -0.007061
9 | YES | 1500.04 | 34.5095 | 0.00053409 | -0.0092239 | -0.008379 | 0.00084497 |  | 33.9736 | -0.0077286
9 | YES | 1249.87 | 34.4549 | 0.00028318 | -0.0081177 | -0.0083374 | -0.0002197 |  | 33.9504 | -0.0065651
9 | YES | 1000.48 | 34.3745 | 0.0002512 | -0.0079231 | -0.0082761 | -0.000353 |  | 33.9224 | -0.0068054
9 | YES | 750.717 | 34.2396 | 0.00031431 | -0.006649 | -0.0081734 | -0.0015243 |  | 33.8828 | -0.0070801
9 | YES | 500.089 | 34.0838 | 0.00031774 | -0.0060806 | -0.0080547 | -0.001974 |  | 34.2379 | -0.0081444
9 | YES | 400.07 | 34.0201 | 0.00044288 | -0.007061 | -0.0080061 | -0.00094515 |  | 34.243 | -0.01144
9 | YES | 299.857 | 33.9736 | 0.00034282 | -0.0077286 | -0.0079707 | -0.00024214 |  | 34.6054 | -0.0064201
9 | YES | 249.854 | 33.9504 | 0.00035442 | -0.0065651 | -0.007953 | -0.001388 |  | 34.5008 | -0.0074806
9 | YES | 200.426 | 33.9224 | 0.0004017 | -0.0068054 | -0.0079317 | -0.0011263 |  | 34.4335 | -0.0075226
9 | YES | 175.452 | 33.8828 | 0.00054129 | -0.0070801 | -0.0079016 | -0.00082147 |  | 34.3508 | -0.0049934
9 | NO | 150.053 | 33.7056 | 0.001202 | -0.027187 | -0.0077666 | 0.019421 |  | 34.4332 | -0.0085487
9 | NO | 125.184 | 33.3311 | 0.0050448 | -0.014332 | -0.0074813 | 0.0068505 |  | 34.3568 | -0.0076828
9 | NO | 100.22 | 32.6854 | 0.00089063 | -0.040199 | -0.0069894 | 0.03321 |  | 34.4949 | -0.0078812
9 | NO | 75.0778 | 32.4244 | 0.0008663 | -0.0060158 | -0.0067905 | -0.00077476 |  | 34.4357 | -0.0066261
9 | NO | 50.2436 | 32.372 | 0.0010826 | -0.010052 | -0.0067506 | 0.0033011 |  | 34.3551 | -0.0093269
9 | NO | 30.2145 | 32.2824 | 0.0050227 | -0.019836 | -0.0066824 | 0.013154 |  | 34.4963 | -0.0079498
9 | NO | 19.9512 | 32.2687 | 0.0011 | -0.0012321 | -0.0066719 | -0.0054398 |  | 34.4326 | -0.0080833
9 | NO | 9.9804 | 32.1901 | 0.0011591 | -0.0096779 | -0.0066121 | 0.0030658 |  | 34.3562 | -0.0076332
9 | NO | 5.0607 | 32.187 | 0.00030421 | -0.0082626 | -0.0066097 | 0.0016529
10 | NO | 5.9902 | 32.1804 | 0.00044399 | -0.0091438 | -0.0066047 | 0.0025392
11 | YES | 750.543 | 34.2379 | 0.00027364 | -0.0081444 | -0.0081721 | -2.7686e-05
11 | NO | 4.8932 | 32.2537 | 0.00044244 | -0.007782 | -0.0066605 | 0.0011215
12 | YES | 749.936 | 34.243 | 0.0003167 | -0.01144 | -0.008176 | 0.0032643
12 | NO | 4.8232 | 32.1579 | 0.00031998 | -0.0075378 | -0.0065875 | 0.00095032
13 | YES | 2214.25 | 34.6054 | 0.00026601 | -0.0064201 | -0.008452 | -0.0020319
13 | YES | 1500.79 | 34.5008 | 0.00029056 | -0.0074806 | -0.0083723 | -0.00089172
13 | YES | 1248.57 | 34.4335 | 0.00021097 | -0.0075226 | -0.0083211 | -0.00079849
13 | YES | 1000.56 | 34.3508 | 0.00025555 | -0.0049934 | -0.0082581 | -0.0032646
13 | NO | 5.0812 | 32.1993 | 0.0020667 | -0.0082436 | -0.0066191 | 0.0016245
14 | YES | 1251.85 | 34.4332 | 0.00022122 | -0.0085487 | -0.0083208 | 0.0002279
14 | YES | 1001.22 | 34.3568 | 0.00028007 | -0.0076828 | -0.0082626 | -0.00057984
15 | NO | 150.754 | 33.845 | 0.0015316 | -0.0073967 | -0.0078728 | -0.00047605
16 | YES | 1500.4 | 34.4949 | 0.00024983 | -0.0078812 | -0.0083678 | -0.00048669
16 | YES | 1250.36 | 34.4357 | 0.0003227 | -0.0066261 | -0.0083227 | -0.0016966
16 | YES | 999.87 | 34.3551 | 0.00026929 | -0.0093269 | -0.0082613 | 0.0010656
17 | YES | 1499.59 | 34.4963 | 0.00024231 | -0.0079498 | -0.0083689 | -0.00041908
17 | YES | 1251.05 | 34.4326 | 0.00031126 | -0.0080833 | -0.0083204 | -0.00023705
17 | YES | 1000.75 | 34.3562 | 0.00039073 | -0.0076332 | -0.0082622 | -0.00062898
17 | NO | 5.0508 | 32.0272 | 0.0028602 | 0.00094986 | -0.006488 | -0.0074378
18 | NO | 4.9489 | 31.1365 | 0.0012558 | 0.14937 | -0.0058094 | -0.15518
19 | NO | 4.9585 | 25.8467 | 0.0018722 | -0.52165 | -0.0017797 | 0.51987
20 | NO | 5.2171 | 26.3558 | 0.0022341 | -1.1601 | -0.0021675 | 1.1579
21 | NO | 4.9316 | 28.4227 | 0.028644 | -0.16286 | -0.0037421 | 0.15912
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE