CTD Data Processing Notes

Sarah Zimmermann, IOS

02 May 2019

Updates:

2019-05-27	Added Chem salt and DIC/Alk analysis notes

2019-05-28	Added nutrients data and text

2019-10-24	Processed Dophin and Union data set.  Added processing summary for this, updated Program Summary text and in the IOS Archive CTD and Chem files.

2021-03-11	Following RBR’s fix to the Ruskin software conversion problem for Rinko data, I reprocessed Cambridge Bay and both Kugluktuk data sets. Paulatuk did not need correction. New *.CTD files made.

2021-03-16	Updated figures with those made from new data

Missing/To do:

O18 analysis text and data if/when analysis performed.

# CROW 2019-031

Contents

Program Summary:	1

Cambridge Bay	3

Kugluktuk	9

Paulatuk	12

Dolphin and Union Strait (Kugluktuk Patrol)  May 2019	19

Preparation for IOS Archive	22

Extra Notes from the field and analysis:	23

# Program Summary:

There were multiple components to this program.

Leg 1, Mike Dempsey  and Chris Clarke went to Cambridge Bay and Kugluktuk, traveling from March 5th to March 16th, 2019.

Leg 2 Mike Dempsey and Jane Eert went to Paulatuk from April 1 to April 10th, 2019

On these trips  the participants met with Rangers for training and discussion of the program in anticipation of the Ranger’s upcoming patrol exercises.   CTD casts were collected in each community.

Leg 3 The Kugluktuk patrol went out to Dolphin and Union Strait two months later and collected data. The data processing steps are discussed below.

5-11 March 2019		In Cambridge Bay (Chris and Mike)

11-16 March 2019		In Kugluktuk (Chris 11 to 14th, Mike 11 to 16th)

1-10 April 2019			In Paulatuk (Jane and Mike)

7 May 2019			Kugluktuk Patrol in Dolphin and Union Strait

For more information see the cruise report:  CROW 2019 Report final v2019-10-23.docx

Figure 1.  All stations conducted during program (2021-03-16 Oxygen corrected).

For more figures see:

2019-31 CROW data.pptx in

N:\SHARE\DATA\2019\2019-31 CROW\Data

# Cambridge Bay

CTD Data:

There are 13 stations in three data sets:

6 Mar 2019	Cambridge Bay		1 stations	065543_20190306_2159_data.txt

7 Mar 2019	Cambridge Bay		3 stations	065543_20190307_2356_data.txt

10 Mar 2019	UN Islands		4 stations	included in next file

11 Mar 2019	Dease Strait		5 stations	065543_20190311_2052_data.txt

Meta data are recorded in:

N:\SHARE\DATA\2019\2019-31_CROW\Data

Comments:

- CBay2 station shows freezing on first downcast, but the upcast and second cast are fine.

Equipment:						Calibration Dates

RBR Concerto 65543 w/ 6Hz sampling rate		25 Jun 2018 (RBR) new

w/ Turner Cyclops Fluorometer sensor #2102849	15 Feb 2018 (3pt-check, IOS), Jan 2013 (new)

Fluorometer set for 0 to 5 ug/l range

w/ JFE Advantech RINKO III Oxygen sensor #5		28 Feb 2019 (2pt-cal, IOS), 25 Jan 2009 (new)

Type-A foil, Reporting in % Saturation

Air thermometer 	SN

Ice Auger#		SN3

Downrigger#		SN6

Kit #			4

Data Collection Method:

Follow protocol laid out in Ranger science manual ver7 2015.doc

Changes to protocol:

Stations typically had a soak then two casts (RBR left in the water).  This was changed for the last five stations (March 11th) to three casts per station.

Processing:

2021-03-11 Update

RBR identified and fixed the problem in Ruskin that was calculating oxygen saturation using the fixed temperature when the CTD was turned on, not the real-time temperature.

The original *.RSK files were corrected using Ruskin 2.10.202005061945 to update the oxygen saturation value and then exported to *.TXT for further processing.

In Ruskin:

- Changed Calibration, 5: Dissolved O2 Saturation term C0 by adding 0.0000001

- Clicked “Store and recompute data”

- Export > Txt.zip

Note the *.RSK is now changed as well.

In MATLAB, existing station scan number list from original processing was used to re-create the rough mat file and then the cleaned mat files.  From here the processing step to export single *.CTD files in IOS header format w/ added note in the header was redone.

Calibration issues:

2-pt Oxygen calibration done at IOS prior to data collection.

New G and H terms are:  -3.32946282	1.181996051 (RSK C0, C1)

Casts:

Using Matlab, the time v. pressure data were plotted and each direction of each cast was separated and given a cast number.

For each station, the best cast was kept, typically the last downcast (ie second or third).

| Station | Name | Cast Number Used | File Name |
| --- | --- | --- | --- |
| 1 | CBay1 | 3 (2nd downcast) | 065543_20190306_2159 |
| 2 | West Arm | 3 (2nd downcast) | 065543_20190306_2159 |
| 3 | Cbay2 | 7 (2nd downcast) | 065543_20190306_2159 |
| 4 | Cbay6 | 11 (2nd downcast) | 065543_20190306_2159 |
| 5 | AZFP | 3 (2nd downcast) | 065543_20190311_0250 |
| 6 | UN10 | 7 (2nd downcast) | 065543_20190311_0250 |
| 7 | UN20 | 11 (2nd downcast) | 065543_20190311_0250 |
| 8 | UN30 | 15 (2nd downcast) | 065543_20190311_0250 |
| 9 | DS18-01 | 21 (3rd downcast) | 065543_20190311_0250 |
| 10 | DS18-03 | 27 (3rd downcast) | 065543_20190311_0250 |
| 11 | DS18-05 | 33 (3rd downcast) | 065543_20190311_0250 |
| 12 | DS18-04 | 39 (3rd downcast) | 065543_20190311_0250 |
| 13 | DS18-02 | 45 (3rd downcast) | 065543_20190311_0250 |

Pressure:

The top 2m of data were removed from all casts. The top 2 m are influenced by proximity to the edge of the drilled ice-hole, the mix of water and ice slush. This can create falsely fresh water but also with and cooling during the cast further and beginning to freeze and reject brine this can create a density inversion along the TS freezing line.

The bottom 0.2m were removed due to poor data resulting from change of flow rate.

Temperature shift and new Salinity:

- Applied 3 scan shift to advance temperature after confirming it was better than 2 scans.

- Smoothed temperature and conductivity to recalculate new salinity. Conductivity smoothing only applied after first 30 scans to prevent overshooting from possible large surface salinity gradients.

Fluorometer

Processed using a low pass filter (pl66.m) with a half amplitude period of 10 scans.

Oxygen (RINKO)

RBR was configured to output percent oxygen saturation (and the unused concentration [mg/l]) and not  volts.  Calibration coefficients have been installed in the CTD.

Oxygen saturation was advanced by 3 seconds (18 scans) and oxygen

concentration [mL/L] was calculated as full saturation concentration

multiplied by percent saturation (DO [ml/l] = DO at full

saturation [ml/l] * %sat). DO at full saturation was calculated using

Garcia and Gordon (1992,1993) Eq.8 using coeefficients in column 1 from

Table 1 (fit to Benson and Krause (1984) data) and in situ temperature.

The final data set has the adjusted oxygen saturation and oxygen

concentration [mL/L].

REFERENCES:

Benson, B.B., and D. Krause, 1984: The concentration and isotopic

fractionation of oxygen dissolved in freshwater and seawater in

equilibrium with the atmosphere. Limnology and Oceanography, 29,

620-632.

Garcia, H.E., and L.I. Gordon, 1992: Oxygen solubility in seawater:

Better fitting equations. Limnology and Oceanography, 37, 1307-1312.

Garcia, H.E., and L.I. Gordon, 1993: Erratum: Oxygen solubility in

seawater: better fitting equations. Limnology and Oceanography, 38,

656.

Removed all pressure reversals

Interpolated to 0.2db bins.  These are NOT centered.

Special edits:

Station		Edit

None

Saved this cleaned dataset to  matlab *.mat file and an ODV *.o4x file

Used script file p_RBR_CROW_2019_CB1.m for processing stations 1.

Used script file p_RBR_CROW_2019_CB2.m for processing stations 2 to 4.

Used script file p_RBR_CROW_2019_CB3.m for processing stations 5 to 13

Science:

No chlorophyll signal yet.

Suggestions:

Figures:

Figure 2. All stations with Cambridge Bay patrol

Figure 3. Just Dease St. and to the north

Figure 4.  Dease St. Section

Figure 5. Just Cambridge Bay before oxygen correction

# Kugluktuk

CTD Data:

There are 3 stations in 1 data set:

13 Mar 2019	Kugluktuk Bay	3 stations	065579_20190313_2052.rsk

Meta data are recorded in:

N:\SHARE\DATA\2019\2019-31_CROW\Data

Comments:

It looks like the sensor may have frozen at the surface before each downcast.  The CTD may be freezing a film of fresh water over the sensors as it is measuring the surface 2 to 3 m of fresh water.

Equipment:						Calibration Dates

RBR Concerto 65579 w/ 6Hz sampling rate		02 May 2016 (RBR), May 2013 (new)

w/ Turner Cyclops Fluorometer sensor #2102820	15 Feb  2018 (3pt check),  Jan 2012 (new)

Fluorometer set for 0 to 5 ug/l range

w/ JFE Advantech RINKO III Oxygen sensor #323 		1-Mar-2018  (2pt cal at IOS),

Type-B foil, Reporting in % Saturation			10 Aug 2017 (new)

Air thermometer 	SN

Ice Auger		SN

Downrigger		#8

Kit			#1

Data Collection Method:

Follow protocol laid out in Ranger science manual ver7 2015.doc

Changes to protocol:  Change from 2 to 3 casts for each station.

Processing:

2021-03-11 Update

RBR identified and fixed the problem in Ruskin that was calculating oxygen saturation using the fixed temperature when the CTD was turned on, not the real-time temperature.

The original *.RSK files were corrected using Ruskin 2.10.202005061945 to update the oxygen saturation value and then exported to *.TXT for further processing.

In Ruskin:

- Changed Calibration, 5: Dissolved O2 Saturation term C0  and C1 terms to coeffiecints from 1 Mar 2018 calibration.

- Corrected Calibration, 5: Dissolved O2 Saturation X6 and X7 terms to 4.166 and 0

- Clicked “Store and recompute data”

- Export > Txt.zip

Note the *.RSK is now changed as well.

In MATLAB, existing station scan number list from original processing was used to re-create the rough mat file and then the cleaned mat files.  From here the processing step to export single *.CTD files in IOS header format w/ added note in the header was redone.

To confirm the correction dealt with the X6 and X7 terms correctly, I processed the 13 March file using the Legacy export of voltage using Ruskin v2.9, and calculated %sat and DO from the voltage and got the same values.

Also, although the Rinko was calibrated in the field during this trip (12 Mar 2019)  the values are compromised due to the Ruskin software and X6 and X7 incorrect settings. The previous March 2018 coefficients agree well with those collected in Oct 2020.  The Mar 2018 calibration will be used to scale the data.

Calibration issues:

2-pt Oxygen calibration was done in Kugluktuk prior to data collection however due to Ruskin software problem and incorrect X6 and X7 coefficients the resulting coefficients are suspect and the previous coefficients from 1 Mar 2018 will be used.

Casts:

Using Matlab, the time v. pressure data were plotted and each direction of each cast was separated and given a cast number.

For each station, the best cast was kept:

| Station | Name | Cast Number Used | File Name |
| --- | --- | --- | --- |
| 14 | KUG-1 | 5 (3rd downcast) | 065579_20190313_2052 |
| 15 | KUG-2 | 11 (3rd downcast) | 065579_20190313_2052 |
| 16 | KUG-3 | 17 (3rd downcast) | 065579_20190313_2052 |

Pressure:

The top 2m of data were removed from all casts. The top 2 m are influenced by proximity to the edge of the drilled ice-hole, the mix of water and ice slush. This can create falsely fresh water but also with cooling during the cast, ice formation and brine rejection  there can be a manufactured density inversion along the TS freezing line.

The bottom 0.2m of data removed from all casts. The very bottom data tends to have density inversion and/or spikes. This removes very little data but cleans up the profile.

Temperature shift and new Salinity:

- Applied a 3 scan shift to advance temperature after confirming this was better than 2 or 4 scans.

- Smoothed temperature and conductivity to recalculate new salinity.  However, due to large surface gradient in salinity, the original conductivity was used for the first 30 scans as the filtered conductivity had artificial spikes due to the filtering ‘overshoot’.

Fluorometer

Processed using a low pass filter (pl66) with a half amplitude period of 10 scans.

Oxygen

The RBR CTD was configured to output percent oxygen saturation (and the unused concentration [mg/l]) from the oxygen sensor and not volts.  Calibration coefficients have been installed in the CTD.  Coefficients G and H, the ones determined with the regular 2-point calibration, are visible as C0 and C1.

Oxygen saturation was advanced by 3 seconds (18 scans) and oxygen

concentration [mL/L] was calculated as full saturation concentration

multiplied by percent saturation (DO [ml/l] = DO at full

saturation [ml/l] * %sat). DO at full saturation was calculated using

Garcia and Gordon (1992,1993) Eq.8 using coeefficients in column 1 from

Table 1 (fit to Benson and Krause (1984) data) and in situ temperature.

The final data set has the adjusted oxygen saturation and oxygen

concentration [mL/L].

REFERENCES:

Benson, B.B., and D. Krause, 1984: The concentration and isotopic

fractionation of oxygen dissolved in freshwater and seawater in

equilibrium with the atmosphere. Limnology and Oceanography, 29,

620-632.

Garcia, H.E., and L.I. Gordon, 1992: Oxygen solubility in seawater:

Better fitting equations. Limnology and Oceanography, 37, 1307-1312.

Garcia, H.E., and L.I. Gordon, 1993: Erratum: Oxygen solubility in

seawater: better fitting equations. Limnology and Oceanography, 38,

656.

Salinity

No edits

Removed pressure reversals

Interpolated to 0.2db bins, however these are NOT centered.

Saved this cleaned dataset to a matlab *.mat files and an ODV *.o4x file

2019-31_CROW_Kug_cleandowncasts.mat and data_from_ 2019-31_CROW_Kug.txt

Used script file p_RBR_CROW_2019_Kug.m for processing and p_RBR_CROW_2019_KugRawV.m to confirm fixed RSK file was correct.

Science:

Clearly fresh river water was capping the 3 stations.

Suggestions:

Figures:

Figure 6.  All stations w/ Kugluktuk patrol

Figure 7.  Kugluktuk stations showing the top 10m

# Paulatuk

CTD Data:

There are 8 stations, all with 3 or more casts per station in 2 data sets:

5 Apr 2019  Paulatuk, South West Darnley Bay	6 stations	065541_20190405_2348.rsk

7 Apr 2019  Paulatuk, Darnley Bay		2 stations	065639_20190408_0132.rsk

Both stations on 7 April have water samples

Meta data are recorded in:

N:\SHARE\DATA\2019\2019-31_CROW\Data

Comments:

The first CTD used had problems with external channels and there are no good oxygen and chlorophyll data for stations 1 to 6.

The next field-day used a different CTD with different external sensors.  The second CTD had a few dropouts with these external channels at the start of the first cast, but as there were 3 or more casts per station, there was no problem in choosing a good cast without dropouts.

Equipment:						Calibration Dates

RBR Concerto 65541 w/ 6Hz sampling rate		12 Mar 2018, Jan2013 (new)

w/ Turner Cyclops Fluorometer sensor #21100027	15 Feb 2018 (3-pt check, IOS), Feb 2017 (new)

w/ JFE Advantech RINKO III Oxygen sensor #9 		04-Apr-2019  (2pt cal, in field),

Type- Afoil, Reporting in % Saturation			01 Jun 2009 (new)

RBR Concerto 65639 w/ 6Hz sampling rate		27 Mar 2018, Mar2014 (new)

w/ Turner Cyclops Fluorometer sensor #21100322	15 Feb 2018 (3-pt check, IOS), May 2017 (new)

Fluorometer set for 0 to 5 ug/l range

w/ JFE Advantech RINKO III Oxygen sensor #259 		1 Mar 2018 (2pt cal, IOS), 24 May 2016 (new)

Type-B foil, Reporting in Volts				9 Apr 2019 (2pt cal, field)  confirmed the

current values were good w/in 1 or 2%

Air thermometer 	SN

Ice Auger#		SN

Downrigger#		SN

Kit #			SN

Data Collection Method:

Follow protocol laid out in Ranger science manual ver7 2015.doc

Changes to protocol:

Added water sample collection by hanging one Niskin on the line and closing with a messenger.

Performed 3 casts per station instead of 2.

Processing:

2021-03-11 Update:  The problem with the Ruskin software does not affect these data and no re-processing was needed.

Calibration issues:

RINKO III SN9 had a 2-pt oxygen calibration done in Paulatuk prior to data collection.

New G and H terms are: -3.530317,1.032586 (RSK terms C0 and C1)

However, there was some failure with this system and no oxygen data were collected.

RINKO III SN259 also had a 2-pt oxygen calibration the day after it was used for data collection.  The calibration agreed within 1 to 2% of the prior (1 Mar 2018) calibration so no changes were made.

Casts:

Using Matlab, the time v. pressure data were plotted and each direction of each cast was separated and given a cast number.

For each station, the best cast was kept which was typically the last downcast (i.e. the third)

List of Stations

| Station | Name | Cast Number Used | File Name |
| --- | --- | --- | --- |
| 17 | CMPA2 | 3 (2nd downcast) | 065541_20190405_2348 |
| 18 | CMPA1 | 13 (3rd downcast) | 065541_20190405_2348 |
| 19 | CMPA4 | 19 (3rd downcast) | 065541_20190405_2348 |
| 20 | CMPA3 | 25 (3rd downcast) | 065541_20190405_2348 |
| 21 | CMPA5 | 29 (2nd downcast - deeper than next 2 casts) | 065541_20190405_2348 |
| 22 | CMPA6 | 39 (3rd downcast) | 065541_20190405_2348 |
| 23 | BPT_01 | 9 (4th downcast) | 065639_20190408_0132 |
| 24 | BPT_HC2 | 21 (3rd downcast) | 065639_20190408_0132 |

Pressure:

The top 2m of data were removed from all casts. The top 2 m are influenced by proximity to the edge of the drilled ice-hole with the mix of water and ice slush. This water mix can create falsely fresh water, also cooling during the cast can freeze surface water and reject brine creating density inversions along the TS freezing line.

The bottom 0.2m were removed due to poor data resulting from change of flow rate.

Temperature shift and new Salinity:

- Applied 3 scan shift to advance temperature after confirming it was better than 2 scans.

- Smoothed temperature and conductivity conductivity using a low pass filter (pl66) with a half amplitude period of 6 scans to recalculate new salinity.  The conductivity smoothing was only applied after first 30 scans to prevent overshooting from possible large surface salinity gradients.

Fluorometer

Missing for first 6 stations due to CTD problems.

For the last two stations, processed using a low pass filter (pl66) with a half amplitude period of 10 scans.

Oxygen (RINKO)

Missing for first 6 stations due to CTD problems.

For the last two stations:

The RBR CTD was configured to output volts.  The data were advanced by 3 seconds (18 scans).  This is roughly a 1.5m offset at a typical lowering (and raising) speed of 0.5m/sec.  This brings the oxygen data in line with temperature.

Factory coefficients A through F and the two determined from the 2-pt calibrations, coefficients G and H, were used to calculate oxygen saturation using RINKO’s provided script for Type-B foil sensors.

Oxygen concentration [mL/L] was calculated as full saturation concentration

multiplied by percent saturation ( DO [ml/l] = DO at full

saturation [ml/l] * %sat). DO at full saturation was calculated using

Garcia and Gordon (1992,1993) Eq.8 using coeefficients in column 1 from

Table 1 (fit to Benson and Krause (1984) data) and in situ temperature.

The final data set has the adjusted oxygen saturation and oxygen

concentration [mL/L].

REFERENCES:

Benson, B.B., and D. Krause, 1984: The concentration and isotopic

fractionation of oxygen dissolved in freshwater and seawater in

equilibrium with the atmosphere. Limnology and Oceanography, 29,

620-632.

Garcia, H.E., and L.I. Gordon, 1992: Oxygen solubility in seawater:

Better fitting equations. Limnology and Oceanography, 37, 1307-1312.

Garcia, H.E., and L.I. Gordon, 1993: Erratum: Oxygen solubility in

seawater: better fitting equations. Limnology and Oceanography, 38,

656.

Removed all pressure reversals

Interpolated to 0.2db bins.

Special edits:

Station		Edit

NA		NA

Saved this cleaned dataset to a matlab *.mat files and an ODV *.o4x file

2019-31_CROW_Paulatuk1_cleandowncasts.mat

2019-31_CROW Paulatuk2_cleandowncasts.mat

data_from_2019-31_CROW_Paulatuk1.txt

data_from_2019-31_CROW_Paulatuk2.txt

Used script file p_RBR_CROW_2019_Paul1.m for processing. Stations 1 to 6

Used script file p_RBR_CROW_2019_Paul2.m for processing. Stations 7 to 8

These both call p_RBR_generic_2019.m that uses ‘readtable’ to import data.

Bottle data:

At the last two stations, BPT_01, BPT_HC2, water samples were taken at 4 and 3 depths respectively.  Depths were essentially at bottom, middle and near surface.

Samples were taken for Nutrients (Nitrate+Nitrate (NO3), Silicate(SiO4), Phosphate (PO4)), Dissolved Inorganic Carbon (DIC), Alkalinity, Oxygen Isotope (δ18O), and Salinity.  All samples were collected in duplicate.  The duplicate set of salinities may be used for other analysis such as Strontium.

A single 5L Niskin was attached to the line above the RBR CTD.  The CTD was lowered to the desired depth, a “messenger” was attached to the line and released, closing the Niskin.  The line was pulled up and the Niskin removed and sampled before being put back on the line for the next sample.  Bottle depth was determined to be maximum CTD depth minus 1.5 m to account for the distance between Niskin and CTD sensors.  The CTD data corresponding to the bottle depth was recorded in a spreadsheet along with the water sample chemistry.

Nutrients:

Nutrient samples were collected, unfiltered, in new 15 mL polystyrene tubes after rinsing 3 times with sample water and then frozen until analysis.  Samples accidentally thawed during the flight back to the lab at IOS (1 day) and were refrozen at -20C until analysis.  The samples had been with cold packs that had remained frozen so its expected the samples had not warmed much or for very long. Samples thawed for 2 hours at ~28C and let equilibrate to room temperature before analysis 14 May 2019 (5 week after collection).  Analysis was performed on a Seal Analytical AA3 with corrections made for salt effects and to the certified reference material (CRM).  CRM used was Kanso Lot CA and CD.  All samples were taken in duplicate however one tube cracked (sample #5) and was not used.  Samples had a precision, calculated as the Standard Pool, of:

NO3: 0.18 µM (n = 6, min = 5.34 µM, max = 8.65 µM)

SiO4:  0.37 µM (n = 6, min = 11.82µM, max = 18.81 µM)

PO4:  0.023 µM (n = 6, min = 1.093 µM, max =1.307 µM)

Salinity:

Salinity samples were collected into 200 mL type II glass bottles with screw caps and

disposable plastic inserts after rinsing 3 times with sample water.   Samples were kept unfrozen and brought back to IOS for analysis.  Samples were  analyzed May 15, 2019 (5 weeks after collection) in a temperature-controlled lab on the Guildline AutoSalinometer Model 8400B (SN: 69086), which  was standardized with IAPSO standard seawater, Batch #162, exp 20 May 2022.  No replicates as second set of salinities may be run for a separate analysis.  Removing 1 outlier, comparison with ctd (n=6):  mean -0.002 PSU, standard deviation 0.004 PSU.

O18 Not analysed yet, below is draft text.

O18:

Oxygen Isotopes Samples were collected into 30 ml glass vials after rinsing 3 times with sample water. Once at room temperature, the caps were retightened and kept unfrozen until analysis.  Samples were analyzed at Oregon State University using the Thermo DeltaPlusXL mass spectrometer connected to a H2O-CO2 equilibration unit. Samples were analyzed June to July 2010 (10 to 11 months after collection).

DIC and Alkalinity

DIC and Alkalinity were analysed out of the same sample.  A glass 330ml (beer) bottle was rinsed three times and filled, head space given,  poisoned with about 70ug powdered mercuric chloride and sealed  with a metal crimp cap.  Samples were kept cool (4C) until analysis at the Institute of Ocean Sciences.  DIC was analysed using the VINDTA 84, a coulometer system.  Dickson CRM batch 178 standard water was used.  Alkalinity samples were analyzed using an automated potentiometric titration system with an open cell type and non-linear least squares endpoint determination for the total alkalinity measurement.  Dickson CRM batch 173 standard water was used.  Samples were analysed May 2, 2019, less than 4 weeks after collection.  All samples had been collected in replicate.

Standard deviation of replicates after removing questionable sample #6:

DIC: 	1.05umol/kg, n=6

Alkalinity: 1.60 umol/kg, n=6

Table 1. CTD data at bottle stops

| yyyy-mm-dd hh:mm | Station and Cast | Sample No. | CTDPres [dbar] | CTDTemp [ITS-90 C] | CTDCond [mS/cm] | CTDSalt [PSS-78] | CTDOxy [mL/L] | CTDFluo [mg/m3] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2019-04-07 | BPT-01 Cast 23 | 1 | 68.4 | -1.671 | 25.665 | 32.254 | 7.65 | 0.45 |
| 2019-04-07 | BPT-01 Cast 23 | 2 | 41.6 | -1.668 | 25.655 | 32.253 | 7.64 | 0.47 |
| 2019-04-07 | BPT-01 Cast 23 | 3 | 12 | -1.682 | 25.605 | 32.216 | 7.80 | 0.71 |
| 2019-04-07 | BPT-01 Cast 23 | 4 | 3.8 | -1.742 | 25.496 | 32.135 | 8.27 | 1.58 |
| 2019-04-07 | BPT-HC2 Cast 24 | 5 | 91.2 | -1.640 | 25.742 | 32.313 | 7.47 | 0.24 |
| 2019-04-07 | BPT-HC2 Cast 24 | 6 | 21.4 | -1.735 | 25.490 | 32.110 | 8.04 | 0.79 |
| 2019-04-07 | BPT-HC2 Cast 24 | 7 | 2.6 | -1.721 | 25.235 | 31.754 | 8.11 | 0.87 |

| yyyy-mm-dd hh:mm | Station and Cast | Sample No. | CTDPres [dbar] | Salt | QF | NO3 [mmol/m3] | QF | SiO4 [mmol/m3] | QF | PO4 [mmol/m3] | QF | DIC [umol/kg] | QF | Alk [umol/kg] | QF | O18 [‰ VSMOW] | QF |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2019-04-07 | BPT-01 Cast 23 | 1 | 68.4 | 32.26 |  | 7.80 | 6 | 17.38 | 6 | 1.282 | 6 | 2201.06 | 6 | 2278.50 | 6.00 |  |  |
| 2019-04-07 | BPT-01 Cast 23 | 2 | 41.6 | 32.26 |  | 7.55 | 6 | 16.89 | 6 | 1.255 | 6 | 2200.06 | 6 | 2279.03 | 6.00 |  |  |
| 2019-04-07 | BPT-01 Cast 23 | 3 | 12 | 32.21 |  | 6.88 | 6 | 15.47 | 6 | 1.212 | 6 | 2196.77 | 6 | 2269.37 | 6.00 |  |  |
| 2019-04-07 | BPT-01 Cast 23 | 4 | 3.8 | 32.14 |  | 5.34 | 6 | 12.28 | 6 | 1.093 | 6 | 2189.16 | 6 | 2288.52 | 6.00 |  |  |
| 2019-04-07 | BPT-HC2 Cast 24 | 5 | 91.2 | 32.32 |  | 8.65 |  | 18.81 |  | 1.307 |  | 2202.62 | 6 | 2269.48 | 6.00 |  |  |
| 2019-04-07 | BPT-HC2 Cast 24 | 6 | 21.4 | 32.11 |  | 6.10 | 6 | 13.44 | 6 | 1.147 | 6 | 2194.23 | 6,3 | 2375.61 | 6,3 |  |  |
| 2019-04-07 | BPT-HC2 Cast 24 | 7 | 2.6 | 31.83 |  | 5.47 | 6 | 11.82 | 6 | 1.101 | 6 | 2155.04 | 6 | 2251.59 | 6.00 |  |  |

Table 2. Chemistry at bottle  stops.

Science:

Suggestions:

Figures:

Figure 8. In Paulatuk’s Darnley Bay.

# Dolphin and Union Strait (Kugluktuk Patrol)  May 2019

CTD Data:

There are 3 stations in 1 data set:

7 May 2019	Dolphin and Union Strait	3 stations	065579_20190910_2145.rsk

Meta data are recorded in:

N:\SHARE\DATA\2019\2019-31_CROW\Data

Comments:

The Kugluktuk Rangers went out in May and collected 3 stations across Dolphin and Union Strait.

Equipment:						Calibration Dates

RBR Concerto 65579 w/ 6Hz sampling rate		02 May 2016 (RBR), May 2013 (new)

w/ Turner Cyclops Fluorometer sensor #2102820	15 Feb  2018 (3pt check),  Jan 2012 (new)

Fluorometer set for 0 to 5 ug/l range

w/ JFE Advantech RINKO III Oxygen sensor #323 		1-Mar-2018  (2pt cal at IOS),

Type-B foil, Reporting in % Saturation			10 Aug 2017 (new)

Air thermometer 	SN

Ice Auger		SN

Downrigger		#8

Kit			#1

Data Collection Method:

Follow protocol laid out in Ranger science manual ver7 2015.doc

Changes to protocol:  Had 2 casts for each station.

Processing:

2021-03-11 Update

RBR identified and fixed the problem in Ruskin that was calculating oxygen saturation using the fixed temperature when the CTD was turned on, not the real-time temperature.

The original *.RSK files were corrected using Ruskin 2.10.202005061945 to update the oxygen saturation value and then exported to *.TXT for further processing.

In Ruskin:

- Changed Calibration, 5: Dissolved O2 Saturation term C0  and C1 terms to coeffiecints from 1 Mar 2018 calibration.

- Corrected Calibration, 5: Dissolved O2 Saturation X6 and X7 terms to 4.166 and 0

- Clicked “Store and recompute data”

- Export > Txt.zip

Note the *.RSK is now changed as well.

In MATLAB, existing station scan number list from original processing was used to re-create the rough mat file and then the cleaned mat files.  From here the processing step to export single *.CTD files in IOS header format w/ added note in the header was redone.

Calibration issues:

2-pt Oxygen calibration was done in Kugluktuk prior to data collection however due to Ruskin software problem and incorrect X6 and X7 coefficients the resulting coefficients are suspect and the previous coefficients from 1 Mar 2018 were used.

Casts:

Using Matlab, the time v. pressure data were plotted and each direction of each cast was separated and given a cast number.

For each station, the best cast was kept:

| Station | Name | Cast Number Used | File Name |
| --- | --- | --- | --- |
| 25 | DU1_2019 | 3 (2nd downcast) | 065579_20190910_2145 |
| 26 | DU2_2019 | 7 (2nd downcast) | 065579_20190910_2145 |
| 27 | DU3_2019 | 11 (2nd downcast) | 065579_20190910_2145 |

Pressure:

The top 2m of data were removed from all casts. The top 2 m are influenced by proximity to the edge of the drilled ice-hole, the mix of water and ice slush. This can create falsely fresh water but also with cooling during the cast, ice formation and brine rejection  there can be a manufactured density inversion along the TS freezing line.

The bottom 0.2m of data removed from all casts. The very bottom data tends to have density inversion and/or spikes. This removes very little data but cleans up the profile.

Temperature shift and new Salinity:

- Applied a 3 scan shift to advance temperature after confirming this was better than 2 or 4 scans.

- Smoothed temperature and conductivity to recalculate new salinity.  However, due to large surface gradient in salinity, the original conductivity was used for the first 30 scans as the filtered conductivity had artificial spikes due to the filtering ‘overshoot’.

Fluorometer

Processed using a low pass filter (pl66) with a half amplitude period of 10 scans.

DU2, Cast 7, interpolated chlorophyll data from 53 to 57 dbar to remove spike.

Oxygen

The RBR CTD was configured to output percent oxygen saturation (and the unused concentration [mg/l]) from the oxygen sensor and not volts.  Calibration coefficients have been installed in the CTD.  Coefficients G and H, the ones determined with the regular 2-point calibration, are visible as C0 and C1.

Oxygen saturation was advanced by 3 seconds (18 scans) and oxygen

concentration [mL/L] was calculated as full saturation concentration

multiplied by percent saturation (DO [ml/l] = DO at full

saturation [ml/l] * %sat). DO at full saturation was calculated using

Garcia and Gordon (1992,1993) Eq.8 using coeefficients in column 1 from

Table 1 (fit to Benson and Krause (1984) data) and in situ temperature.

The final data set has the adjusted oxygen saturation and oxygen

concentration [mL/L].

REFERENCES:

Benson, B.B., and D. Krause, 1984: The concentration and isotopic

fractionation of oxygen dissolved in freshwater and seawater in

equilibrium with the atmosphere. Limnology and Oceanography, 29,

620-632.

Garcia, H.E., and L.I. Gordon, 1992: Oxygen solubility in seawater:

Better fitting equations. Limnology and Oceanography, 37, 1307-1312.

Garcia, H.E., and L.I. Gordon, 1993: Erratum: Oxygen solubility in

seawater: better fitting equations. Limnology and Oceanography, 38,

656.

Salinity

No edits

Removed pressure reversals

Interpolated to 0.2db bins, however these are NOT centered.

Saved this cleaned dataset to a matlab *.mat files and an ODV *.o4x file

2019-31_CROW_KugDU_cleandowncasts.mat and data_from_ 2019-31_CROW_Kug.txt

Used script file p_RBR_CROW_2019_Kug2.m for processing.

Science:

Suggestions:

Figures:

Figure 9. Dolphin and Union stations w/ Kugluktuk patrol

# Preparation for IOS Archive

Use matlab scripts from Roy Hourston to convert *.mat files to IOS header format:

C:\Users\Zimmermanns\Documents\Gmatlab\IOSarchivescripts

C:\Users\zimmermannS\Documents\Gmatlab\fromRoyHourston

Use prntios_fmts_201931CROW_wrapper.m to cycle through the various CTD data sets within 2019-31 (ie Cambridge Bay, Kugluktuk, Paulatuk).  This script has a list of the needed variables and file names  and then calls the function prntios_fmts_201931CROW_single.m to write the CTD files.

Using prntios_fmts_201931CROW_single.m

Keep all the variables up at the top to make it easy to change for each data set

Calls (input)

*.mat	The processed data in structured array with one cast per station (ie keep best dip)

*.txt	The processing information that will be loaded into the IOS Shell header

*.xls	Create a page that is easy to pull sensor information from (and clear for Mike to 			write pre-CROW)

Writes one file per cast as is custom for IOS header format files.

Variables to archive:

Pressure, Depth, Temperature, Conductivity, Salinity, Oxygen %sat, Oxygen ml/l, Fluorometer

0.2 dbar intervals

Tried to follow standard IOS variable names, units and sig figs

Bottom depth was calculated as the max depth plus 5m based on typical cast protocol.

Cruise name changed from 2019-31 to 2019-031

The chemistry has been converted to CHE files and has water sample data as they are available.

For more tips see  How to Archive CROW data v2019-04-29.docx (in G:\Sorted\RBR_Processing\How to put in IOS Archive)

# Extra Notes from the field and analysis:

Follow the oxygen processing trail…

- Do g and h act on percent oxygen saturation and what is the formula?  YES

- Is RBR applying g and h to the percent oxygen saturation or only when calculating ml/l? BOTH

- Can I get the same numbers or is something quite diff when I’m using the Garcia and Gordon formula? YES it’s the same (see:   G:\Sorted\myCROW\Data\RINKOCalTest)

How does Ruskin deli er oxy sat:

Cal terms in the CTD, Rinko is putting out volts and CTD converts.

CTD is set up for a specific RINKO sensor w/ its cal info.

If cal terms are in the RBR CTDs can Rinko be swapped btw CTDs?-NO SWAPPING w/out changing ALL Coef (Xs and C0 cnc C1)

What happens when Type A Rinko and Type B Rinko are swapped? – PROBLEMs Send your data to RBR and see if they can correct this.

From: Dempsey, Mike <Mike.Dempsey@dfo-mpo.gc.ca> 
Sent: 2019–April-08 8:13 PM
To: Zimmermann, Sarah <Sarah.Zimmermann@dfo-mpo.gc.ca>
Cc: Eert, Jane <Jane.Eert@dfo-mpo.gc.ca>
Subject: RE: CROW log 2019

Hi Sarah,

Some background for the data files in Dropbox.

We had issues with CTD 65541. It was sent to RBR last year for updating to logger 3 mod. Unfortunately it was sent back with the default period set to 11 s. I didn't catch that at IOS before it was sent to Paulatuk last June. ( we had a similar experience in Kugluktuk with a new logger 3 mod. Heidi Swanson was using the CTD and it was stuck at 6s. Ryan Flagg was talking with her there and said he had seen the same problem and set up a Go To meeting with RBR to fix it.) The first file for 65541 which was Darcy Mcnicholl's winter sampling was done at this rate. There are also error7's ( no response received) on channel 4 (Chl a). The CTD had been set up for auto range on the fluorometer which I assumed caused the errors.

I had a Go To meeting with RBR and they set the rate option working again (it was greyed out). I tested it on the bench and did a fresh cal on the Rinko. It appeared to be ok so we used it on 5 April for CMPA1-6. The downloaded data has drop outs on the external channels 4 & 5 (fluor and DO).

we had brought CTD 65639 as a spare. It had been set up as a fast Fl/v with voltage only on channel 5.I tried to get RBR to change ch 5 to DO but they said that it would be a $1500 upgrade, so we used it with channel 5 as voltage only. As a result of this, we used 65639 with Rinko 259 and cyclops 32 2 on 7 April to BPT_01  and BPT_HC2.

RBR has since sent me an email saying they will modify all our remaining Concertos that have ch5 as voltage to DO at no cost as a "goodwill gesture" but will charge for any new orders. We have a Concerto at RBR right now that is being upgraded to logger 3 mod and it will have channel 5 changed to Rinko. Tomorrow early I have arranged a GoTo meeting to reset ch 5 to Rinko. After the GoTo meeting, I ill run a field cal in a bucket to calculate the latest G and H coefficients for Rinko 259.

The data from 7 April also has small drop outs on channel 4 & 5. I am bringing 65541 back for evaluation and repair. Depending on what we find out (possible external voltage card , firmware , other h/w problems, or???), we can swap out 65639 in Paulatuk with a fresh CTD when one is available.

Always interesting.

Thanks.

Mike