## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2019-038

Agency: OSD

Location: Baynes Sound

Project: Coastal Ecosystem Research Program

Party Chief: Sutherland T.

Platform: Vector

Date: 13 April 2019 – 16 April 2019

Processed by: Germaine Gatien

Date of Processing: 26 June 2019 – 20 August 2019

Number of original HEX files:  	9		Number of CTD files:  9

Number of original ROS files: 	4		Number of CHE files:  4

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (3685), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (692) and an altimeter (43281). A Biosperical/Licor Surface PAR (#20518) was mounted on the boat deck.

Seasave version 7.23.2 was used for CTD acquisition and 7.26.7.110 was used for TSG acquisition.

The data logging computer was water Properties #32.

The deck unit was a Seabird model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model,  8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) Kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

This cruise immediately followed 2019-016 and used the same equipment. As there was more information available for the earlier cruise the same parameters were chosen for processing steps.

The Daily Science Log entries are included in the same book as 2019-016 though copies have been made of the relevant pages to go in the folder for this cruise. Rosette sampling logs were available.

The event numbers were restarted after a series of grab samples. The normal protocol is to use a continuous set of event numbers for all logged activities.

There is confusion about which bottles were sampled during event #6. The log indicates that 6 bottles were fired, with Niskin #3 fired twice. In fact 7 Niskin bottles were fired. Based on a comparison with data from the CTD DO sensor it is believed that there was no sample taken from Niskin #4.

One salinity sample was mislabelled as #5. The log indicates that it was originally planned as #5 but changed to #1. It is clear in comparison to CTD data that it was not a surface sample.

This cruise was not well suited to calibration sampling since casts were all in shallow water in an area with large vertical gradients. The comparison with bottle salinity and dissolved oxygen showed evidence of incomplete flushing of the bottles. Dissolved oxygen values were in a very narrow range making a reliable fit impossible. Recalibration of salinity and dissolved oxygen were based on the results of cruise 2019-016 which immediately preceded this cruise and used the same equipment.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.50 mL/L from 0 to 20db

±0.20 mL/L from 20 to 60db

The pH:SBE:Nominal data should be used with caution; no field calibration data are available. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained. There was some confusion in both.

- The event numbers  in the Daily Science Log were restarted after the GRAB casts, while the first rosette cast was called #32 in the Rosette log and #2 in the Daily Science Log. The numbers used in the Daily Log book will be used for the processed files.

- The log notes that during event #6 Niskin Bottle #3 was fired twice at 20m. In fact, that is not what happened – bottles 3 and 4 were fired at 20m. When the bottle file is prepared, care will be needed to assure that the DO samples are assigned to the right bottles as those gathering samples may not have understood what happened.

- Dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the dissolved oxygen , temperature, conductivity and pressure sensors was obtained.

- The configuration files used at sea had an error in the transmissivity serial number but the parameters were correct. The dissolved oxygen parameter E was change to 0.00365 based on tests run during 2019-038. This is of no importance for this cruise since all casts are shallow, but was changed to avoid confusion in future. It was saved as 2019-038-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2019-038-ctd.xmlcon.

The ROS files were converted to IOS format. They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files and no significant outliers were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. There is some confusion over event #6 where 2 bottles were fired at 20m. The log has a note that there were 2 firings at 20m and indicates that samples 16, 17 and 18 came from Niskins 4, 5 and 6. To get sampling from the correct depths those samples should have come from Niskins 5, 6 and 7. Samples 15 and 16 are not as close as expected if they both came from 20m, so it is assumed that Niskin #4 was not sampled. No sample number was assigned to Niskin #4. (This was checked after the bottle files were put together and the choice made looks appropriate, though far from certain so 3 flags were later added to the 3 samples of uncertain origin and the bottle file was then rebuilt.)

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-038-bot-hdr.txt which will be updated as needed during processing. (* indicates date of creation for the file.)

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2019-038oxy.xlsx which includes flags and comments. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-038oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file QF2019-038 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab 100 to 101 days after collection. The files were simplified and saved as 2019-038sal.csv.

That file was then converted to individual SAL files.

The SAL and OXY files were merged with CST files in 2 steps.

After the 2nd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were compared to the rosette log sheets to look for omissions. None were found.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were only 4 bottles and 2 had very noisy CTD salinity data during the 10s window. Using the 2 “quiet” bottles, COMPARE showed that both CTD salinity channels were reading higher than the bottles with the primary high by 0.003 and 0.017psu while the secondary was high by 0.005 and 0.020psu. Given that these were bottles fired at the bottom, the effect of poor flushing would be to make the CTD look like it was reading higher than it really would be, though the error is smaller in the lower gradients found at depth. These comparisons are not suitable for calibration purposes.

For full details for the COMPARE run see file 2019-038-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are only 24 bottles of which 8 are close enough to the surface that incomplete bottle flushing and high vertical gradients lead to the differences being major outliers in the fit. Those below 5m have a very narrow range of values so the fit is quite flat. Choosing the same offset as used for 2019-001 and rejecting bottles from above 6m, and a few outliers based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0248 + 0.0275

The slope for 2019-001 was 1.0299. It was 1.03 for the 2019-016 casts in western Juan de Fuca when the offset was forced to the 2019-001 value. A lower slope was also found for 2019-016 when more bottles were included and that was likely due to incomplete flushing of bottles. Niskin bottles fired near the mouth of Juan de Fuca Strait are generally better flushed.

The only significant outliers are from very close to the surface where the DO vertical gradients are high and flushing issues are sufficient to explain them; no quality flag changes are justified..

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-038-ctd. The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. As usual the T, C and Sal channel pairs track well on downcasts and upcasts are much noisier. The Dissolved Oxygen, pH, transmissivity, fluorescence, pH and PAR profiles look normal. The PAR values are much lower than the SPAR values at the beginning of casts. The altimetry was sometimes spiky near the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

When this equipment was used during 2019-016 the best choice for aligning dissolved oxygen with temperature was found to be +2.5s.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration..

There were no casts deep enough to make a comparison of sensor pairs worth doing. Below are the results of 4 previous cruises using the same sensors

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2018-01-0043 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, Noisy |
| 2018-01-0062 | 1000 | -0.0006 | -0.0002 | -0.0021 | High, XNoisy |
| 2018-029-0007 | 300 | +0.0001 | -0.0005 | -0.005 | FHigh, FSteady |
| 2018-029-0093 | 300 | +0.0001 | -0.0004 | -0.0045 | High, V.Steady |
| 2018-029-0105 | 300 | +0.0001 | -0.0004 | -0.004 | High, V.Steady |
| 2018-030-0050 | 300 | -0.0008 | +0.00003 | +0.0009 | High, Steady |
| 2018-030-0086 | 300 | -0.0013 | -0.00003 | +0.0008 | F.High, Noisy |
| 2019-001-0020 | 1000 | -0.0009 | +0.00007 | +0.0019 | High, XNoisy |
| 2019-001-0037 | 1000 | -0.0009 | +0.00008 | +0.0019 | High, XNoisy |
| 2019-001-0049 | 1000 | -0.0005 | +0.00010 | +0.0019 | High, XNoisy |
| 2019-016-0042 | 325 | +0.0009VN | +0.00020 | +0.0012N | High, Steady |
| 2019-016-0055 | 325 | +0.0004 | +0.00016 | +0.0012 | High, Noisy |
| 2019-016-0084 | 325 | +0.0005 | +0.00013 | +0.0010 | Moderate, Noisy |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference list was printed and checked against the log book. No problems were found.

The cruise tracks (event #s and stn names) were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±3m the readings were checked. There were 2 such casts though neither was significantly high, being close to 4m, so these are not signs of serious problems. Plots show that the altimetry headers were reasonable for both, but that the altimetry value was varying during the stop at the bottom so may be a little less reliable than for other casts. No adjustments were made to the header entries.

A surface check was run and shows an average surface pressure for the cruise was 2.3db which is reasonable for the Vector and this area.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence vertical offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

pH

Tests were run during processing of 2019-016 found the best setting to align pH with temperature was +50 records. SHIFT was run to advance pH by 50 records.

Conductivity

Tests were run for 2019-016s to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. The best choice for both channels was -0.8 records.

SHIFT was run twice on the 2019-038 casts using -0.8 records for the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The pressure, temperature and conductivity sensors have been used on 4 cruises since they were last serviced and the DO sensor #3234 was used for 3 of those cruises.

- During 2018-030 the primary salinity was found to read low by 0.002, the secondary by 0.001.

- During 2018-032 there were problems with the salinometer so the comparison was not trusted.

- During 2019-003 the primary sensor was low by 0.002psu and the secondary low by 0.0004psu.

- During 2018-032 DO was recalibrated using a slope of 1.0324 and offset of 0.0096.

- For 2019-003 the results from 2019-001 were used – slope 1.0299 offset 0.0275.

- During 2019-016 the primary salinity was found to be low by about 0.0008psu and the secondary high by 0.0005psu. Using all casts the primary was lower than the secondary by 0.0011psu. The correction used for dissolved oxygen was slope 1.03/0.0275.

- There were no pressure problems noted in any of the cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.  All data fell within the climatology.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary channels were selected for 2019-016. The secondary channels were chosen for editing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. Most of the data removed are from near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise.

After editing T-S plots were examined for all casts. There are some small unstable features which could be real as this is an area of active mixing. No further editing was applied.

##### Recalibration

There is insufficient evidence from this cruise to determine what recalibration is needed. The same correction will be applied to channel Oxygen:Dissolved:SBE as applied for cruise 2019-016.

File 2019-038-recal1.ccf was prepared to apply the following correction to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.03 + 0.0275

This correction was first applied to the SAM and MRG files.

COMPARE was rerun for dissolved oxygen. Using the same data as for the original fit the average difference shows the CTD to be higher than the bottles by 0.02mL/L.  The standard deviation is 0.04mL/L. See file 2019-038-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When the same outliers were removed as in the first run of COMPARE the CTD DO was higher than the titrated samples by an average of ~0.07mL/L (standard deviation of 0.1L/L). The differences are highest at the upper end of the range where vertical gradients are high so the slow response rate of the DO sensor leads to CTD DO values that are a little too high. Incomplete flushing of Niskin bottles also lead to the CTD DO looking high. The differences are small given the limitations in the comparison method.

A plot of differences versus pressure was used to make an estimate of errors at different depths. The results fit well within the estimates for 2019-016.

No further recalibration is necessary.

See 2019-038-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Conductivity, Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed in editing

temperature and salinity.

Dissolved oxygen was calibrated using the method described in SeaBird Application Note

#64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when

it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration) a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will

be due to problems with flushing of Niskin bottles and/or analysis errors, so

the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.50 mL/L from 0 to 20db

±0.20 mL/L from 20 to 60db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data are available.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see document: 2019-038_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S space, but those are very small and may be real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~80% to 140% with the highest values in the north and lowest in the south. There is no obvious problem with the dissolved oxygen data but also no evidence that it is well calibrated.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-038-bottles-final.xls. The entries were compared with the rosette log sheets. It was noted that the bottle from event #6 that was not sampled had not been removed from the file. But it was left in on purpose so others may “second guess” the choice of which Niskins were likely closed.

Standards check, cross-reference listing  and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

##### Thermosalinograph

The thermosalinograph data were processed with the data from cruise 2019-016. The file was originally saved as 2019-016-0002. The final file was renamed as 2019-038-0001.tob.

For details on TSG processing see document “2019-016_Processing_Report.doc” archived with the 2019-016 data.

Data from this cruise were used for a quick assessment of whether the recalibration applied was appropriate but the near-surface data were too noisy to be useful.

The track plot was added to the end of this report. There are no temperature and salinity data for the last 10 hours as there was no flow in the loop. Times and positions were available.

PARTICULARS

6. Log shows 6 bottles fired with Niskin #4 fired twice. In fact, 7 bottles were fired. Source of samples                #16, 17 and 18 is not certain.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1201DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1396DR |  | 2Feb2016 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3685 |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |