## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19 March 2025 | Updated channel names & formats in TOB files. G.G. |
| 25 May 2022 | Added 0 to empty flag channels. G.G. |

## PROCESSING NOTES

Cruise: 2019-043

Agency: WVL, Ecosystem Sciences Division, Vancouver, B.C.

Location: Baynes Sound

Project: Baynes Sound

Party Chief: Sutherland T.

Platform: Vector

Date: 8 June 2019 – 11 June 2019

Processed by: Germaine Gatien

Date of Processing: 14 April 2020 – 16 April 2020

Number of original HEX files: 11	 	Number of CTD files: 10

Number of original ROS files: 10		Number of CHE files: 9

Number of original TSG files:  1	  	Number of TOB files:

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#953DR), a SBE 43 DO sensor (#3791), a SeaPoint Fluorometer (3642), a Biospherical QSP-400 PAR sensor (#4565), an SBE pH sensor (691) and an altimeter (43281). A Biosperical/Licor Surface PAR (#20518) was mounted on the boat deck.

SeaBird SBE21 Thermosalinograph #2488 in use during this cruise.

Seasave version 7.26.2.13 was used for CTD acquisition and 7.26.7.110 was used for TSG acquisition.

The data logging computer was 10.249.56.137.

The deck unit was a Seabird model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model,  8400B Portasal, serial # 68572.

The oxygen kit was SIO (nsB-6009) Kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science and Sampling Logs were generally in good order. The transmissometer serial number was entered incorrectly in the log and in the configuration file, but the parameters in the configuration file were correct.

The secondary sensors were selected for archiving because the primary channels were very noisy. The secondary conductivity sensor was found to be damaged in December 2019 but there was no sign of serious problems with the sensor during this cruise.

The casts were all very shallow and the comparison with bottles too noisy to enable an estimate of dissolved oxygen sensor accuracy.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples.

The pH:SBE:Nominal data should be used with caution; no field calibration data are available. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

There was no flow meter, intake thermistor or loop sampling. The thermosalinograph data were compared with data from co-incident CTD casts. The comparisons had a lot of scatter, so recalibration of salinity and derivation of a proxy for intake temperature were based on observations during 2019-016 and 2019-050 when the same equipment was used.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained and were generally in good order. The serial number for the transmissometer was entered wrong, both in the log and in the con file. It should have been #983. There were no photos available to confirm this, but #953 has not been in service for many years and the parameters entered are those for #983.

- Dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- There history was found for the use of the pressure, temperature, conductivity and dissolved oxygen sensors.

- The configuration files used at sea had an error in the transmissivity serial number but the parameters were correct. It was saved as 2019-043-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2019-043-ctd.xmlcon.

The ROS files were converted to IOS format. They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files and no significant outliers were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-043-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2019-043oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-043oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file QF2019-043SAL*.xlsx. The analyses were carried out in a temperature-controlled lab 37 to 39 days after collection. The files were simplified and saved as 2019-043sal.csv.

That file was then converted to individual SAL files.

(* - indicates date of creation for the files from analysts.)

The SAL and OXY files were merged with CST files in 4 steps.

Then the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. No problems were found.

##### Compare

Salinity

Compare was run with pressure as reference channel.

All the casts are very shallow, so results of such comparisons are not very useful. However, the average differences between the 2 salinity channels during stops were not large, ranging from the primary being higher by 0.002psu to the secondary being higher by 0.003psu, and the average difference including all salinity bottle stops shows the primary being lower than the salinity by an average of 0.0003psu.

There were 4 outliers in the primary comparison with 3 having a high standard deviation in the CTD data and 3 in the secondary with 2 having a high standard deviation in the CTD data. Excluding those outliers the primary salinity is low by an average of 0.0005psu and the secondary is high by 0.0004psu. This is similar to differences found during 2019-050 with no suggestion of major trouble with either sensor pair.

For full details for the COMPARE run see file 2019-043-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter in the plot as expected in shallow sampling..

The initial fit using all casts but forcing the offset to =0 and excluding outliers based on residuals was:

CTD DO Corrected ≈ CTD DO * 1.04 (R² = 0.41)

This is as close as can be expected to the fit used for 2019-050:

CTD DO Corrected = CTD DO * 1.0352 (R2 = 0.91)

For details see 2019-043-DO-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-043-ctd. The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present..

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

When this equipment was used during 2019-050 the best choice for aligning dissolved oxygen with temperature was found to be +2.5s. Tests are not useful for these casts due to the shallowness and many stops for bottles. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

There were no 2019-043 casts deep enough for differences between channels to be reliable. During the previous cruise the following results were found.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2019-050-0053 | 350 | +0.0004 | +0.00006 | 0.0000 | High, Steady |
| 2019-050-0066 | 350 | +0.0006 | +0.00006 | +0.0001 | High, Moderate |
| 2019-050-0101 | 340 | +0.0006 | -0.0003 N | -0.0038 | VHigh, VSteady |

There was a significant change in differences between casts #66 and 101.

During 2019-050 the secondary channels were selected for most casts because the primary channels were quite noisy. However, there was one poor cast near the end of the cruise for which the primary channels were selected and, as shown above, the differences between channels increased towards the end of the cruise. So differences were examined more closely for 2019-050 and the change to larger differences was found to be gradual and late in the cruise.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference list was checked against the log book and no significant discrepancies were found. The cruise tracks (event #s and stn names) were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth. Where that number exceeded ±3m the readings were checked. The altimetry values all look appropriate, but there were 3 cases in which the depth in the log book produces a better result than the one in the headers. Appropriate changes were made to the CLN files to events #1, 38 and 39 to the CLN files and to the MRG and MRGCLN2 files for event #38.

A surface check was run and shows an average surface pressure for the cruise was 2.0db which is typical for the Vector in protected waters.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. A plot of the one cast with no stops for bottles (#39) show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. For cast #39 the results look ok.

pH

Based on tests run during 2019-50 SHIFT was run to advance pH by 35 records. The results look reasonable for cast #39.

Conductivity

The best choice for the primary channel during 2019-050 was -0.7 records and for the secondary channel, -0.1records. Tests run on 2 casts show those settings are also best for these data, though differences were small between different settings and nothing seemed to improve the primary channels much.

SHIFT was run twice on all SBE911 casts using those settings. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The pressure sensor has been used for 7 cruises since the last factory calibration and no problems have been noted.

The temperature, conductivity and dissolved oxygen sensors were used during 2019-050 when the primary salinity was found to be very noisy and low by about 0.001 while the secondary salinity was low by 0.0022psu. After this cruise the secondary conductivity cell was replaced due to high residuals.

The dissolved oxygen sensor was used during 2019-050.

Historic ranges – Local climatology was not available at the time of processing.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were a post-cruise calibration available with comments that the secondary conductivity cell was replaced due to high residuals. The primary salinity appears to have drifted downwards by roughly 0.004psu.The secondary had drifted downwards by ~3.1psu.

##### DETAILED EDITING

Before choosing which T/S pair to edit and archive a few investigations were made.

- A general examination of some T-S plots suggests that the secondary data have many fewer unstable features than the primary. This was also the case during 2019-050.

- In December it was noted at the factory that the secondary conductivity cell needed replacement. The post-cruise calibration report shows secondary conductivity was very high which would lead to extremely low salinity values. The primary salinity appears to have drifted up by about 0.004psu.

- The bottle comparisons from 2019-050 does show a shift towards the end of 2019-050 to somewhat lower salinity, on the order of -0.003psu; some of that might have been related to incomplete flushing of Niskin bottles.

- The differences between sensors while the CTD was stopped for bottles are variable from this cruise but never more than 0.003psu with the secondary looking lower than the primary early in the cruise and slightly higher later.

While it is possible that the secondary sensors were drifting somewhat, there is no evidence of the very bad values observed at the factory. The comparison with the primary salinity highlights that the very large change in conductivity must have been due to cell damage that occurred or worsened significantly after 2019-043.

While the primary sensors may have a better calibration, the data are very noisy while the CTD was in motion. If we choose to archive secondary channels, there will likely be some error due to calibration drift, while if we choose the primary there will be data loss due to poor resolution of gradients leading to many instabilities. The secondary sensors were chosen for editing.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. All of the data removed are from near the top and bottom of casts. Salinity was cleaned lightly to remove unstable features that appear to be due to small misalignments. All files required some editing.

After editing T-S plots were examined for all casts. There are some small unstable features which could be real as they come from areas of active mixing. No further editing was applied.

##### Recalibration

There is no evidence that the pressure channel requires recalibration.

Dissolved Oxygen will be recalibrated by applying the fit found for cruise 2019-050 as discussed in section 4.

The damage to the conductivity cell appears to have occurred after this cruise. The comparison with bottles during 2019-050 suggested the secondary salinity was low by 0.002psu though some of that may have been due to incomplete flushing of bottles.  From this cruise it appeared to be high by 0.0004psu, though that comparison was poor due to the shallow sampling. For 2019-050 recalibration was done by adding 0.0012psu to the secondary channel to make it match the primary. Since there was some apparent drift of the secondary salinity relative to the primary during the latter part of cruise 2019-050, the correction to the secondary salinity will be set to +0.0035psu for this cruise.

File 2019-043-recal1.ccf was prepared to add 0.0035psu to channel Salinity:T1:C1 and apply the following correction to SBE Dissolved Oxygen:

CTD DO Corrected = CTD DO * 1.0352

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity and dissolved oxygen and showed the correction was done correctly.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The comparison of bottle data with downcasts was skipped since the casts are so shallow, so the results tend to be very noisy. Tests run during the previous cruise with the same equipment found no need to recalibrate further.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add header comments about the data processing.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined and no problems found.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~110% to 150% similar to values in the northern part of the Strait of Georgia during 2019-050.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-043-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found..

Standards check was run and no problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files. No problems were found.

##### Thermosalinograph Data

There was 1 thermosalinograph file.

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts and look at sensor history.

a.) Checking calibrations

The configuration file used at sea was checked and no errors were found.

b.) Conversion of Files

The cnv file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. There was a drop-out of salinity for a less than a minute that was likely a time when flow was turned off. No other problems were noted.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

- The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 2db. These were exported to a spreadsheet which was saved as 2019-043-ctd-tsg-comp.xls. There were only 10 CTD casts with data available from 2m; they all overlapped with TSG data.

- The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to 2019-043-ctd-tsg-comp.xls.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median and average differences were both <0.0000º with the largest differences being 0.0001º.

d.) Comparison of T and S from TSG and CTD data

When all casts are included the TSG temperature is higher than the CTD temperature by an average of 0.411C° and a median of 0.396C° with standard deviation of 0.45C°. We do not expect heating in the loop to be particularly large when ambient ocean temperatures are as high as those seen in Baynes Sound.

During 2019-050 it was noted that the difference in temperature varied from one region to another, which was likely due to variable vertical temperature gradients so that small mismatches in depth of TSG and CTD sampling have variable effects. The 3 casts from 2019-050 closest to Baynes Sound had temperatures higher by 0.170, 0.174 and 0.160C°. The median difference for 2019-050 was 0.209C° and it was 0.223C° for Haro Strait which was very well mixed. However, intake temperatures were lower in Haro than in the Strait of Georgia, so we might expect more heating in the loop in Haro.

Salinity is lower by an average of 1.14psu and median of 1.20psu and standard deviation of 0.30psu.

Variability is very high and plots show no evidence of differences being lower when standard deviations over the 2–minute window are lower. For the 3 casts closest to Baynes Sound during 2019-050 the salinity was low by 0.15, 0.52 and 0.27psu, but in Haro Strait where vertical gradients were very low the salinity was low by 0.024psu.

There are few data, large scatter and relatively high vertical gradients. It is likely the TSG was sampling shallower data than the CTD so that temperatures are higher and salinity lower due to that difference. TSG temperature would also be higher due to heating in the loop.

See 2019-043-ctd-tsg-comp.xls for details.

Calibration History

The TSG temperature and conductivity were recalibrated in November 2018 and this is the 3nd known use of it since then. During cruise 2019-016 the TSG Salinity was found to be low by about 0.03psu. The TSG lab temperature was higher than the CTD temperature by about 0.37Cº which was considered unusually high. It was thought that the flow rate must have been lower than usual, though vertical gradients might be a factor as well.  For 2019-050 recalibration was done by subtracting 0.2C from Temperature:Primary and adding 0.03psu to Salinity:T0:C0.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by a median of ~0.40Cº which is much higher than expected given the surface water temperatures were fairly high. The standard deviation in the TSG temperatures were also high reflecting great variability in these waters. The large difference is likely due to high vertical temperature gradients with the TSG likely drawing water from above the shallowest CTD data available. Subtracting 0.2Cº from TSG temperature to obtain an intake temperature looks reasonable based on well-mixed casts during 2019-050.

3. The TSG Salinity is lower than the CTD salinity by a median of 1.2psu, much higher than during the previous 2 cruises. This is likely due to high vertical near-surface gradients and a mismatch of sampling depths. Subtracting 0.03psu from the TSG salinity is a reasonable estimate based on 2 previous cruises.

f.) Editing

Temperature:Primary and Salinity:T0:C0 channels were edited to remove 27 values using CTDEDIT because flow in the loop appeared to have been off.

g.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2019-043-tsg-recal1.ccf to subtract 0.2Cº from Temperature:Primary and to add 0.03Cº to Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan Number and Flag channels.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

The TSG sensor history was updated.

As a final check plots were made of the cruise track which looked fine.

A time-series plot turned up one missed bad salinity point which was padded in the edited file and the final steps were rerun. The plot then looked fine.

The cruise plot was added to the end of this report.

Particulars

1. Test cast – 12 bottles closed and 5 samples taken but only for exercise purposes. No CHE file needed.

6. Sample numbers started at 13.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0443 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 4888 |  | 14Feb2018 |  | Factory |  |  |  |  |
| Conductivity |  |  | 3531 |  | 06Mar2018 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 4700 |  | 13Feb2018 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3791 |  | 13Feb2018 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 983DR |  | 9Aug2017 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  | 3791 |  | 13Mar2019 |  | Factory |  |  |  |  |
| PAR |  |  | 4565 |  | 16Mar2011 |  | Factory |  |  |  |  |
| pH |  |  | 691 |  | 13Dec2018 |  |  |  |  |  |  |
| SPAR |  |  | 20518 |  | 21Mar2016 |  |  |  |  |  |  |
| SeaPoint Fluor. |  |  | 3642 |  |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 31Aug2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 43281 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/21/2488

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 14Nov18 | Factory |  |  |
| Conductivity | 2488 | 14Nov18 | Factory |  |  |