## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2019-053

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B.C.

Chief Scientist: Neville C.

Platform: Sea Crest

Location: Juan de Fuca Strait and Strait of Georgia

Project: Strait of Georgia Juvenile Salmon

Date: 10 September 2019 – 28 September 2019

Processed by: Germaine Gatien

Date of Processing: 28 April 2020 -  12 May 2020

Number of original HEX files: 58 (1 with no data)	Number of CTD files: 57

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #2216, dissolved oxygen sensor #1176, SBE pH sensor 0852 and pressure sensor #464.

# SUMMARY OF QUALITY AND CONCERNS

There was a log book available plus a spreadsheet of events.

The format used for header information in the raw files did not permit automatic conversion into IOS Headers. If the following format is used in the HEX files the information will be entered automatically and this will make data processing quicker and less error-prone:

** Latitude: 49 9.781 N

** Longitude: 123 32.970 W

** Station: 39

** Depth: 378 m

There was a spreadsheet provided with the necessary information which is very helpful.

The configuration file provided contained out-of-date calibration parameters. All instruments had been recalibrated in November 2018 except the fluorometer. The offset used for the fluorometer had been changed from the one used since the last factory check. While this could have been done on purpose if deep fluorescence values seemed too large, in this case it produced significantly negative values for most depths. So the offset on file from the previous factory calibration was used. This produced deep fluorescence values similar to those found during another cruise in the Strait of Georgia in June 2019 using a different type of fluorometer.

Deployment involved a soak at the surface for between 1 and 2.7 minutes before starting the cast. Few soaks lasted for the recommended 2 minutes to allow equilibration of the conductivity sensor. When a dissolved oxygen sensor is in use, it is recommended that there be a soak at 10m before running casts. If that is undesirable for a particular project, then a lengthy surface soak is particularly important.

There are concerns about the reliability of the salinity data so they are reported with only 3 decimal places. There were some unstable excursions in salinity. These are likely due to variable flow rates leading to misalignment of temperature and conductivity. Plumbing or pump problems are possible causes. Some such features appear related to descent rate variations which may again be a result of flow rate problems. The descent rate was generally high and steady for this cruise which may account for the salinity problem being a little less severe than during 2019-052 when this equipment was last used.

For dissolved oxygen there are two concerns. First, because of the deployment method there are likely to be values that are too high around 10m as bubbles are released. Second, the upcast data often look unrealistic with values that vary little and are much lower than downcast values. This could also be due to some problem in the flow that affected the salinity data. While there is no evidence of this affecting the downcast dissolved oxygen data, that is a possibility. As found for salinity, the dissolved oxygen data appear to be somewhat less affected during this cruise than during 2019-052.

There was no calibration sampling during this cruise nor during 2019-052, the only other cruise during which these sensors were used since they were last serviced at the factory. No recalibration was applied.

While the CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples.

Warning: The pH:SBE:Nominal data should be used with caution; no field calibration data are available.  Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were in standard format.

2. Preliminary Steps

A spreadsheet with event information was provided.

The configuration file provided was out of date with calibrations from 2015 and 2017. All instruments had been recalibrated in November 2018 except the fluorometer. The file used at sea had a different offset for the fluorometer than has been in use since the last factory check. While this may could have been done on purpose if deep fluorescence values seemed too large, in this case it produced significantly negative values for most depths. So the offset on file from the previous factory calibration was used. This produced deep fluorescence values similar to those found during another cruise in the Strait of Georgia in June 2019 using a different type of fluorometer.

The configuration file with the most recent calibration parameters was saved as 2019-053-ctd.xmlcon.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

The HEX files were converted using configuration file 2019-053-ctd.xmlcon.

There was no 10m soak. The surface soak varied from 1 to 2.7 minutes.

Examination of a selection of casts shows descent rates to be high.

- There were problems in conductivity and salinity in some casts similar to those noted during 2019-052 using the same equipment, but they seem less severe overall.

- The dissolved oxygen data  near the surface for upcasts and downcasts look significantly different though the full profiles are closer than during 2019-052. There are local maxima around 5 to 10m but they may be a little less of a problem than during 2019-053 perhaps because the soak was usually a little longer than during 2019-052 and the descent rates higher.

- Conductivity looks like it had equilibrated for most casts at the end of the soak period.

- Upcast profiles look quite different from those in downcasts in conductivity and salinity channels and towards the end of the cruise there are a number of casts with very poor upcast temperature. This does not appear to vary with soak period length.

- Fluorescence and pH look normal.

Based on upcast conductivity dropping to low values, it appears that the surface is between -1db and -0.3db but since there would be water still moving in the tubing; this is likely as close to 0 as can be expected for this type of test and it is within the specifications for the instrument.

4. WILDEDIT

Most casts had no spikes. At  least one cast did have spikes in the data but tests showed WILDEDIT made no significant difference with a variety of settings; the spikes involved too many points for WILDEDIT to be effective. It looks like something was wrong with the plumbing. WILDEDIT was not run.

5. WFILTER

Tests were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. WFILTER was run using a cosine filter, size 9 on the pressure, depth, temperature and conductivity channels since that was the best choice for the previous use of this equipment.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results do not look good, but that is likely because the DO data are poor during upcasts.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8). When problems were found in the conductivity channel during 2019-052 tests were run using a variety of settings but none had any better results to (0.04,8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and it is clear that the data have been improved, but many problems remain.

Plots were made of dissolved oxygen concentration data. The data look better than during 2019-052 when the upcasts were very poor, but there are some casts with suspiciously low DO values during upcasts and there are some downcast maxima around 10m that may be due to bubble release.

There was another cruise in this area that left after the end of this one and used a different CTD. An attempt was made to find stations that could be compared but the closest in space were almost a week apart in time and in an area where temporal and spatial variations are very large. The comparison at 10m indicated that temperatures were higher, salinity lower and dissolved oxygen higher than found 6 days later at a cast close to 2019-053-0220. So the low dissolved oxygen values do appear to be reasonable. Another comparison was attempted in an area with somewhat lower variability but further apart in time. This cruise had lower temperatures, higher salinity and much lower dissolved oxygen at 10m in both downcast and upcast. But 2 nearby casts in the same area from the later cruise show that local variability is high enough to make any comparison of little value.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values. Where fluorescence values were negative, pad values were substituted.

A second run of CLEAN was run to do the same for negative DO values.

Merge Headers was used to add latitude, longitude, water depth and station names.

Times were checked against the log and while the file times were expected to be in local time (UTC-7 hours) they were actually 9 hours off. Likely an attempt to change the computer from PST to PDT moved in the wrong direction.

ADD TIME CHANNEL was run to add 9 hours to all start times.

10. Checking Headers

The header check was run. There were negative values for pressure and conductivity that will likely be removed by DELETE.

The speed check turned up a problem with the latitude for 2 casts which were wrong in the spreadsheet provided with 48º instead of 49º. Those were fixed in the spreadsheet and MERGE Headers and ADD Time Channel were rerun and the header check rerun. No further problems were found.

A cross-reference list was produced and checked against the event spreadsheet; one event was missing  that was in the spreadsheet but there was no data acquired for that event.

Since there was no subsurface soak, CLIP was not needed.

The surface check shows a mean value of -0.04db with values ranging from -0.58db to +1.1db. As it starts recording as it passes through the surface this indicates a good pressure calibration. As noted earlier at the end of casts conductivity went to 0 when pressure was between -1.db and -0.4db. Because there is some lag in conductivity due to water remaining in the cell, the pressure is likely well within the expected range.

##### 11. SHIFT

Conductivity

Tests were run to see what alignment parameter for conductivity would improve stability in T-S space.

During 2019-052 a setting of +5.5 records suited most casts but some looked better with +0.6 records. The latter is a value closer to our usual settings for this type of CTD.

The larger setting looks better as judged by appearance on a T-S curve after putting the file through DELETE, but it looks far from perfect, so tests were run a variety of settings. The best overall result was with an advance of +3.5 records though the optimum varies from feature to feature and cast to cast.

This is an unusually large alignment correction suggesting a problem with flow to the sensor. This may account for the unusually large differences between downcasts and upcasts, particularly in salinity and dissolved oxygen.

SHIFT was run apply an advance of 3.5 records to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined. In many places the upcast temperature looks poor so it is hard to judge, but overall the alignment looks reasonably good.

Dissolved Oxygen

This channel was aligned earlier but the effect is impossible to judge due to the poor upcast data in all channels.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017/2018 cruises using this equipment. For cruise 2019-052 a setting of +10 records looked best overall and it also looks appropriate for this cruise.

SHIFT was run on all casts to advance the pH signal by 10 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

There were problems in the conductivity channel so salinity data often look bad. It is frequently difficult to tell good salinity from bad, though the problems during this cruise were not as severe as during 2019-052, possibly because the descent rate was higher and usually quite steady.

CTDEDIT was used to remove some salinity points that look bad and salinity was cleaned. Other editing included removing some near-surface records and records corrupted by shed wakes.

All files needed editing except for event #137.

T-S plots were examined after this step. Some unstable features remain. Some minor instability is expected in this area but not so many large unstable features..

16. Calibration checks

Sensor History – This was the 2md use of the dissolved oxygen, pressure, temperature and conductivity sensors since they were last calibrated. There was no calibration sampling during the previous cruise and no recalibration applied to any channels.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Where available, profiles of temperature and salinity were examined with a 3-standard deviation climatology superimposed. There were some excursions from the climatology but all were in near-shore casts where this climatology is not considered appropriate. Low salinity and high temperature values were found around 20db in the Discovery Islands and Gulf Islands and at some casts around Texada Island temperatures were high near the bottom. These mostly look like cases of strong mixing and do not appear to indicate calibration problems. No excursions were found in the central Strait of Georgia or Juan de Fuca Strait.

Post-cruise calibrations – There were no post-cruise calibrations available.

17. CALIBRATE

There was no calibration sampling and no history of use during other cruises.

No recalibration was applied.

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data. Plots show it was effective in smoothing fluorescence slightly.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

Plots were examined and no spikes were found.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 50% to 135%. Such variability is not unusual in this region where some casts are in tidally mixed regions with small vertical gradients while others have high near-surface DO gradients. The lowest values were in narrow channels and the DO profiles show well-mixed temperature and dissolved oxygen. The highest saturation rates are associated with high vertical gradients where tidal mixing may have been insignificant. The DO profiles occasionally show bulges at 10m but they are not as obvious as during the previous cruise using this equipment, 2019-052.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add the following note to the headers:

Conductivity, Fluorescence:URU:Wetlabs and pH:SBE data are nominal and unedited

except that some records were removed in editing temperature and salinity.

The CTD was soaked at the surface for between 1 minute and 2.7 minutes.

There was no salinity or dissolved oxygen calibration sampling from this cruise.

This was the second use of the sensors since they were last serviced at the

factory so calibration drift is likely low.

There are concerns about the reliability of the salinity data so they are only

reported with 3 decimal places.

The problems with salinity are somewhat random unstable excursions that derive

from the conductivity channel and are likely due to variations in the alignment

with temperature data. This may be due to variable flow rates possibly due to

a pump/plumbing problem. The temperature data look fine.

Dissolved oxygen values are likely to be high around 10 to 20m as bubbles are

released as there was no initial 10m soak.

While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

A cross-reference listing was produced.

A header check was run on the CTD files and no problems were found.

The sensor history was updated.

Track plots were produced and look fine; they were added to the end of this report.

Plots of CTD casts were examined and no further problems were found.

Particulars

196 – No data acquired due to low battery.

CRUISE SUMMARY

| Cruise ID#: 2019-053 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 18 June 2018 End: 6 July 2019 |  |  |  |  |  |  |
| Location: Sea Crest |  |  |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 404 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2019-053

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 8Nov2018 | Factory |  |  |
| Conductivity | 1764 | 21Nov2018 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 3Nov2018 | Factory |  |  |
| pH | 0852 | 13Dec2018 | Factory |  |  |
| Pressure | 0464 | 26Nov2018 | Factory |  |  |