## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2019-067

Agency: IOS, Ocean Sciences Division, Sidney BC

Chief Scientist: Vagle S.  		Platform: Vector

Location: Strait of Georgia		Project: OPP- Marine Environmental Quality Program

Date: 30 May 2019 –2 June 2019

Processed by: Germaine Gatien

Date of Processing: 17 April 2020 – 24 April 2020

Number of original HEX files: 34	Number of CTD files: 34

Number of raw TSG files: 1		Number of processed TSG files: 1

# INSTRUMENT SUMMARY

Three CTDs were used during this cruise:

A SeaBird Model SBE-25 CTD (s/n 0334) was used with temperature sensor #4484, conductivity sensor #2128, Wetlabs ECO Fluorometer #2215, dissolved oxygen sensor #1483 and pressure sensor 0482.

SeaBird Model SBE911 (s/n)

# SUMMARY OF QUALITY AND CONCERNS

There was no personnel list in the log book. Daily Science Log Book. While there were some notes about problems encountered, there was no mention of the change of CTDs at event #21. The equipment list lacks any information about CTD #0443 which was used for the last CTD cast; while there was no room for 3 entries in the log, there is space for notes elsewhere, so a list could have been created or pasted in.

The file names were non-standard with location information in the names for the SBE25 casts and SBE911 casts having the wrong event numbers.

The serial number was incorrect in the log and configuration file for the transmissometer used on the SBE911+ CTDs, but the parameters were correct in the configuration file.

Header information including positions, station names and water depths were entered in the raw files for the SBE25 casts in a format that enabled easy conversion into IOS Header format. That was very helpful.

It is recommended that there be a wait of at least 30s at the surface after the 10m soak to allow the water to settle from the effects of the CTD passing through.

During a previous cruise the SBE25 pressure was found to be high by between 3 and 4db, so the offset in the configuration file for this cruise was changed from 0 to -3.2db.

The upcast was used for cast #6 because the downcast data looked unreliable in the top10m and the CTD was not mounted in a rosette so the upcast data quality is reasonable good.

There was no calibration sampling.

Thermosalinograph salinity was recalibrated and a proxy for intake temperature was created based on a comparison with co-incident CTD casts.

# PROCESSING SUMMARY

General

There were 3 CTDs used for this cruise, an SBE25 and 2 SBE911+ CTDs.

- No configuration file was provided for the SBE25, so it was assumed to be the same as the one used for cruise 2019-045 when the same equipment was used.

- Configuration files were provided for the SBE911s. These were checked and the only error found was the serial number of the transmissometer that was used on both of those CTDs. This was incorrect in the log as well, but the parameters in the configuration files were correct. CTD #1222 had not been used since it was last serviced. CTD #0443 had been used for 6 cruises before this one and 2 that have been processed and occurred after this one.

PART 1 – SBE25 Processing – Casts #2 - #19

1. Seasave - This step was completed at sea. The file names were not in standard format. They were based on station names.

2. Preliminary Steps

The Daily Log was obtained. There were no significant problems noted during this part of the cruise.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

No configuration file was provided but it is believed to be the same one as used during 2019-044 and 2019-045 which bracketed this cruise. The file for 2019-045 was used as it had offsets adjusted to pressure and fluorescence based on the results of 2019-044. The file was saved as 2019-067-ctd1.xmlcon.

The HEX files were converted using configuration file 2019-067-ctd1.xmlcon.

Plots show that the channels all produced reasonable values.

There is the usual steppy pressure which leads to unusual profiles. The descent rate is extremely noisy, so the files were reconverted without that channel.

After conversion the file names were changed to standard format.

4. WILDEDIT

The only spikes noted in the data occurred at the beginning or end of the casts or included many points, and will be removed in the normal course of editing. So WILDEDIT was not run.

5. FILTER

Normally pressure is filtered later in processing when running DELETE, but the poor resolution of this instrument means it is necessary to do this early to make sense of the other data which update more often than pressure. So program FILTER was used to apply a low-pass filter with a time constant of 0.5 seconds to pressure. The results were excellent.

Next the temperature and conductivity were examined and the usual approach of applying a cosine filter size 5 did a good job of removing small reversals.

6. ALIGNCTD

Based on tests run for other cruises using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

An initial header check showed the presence of many spikes, but examination of plots showed the negative values occurred at the beginning or end of casts or in the soak period.

The next step is to remove the data collected during soaks at 10m. Plots were made to determine the approximate scan # when the CTD is just about to do its full cast. This varied from 1720 to 2780 records.

CLIP was run to remove the records. Plots were examined after this step and the results looked good.

10. Checking Headers

The header times are in PST so ADD TIME CHANNEL was run to add 8 hours. (This step was accidentally missed for casts #2-4 and 10-19, so a text editor was used later on the REO and CTD files to correct the time.)

A cross-reference list was produced and compared with the log book entries. There were no significant discrepancies. Track plots looked ok but won’t be added to the end of this report until the SBE911 casts have also been processed..

HEADER CHECK was run. There remain some negative values but they appear to be all from the upcast. This will be checked after DELETE is run.

When conductivity drops suddenly to ~0 the pressures were studied and they ranged from -0.15db to +0.3db with a median of <0.1db, so the pressure calibration looks well within expected accuracy.

##### 11. SHIFT

Conductivity

Tests on 4 casts from 2019-044 showed that a shift of +0.5 records made the best improvement to stability in T-S space. That setting also worked well when tests were run on casts from this cruise.

A shift of +0.5 records was applied to all casts from 2019-067.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. The alignment looks good. No further adjustment was made. There are some noisy sections of dissolved oxygen but all appear to be during upcasts.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were warnings for 1 cast but they pertain to ends of upcasts.

13. DETAILED EDITING

An initial examination of T-S plots showed a few casts required editing near the top and/or bottom of the casts. Cast #6 had unstable features in the top 10m. This could be due to the very short wait at the surface before running the full cast. The CTD might have passed through water stirred up by the CTD as it rose from the 10m soak; it is best to wait at least 30s to give the water time to return to normal conditions. The upcast data were examined to see if they would be a better choice; they have more unstable features throughout the water column, though those at the surface are not as significant as in the downcast. However, the overall quality is poor, so the downcast was chosen for archiving.

CTDEDIT was used to remove records some records near the top and/or bottom of casts and to clean salinity where unstable features looked likely to be caused by misalignment of T and C.

Notes of editing details were made in the headers.

Editing was applied to casts #6, 7, 8, and 9.

T-S plots were examined after this step and no significant problems were found; there are small unstable features but this is an area where some instability is expected.

16. Other calibration checks

Sensor History – These sensors were used for 2 previous cruises since the last service. There was no sampling during the first. During the second the comparison was not considered suitable for recalibration but showed salinity to be high by ~0.008psu. During 2019-045 which occurred after this cruise, there were similar results. Systematic errors may have contributed to the CTD salinity looking high.

Historic Ranges – All salinity fell within the local climatology. Temperature between 20m and 60m was occasionally above the range maximum at events from Jordan River to Station Tran5. This is likely a real excursion rather than a calibration problem. The climatology is old.

Post-cruise calibrations – None were available.

17. CALIBRATE

Pressure was adjusted in conversion.

There is insufficient information to justify recalibration of dissolved oxygen or salinity.

The offset used for the fluorescence in conversion looks appropriate.

No recalibration was applied.

18. Fluorescence Filter

The fluorescence data was not filtered as it was not very spiky.

19. Bin Average, Remove, Derive DO in mass units, Reorder

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 74 to 95% with the highest value at station Swiftsure. The casts in Juan de Fuca and Haro Straits were well-mixed near the surface which is associated with low saturation. The offshore value may suggest that the DO values are slightly low, but there is insufficient information to justify recalibration.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and fix an error in the headers and to add comments.

A cross-reference listing was produced.

A header check was run; a problem was found in the times for some casts – the ADD TIME CHANNEL step had been missed for casts 2-4 and 10-19.

A text editor was used to correct the header times by adding 8 hours to those casts.

A standards check was run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

PART 1I – SBE911+ Processing – Casts #21 - #31

##### 21. Preliminary Steps

SEASAVE was completed at sea; the raw data files have extension HEX.

There are 9 casts using CTD #1222 and 1 using CTD #0443.

The log book mentions that archiving did not start until cast #23 was at 10m of the downcast, and the difference between the 2 temperature channels was 0.3Cº during cast #255.

The cruise summary sheet was completed.

##### 22. Conversion of Full Files from Raw Data

Files for casts #21 to 30 were converted using 2019-067-ctd2.xmlcon.

The file for casts #31 was converted using 2019-067-ctd3.xmlcon.

The Tau function was selected but not the hysteresis function as there was no deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

The primary channels (T, C and S) from CTD #1222 have bad data. This explains the large difference in temperature noted in the log book. The log does not record what pump the DO sensor was mounted on, but the downcast and upcast are very different, so it is obvious that it was mounted on the primary pump.

For cast #31, with a different CTD, the data look fine.

Transmissivity, Fluorescence, pH and PAR profiles look normal. The altimetry was fine near the bottom.

##### 23. WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

The few spikes noted were removed by this step.

##### 24. ALIGN DO

The sensor used for CTD #1222 did not produce good data, likely due to a pump or plumbing problem, so the choice of alignment parameter is irrelevant. For CTD #0443 the sensor was one that was used on other recent cruises and a setting of +2.5s looked best. ALIGNCTD was run on all casts using +2.5s.

##### 25. CELLTM

CELLTM was run using the default setting (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and this step does improve the data.

##### 26. DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on the final cast to find differences between the pairs of temperature, conductivity and salinity channels. The shaded lines show results from the cruise that followed immediately and used the same equipment.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2019-067-0031 | 207 | +0.0006 | -0.00005 | -0.0012 | High, F.Steady |
| 2019-050-0053 | 350 | +0.0004 | +0.00006 | 0.0000 | High, Steady |
| 2019-050-0066 | 350 | +0.0006 | +0.00006 | +0.0001 | High, Moderate |

The differences are reasonably close to those from the later cruise.

##### 27. Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### 28. Checking Headers

The header check was run. One station name was missing, so that was corrected.

The cross-reference list was printed and checked against the log book.  No errors were found.

Times are in UTC.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check was made by subtracting maximum depth sampled plus altimetry header from the water depth and problems were found in most casts:

- For 5 casts the difference was large; when log depths were used rather than header depths, the results were much better.

- For cast #24 the log notes that depth was changing during the cast; the CTD did not get within 15m of the bottom so there is no altimetry entry. The header depth entry is not appropriate for the position when the ship was at the bottom but the log entry looks appropriate.

- For cast #28 the header entry was missing but was found in the log and looks good.

- For event #29 there was no log or header entry for depth, so a value was estimated based on maximum depth sampled plus altimetry header.

Water depth was adjusted or added for 8 casts. Details may be seen in document 2019-067-altimeter-ctd.xlsx.

A surface check was run and shows an average surface pressure for the cruise was 2.5db which is reasonable. The lowest value was 1.6db. The 2 casts at the end (#30 and #31) were fairly close together but used different CTDs; they have surface pressure 1.66 and 1.60db which proves nothing but at least does not suggest a problem with pressure. At the end of cast #23 the CTD appears to have reached the surface; the pumps were off but transmissivity suggests that the surface was at about 0.4db.

Plots were made to compare casts #30 and 31. There are significant differences between the secondary temperature and salinity but this is an area of rapid change due to tidal mixing. The 2nd cast looks as though it has mixed more than the 1st. The primary sensors look totally different, which is expected given the problems noted earlier.

##### 29. Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. The only cast with reasonable DO data was #31. The alignment looks ok, for that cast. No further alignment will be applied to the DO concentration channel,

pH

A setting of +35 was used when CTD#0443 was used during 2019-050 which followed this cruise. That setting also appears to work well for the casts using CTD #1222.

SHIFT was run to advance the pH channel by +35 records for all SBE911 casts.

Conductivity

- For casts #21 to #30

This sensor has no history from other cruises. The primary channels are bad so only the secondary was studied. Settings of +0.3 to +0.8 made slight improvements. A setting of +0.6 records was chosen as best overall, but the differences are very slight.

SHIFT was run using +0.6records for the secondary conductivity. Salinity was recalculated.

- For cast #31 the parameters found cruise 2019-050 were applied.

SHIFT was run applying with the setting for the primary channel being -0.7records and for the secondary channel, -0.1records. Salinity was recalculated for both channels.

Cast #31 was too well mixed to see any difference in T-S space before and after the shift.

##### 30. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were many warnings but most referred to bottom and upcast data so are of no concern. The one exception was at the surface for event #25. This does not seem to have led to any significant loss of data.

NOTE: The log file for this run was appended to the run for the SBE25.

##### 31. Other Comparisons

Previous experience with these sensors –

- CTD #1222 There is no history for any of the sensors on this CTD.

- CTD #0443 The pressure sensor has been used for many cruises with no indication of calibration drift. The temperature and conductivity sensors were recalibrated in March 2018 but have not been used on any cruises before this one. They were use during cruise 2019-050 which immediately followed this cruise. The primary salinity was found to be low by 0.001psu and the secondary salinity low by 0.0022psu. Problems developed with the secondary conductivity towards the end of 2019-050 and the cell needed replacement in January 2020. The dissolved oxygen sensor was also used during 2019-050 and recalibrated by applying an offset of 1.0352.

Historic ranges – All secondary temperature and salinity fell within the local climatology.

Repeat Casts – There were no repeat casts and nearby casts are too shallow to provide a reasonable comparison.

Post-Cruise Calibration – There were post-cruise calibrations available for CTD 0443 but there were other cruises between this one and the factory visit. The secondary conductivity sensor was found to have been damaged, but a bottle comparison shows little drift for most of the cruise that followed this one.

##### 32. DETAILED EDITING

For most casts only secondary channels are useful. For cast #31 the secondary channels were selected for archiving since they were found to be best during 2019-050.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes or ship effects. Data were only removed from near the top and bottom of casts. Salinity was cleaned very lightly in a few casts.

All files required some editing except for cast #26.

After editing T-S plots were examined for all casts. There are some small unstable features which could be real as they come from areas of active mixing. No further editing was applied.

##### 33. Recalibration

- For casts #21 to 30 we have no sensor history since the last factory calibration, no calibration sampling and no-post cruise calibration so no recalibration will be applied. Pressure at the surface at the end of one cast was ~0.4db.

- For cast #31 recalibration was run using file 2019-067-recal1.ccf was prepared to add 0.0012psu to channel Salinity:T1:C1 and apply the following correction to SBE Dissolved Oxygen:

CTD DO Corrected = CTD DO * 1.0352

CALIBRATE was then run on the EDT files using the same recalibration file.

##### 35. Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### 36. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

There were no warnings.

On-screen T-S plots were examined. There are only the small instabilities that are expected in this region.

Profile plots were examined to see if there any problems. None were noted.

##### 37. Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts except #31 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE,  Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- For casts #31 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add header comments about the data processing.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile and T-S plots were examined. There are some unstable features in T-S space where waters are well-mixed, but those are very small and may be real.

The sensor history was updated.

##### 38. Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated for cast #31 and plotted. The surface value was ~75%. This cast is in Haro Strait and quite well mixed so a low value is expected. It is close to the saturation values seen using the SBE25 at the eastern end of Juan de Fuca Strait.

REPORTS FOR ALL CASTS – SBE25 and 2 SBE911+ CTDs

Header check, Cross-reference lists and track plots were prepared for all CTD casts.

The track plots were added to the end of this report.

PART I1I –Thermosalinograph

##### 39. Thermosalinograph Data

There was 1 thermosalinograph file.

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only methods available for checking calibration is to compare with the CTD casts during this and other recent cruises using the same system. The TSG was used during 2019-016 in April and during 2019-050 in June.

a.) Checking calibrations

The configuration file used at sea was not available at the time of processing, but it is presumed to be the same as the one used during 2019-050.

b.) Conversion of Files

The cnv file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. There is a 20-min section at the beginning with near-zero salinity and very high temperature that was followed by a period of equilibration. No other problems were noted.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

- The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 2db. These were exported to a spreadsheet which was saved as 2019-067-ctd-tsg-comp.xls. There were 21 CTD casts with data available from 2m; all of them overlapped with TSG data.

- The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to 2019-067-ctd-tsg-comp.xls.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were both 0.0000º but one cast had differences of 0.0046 º latitude and 0.0074º longitude. Examination of the TSG file showed that the positions were changing rapidly so that a small difference in times between the 2 systems would be significant. The TSG clock appears to have worked well.

d.) Comparison of T and S from TSG and CTD data

When all casts are included the TSG temperature is higher than the CTD temperature by an average of 0.26 C° and a median of 0.23C° with standard deviation of 0.18C°.  Salinity is lower by an average of 0.027 and median of 0.007psu and standard deviation of 0.079psu. The results differed significantly between Haro Strait and Juan de Fuca Strait, especially for salinity. In Haro Strait the standard deviations were large reflecting the high temporal variability, but the median value is close to that found during 2019-016 in April. In Juan de Fuca the average and median values are close.

|  |  | TSG TEMP-CTD TEMP | TSG SAL-CTD SAL |
| --- | --- | --- | --- |
| All casts | average | 0.2641 | -0.0273 |
|  | median | 0.2345 | -0.0070 |
|  | std dev | 0.1840 | 0.0788 |
| Juan de Fuca Strait | average | 0.2331 | -0.0057 |
|  | median | 0.2345 | -0.0055 |
|  | std dev | 0.0420 | 0.0065 |
| Haro Strait | average | 0.3242 | -0.0695 |
|  | median | 0.2003 | -0.0312 |
|  | std dev | 0.3119 | 0.1322 |

See 2019-067-ctd-tsg-comp.xls for details.

Calibration History

The TSG temperature and conductivity were recalibrated in November 2018 and this was the 2nd known use of it since then. During cruise 2019-016 in April the TSG Salinity was found to be low by about 0.03psu. The TSG lab temperature was higher than the CTD temperature by about 0.37Cº which was considered unusually high. It was thought that the flow rate must have been lower than usual.

Cruise 2019-050 used this equipment immediately after this cruise. The temperature was found to be high by a median of 0.21Cº. There was great variability in salinity differences between different regions which was thought to be due to variable near-surface vertical gradients.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than that of the CTD casts by a median of 0.23Cº using either all casts or only the Juan de Fuca casts. On the cruises before and after this one it was 0.37Cº (2019-016) and 0.21Cº (2019-050). The 2019-050 result likely reflects the intake water being a little warmer so there is a little less heating in the loop. The 2019-016 cruise had some odd results that might have been affected by flow rate in the loop being lower. Applying a correction of -0.23Cº looks reasonable.

3. The TSG Salinity is low by a median of 0.0070psu using all casts, 0.0055psu in Juan de Fuca Strait and 0.0312 in Haro Strait. Only in Juan de Fuca Strait is the standard deviation low. The salinity correction applied for 2019-016 TSG salinity was +0.03psu. For 2016-050 the comparison was not trusted due to great variability among regions and large near-surface salinity gradients. Applying a correction of +0.007psu looks reasonable given the amount of drift expected from conductivity sensors.

f.) Editing

Editing was applied to the beginning of the file to remove temperature and salinity data during the initial 20 minutes during which flow appears to have been off or still equilibrating.

g.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2019-050-tsg-recal1.ccf to subtract 0.23 from Temperature:Primary and to add 0.007Cº to Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan Number and Flag channels.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. Some header items were corrected at this stage as well.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

PARTICULARS – notes from logs

10. Winch stopped at 10m on way up by accident.

23. Forgot to archive until ~6m.

24. Forgot to archive; restarted. Cast on slope – changing bottom depth.

25. Temperature sensors are 0.3 out.

29. Test cast after service. Swapped out pump cable.

31. Swapped out CTD 1222 for CTD 443 for this cast.

CRUISE SUMMARY

| Cruise ID#: 2019-067 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 30 May 2019 End: 2 June 2019 |  |  |  |  |  |  |
| Location: Strait of Georgia Zooplankton |  |  |  |  |  |  |
| Chief Scientist: Vagle S. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| CTD 1 | SEABIRD | 25 | 334 | No | Yes |  |
| CTD 2 | SEABIRD | 911+ | 1222 | Yes | Yes |  |
| CTD 3 | SEABIRD | 911+ | 0443 | Yes | Yes |  |
| TSG | SEABIRD | 21 | 2488 | n/a | Yes |  |

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2019-067

| Calibration Information - 0334 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 22Feb2017 | Factory |  |  |
| Conductivity | 2128 | 23Feb2017 | Factory |  |  |
| ECO Fluorometer | 2215 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1483 | 2Mar2017 | Factory |  |  |
| Pressure | 0482 | 28Feb2017 | Factory |  |  |

Make/Model/Serial#:	SEABIRD/SBE911+/1222	Cruise ID#:		2019-067

| Calibration Information – 1222 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 03Jan2019 | Factory |  |  |
| Conductivity | 4395 | 01Jan2019 | Factory |  |  |
| Secondary Temp. | 4883 | 01Jan2019 | Factory |  |  |
| Secondary Cond. | 3184 | 28Dec2018 | Factory |  |  |
| Transmissometer | 983DR | 9Aug2017 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 01Nov2018 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | Factory |  |  |
| pH | 691 | 13Dec2018 |  |  |  |
| SeaPoint Fluor. | 3642 |  |  |  |  |
| Pressure Sensor | 1222 | 15Mar2019 | Factory |  |  |
| Altimeter | 43281 |  | Factory |  |  |

Make/Model/Serial#:	SEABIRD/SBE911+/0443 Cruise ID#:		2019-067

| Calibration Information – 0443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4888 | 14Feb2018 | Factory |  |  |
| Conductivity | 3531 | 06Mar2018 | Factory |  |  |
| Secondary Temp. | 4700 | 13Feb2018 | Factory |  |  |
| Secondary Cond. | 4513 | 13Feb2018 | Factory |  |  |
| Transmissometer | 983DR | 9Aug2017 | Factory |  |  |
| SBE 43 DO sensor | 3791 | 13Mar2019 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | Factory |  |  |
| pH | 691 | 13Dec2018 |  |  |  |
| SeaPoint Fluor. | 3642 |  |  |  |  |
| Pressure Sensor | 0443 | 31Aug2017 | Factory |  |  |
| Altimeter | 43281 |  | Factory |  |  |

Make/Model/Serial#:	SEABIRD/SBE21 /2488 Cruise ID#:		2019-067

| Calibration Information – TSG - 2488 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 14Nov18 | Factory |  |  |
| Conductivity | 2488 | 14Nov18 | Factory |  |  |