## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 January 2021 | ### Transmissivity:Green was recalibrated to correct an error in the original conversion. S.H. |
| 10 July 2020 | Bottle files were prepared for events 135, 136, 139, 144, 145. Transmissivity channels were reversed so that Red precedes Green. See end of report for details. |

## PROCESSING NOTES

Cruise: 2019-069

Agency: OSD

Location: BC Coast

Project: Moorings and Chatham Sound

Chief Scientist: Spear D.

Platform: John P. Tully

Date: 31 July 2019 –  12 August 2019

Processed by: Germaine Gatien

Date of Processing: 9 Oct. 2019 – 14 February 2020

Number of original HEX files: 108 			Number of CTD files: 107

Number of rosette files:	61				Number of bottle casts processed: 52

Number of original TSG files: 7 (only 2 usable)		Number of processed TSG files:  0

# INSTRUMENT SUMMARY

CTD #0506 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (#1883DG & 1185DR), a SBE 43 DO sensor (#3234), a SeaPoint Fluorometer (#3640) and an altimeter.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully Laptop.

The deck unit was a Seabird model 11+ #425.

A Guildline model 8400B Autosal serial # 68572 was used to analyze salinity samples.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book had no information about the science personnel or the thermosalinograph. There were good notes about problems encountered. The rosette log sheets were in good order with comments about problems encountered. This was especially helpful because there were many changes in station order made after sampling logs had been prepared and several cases of bottles fired accidentally and replaced by extra bottles being fired. One event # (92) was entered incorrectly in the rosette log but the Daily Science Log was correct.

There were 2 WetLabs CStar transmissometers in use during this cruise:

Channel Transmissometer refers to sensor #1883DG (530nm - green)

Channel Transmissometer2 refers to sensor #1185DR (650nm - red)

For comparison with other Institute of Ocean Sciences cruises, note that the

transmissometer wavelength is 650nm unless otherwise stated.

The configuration file used at sea had the transmissometer parameters reversed from the positions in which they were actually mounted. This was corrected.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The primary temperature and salinity channels were chosen for most casts, but for event #67 there was a problem with the primary conductivity so the secondary channels were prepared for archival. For event #168 archiving started when the CTD was down to 132db, so upcast data were prepared for archival and fewer decimal places are shown for temperature and salinity due to the lower quality expected from upcast data.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.50 mL/L from 0 to 50db

±0.30 mL/L from 50db to 250db

±0.06 mL/L from 250db to 1000db

±0.01 mL/L below 1000db

There were 7 Thermosalinograph files that covered the time of the cruise. Only 2 of the files contained useful data. However, the positions were found to be bad. Processing was suspended and no files are to be archived at this time. If appropriate positions are found, data will prepared for the archive.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity, and NH4 data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Since their last calibration most of these sensors have been used on 9 previous casts (just 8 for the DO sensor).

- There were 2 transmissometers mounted on the CTD and it was reported that they were mounted in reverse order to that indicated in the configuration file. After that was corrected the file was saved as 2019-069-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2019-069-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and there were some spikes and noisy patches in both salinity channels but they were not bottles that were sampled for salinity and were not large enough to affect average salinity significantly.

A preliminary header check was run and a problem was found in casts #82-169. They had been converted with the wrong configuration file, so the steps above were repeated for those casts.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Bottles were removed that have not been sampled.

There were many adjustments needed due to station order having changed after sample numbers were entered in the log sheets.

There were also several casts with sampling but no sample numbers so a CHE file is not needed. In a few cases bottles were fired accidentally at the wrong depth, so another bottle was fired at the planned depth. Good notes were available concerning those issues.

The addsamp file was then sorted on Event # and Sample #.

The file list was adjusted to remove events 52, 80, 82, 121, 135, 136, 139, 144 and 145.

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-069-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF 2019-069 OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2019-069oxy.csv. That file was converted into individual *.OXY files.

There were no DO comments starting with “ALL”.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2019-069 CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2019-069chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2019-069SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 38 to 45 days after collection. The files were simplified and saved as 2019-069sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2019-069*.xlsx. The file was simplified, saved as 2019-069nuts.csv and converted to individual NUT files.

AMMONIUM

NH4 data were obtained in spreadsheet QF_NH4_2019-69*.xlsx. This includes a precision study. The file was simplified and saved as 2019-069NH4.csv. This file was converted to NH4 files.

The nutrient and salinity data for event #92 were incorrectly identified as from #91. The files were renamed.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Some problems were found:

- The Chlorophyll samples 375-379 from event #104 were mislabelled as 103. That was fixed and the merge process was repeated.

- The NH4 data were missing from event 131. This was due to the samples being miss-assigned to event 130 in the analysis spreadsheet. This was corrected and the merge process was repeated.

Bottle data were once again exported to a spreadsheet and the corrections were checked and found to be ok.

##### Compare

Salinity

Compare was run with pressure as reference channel. 
Most of the samples are shallow and many, including all deep samples, come from the bottom of casts. It is unfortunate that more bottles were not fired below 500m during the deepest casts.

Outliers were excluded based on standard deviation > 0.0008 in the CTD salinity and differences between bottles and CTD salinity >0.015psu. The primary salinity was found to be low by 0.0051 and the secondary high by 0.0007psu with standard deviations of 0.006psu.

There were no extreme outliers, with the largest differences being right at the surface where there were large vertical salinity gradients so poor matches to the CTD data are expected. If bottles are not flushed well then we expect the CTD to look lower than the bottle samples except for bottles fired at the bottom, for which the opposite result is expected, especially if local gradients are high. Looking at only bottles from below 250db the primary CTD salinity is low by an average of 0.0064psu (standard deviation 0.0031) and the secondary is low by an average of 0.0004 (standard deviation 0.0034). The slopes near the bottom for these casts were not very high and in a few cases there was some vertical movement of the CTD before firing, so flushing errors should be fairly low. It is surprising that the results with deep bottles show CTD salinity to be relatively lower than in the comparison that included some shallower bottles that were not fired at the bottom of casts.

There were some comments about analysis problems and there was a 1.5month delay in analysis which may account for bottles reading a little high due to desorption of glass and/or evaporation of samples. These potential errors would make the sample salinity read a little high and thus the CTD may appear to read lower than it really is.

For full details for the COMPARE run see file 2019-069-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The largest outlier had a high standard deviation in the CTD data and is associated with the CTD being lowered during the upcast to fire a bottle that had been missed. The CTD DO does not appear to have had sufficient time to equilibrate, this being a level of high DO gradient.

When outliers were excluded based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0357 + 0.009

The fit when the sensors were last used during 2019-014 was:

CTD DO Corrected = CTD DO * 1.035 + 0.013

Values flagged 3 by the analyst were investigated. A few were minor outliers but not by enough that would warrant a change to the flag or added comment.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The ratio of FL/CHL versus CHL had the shape typical of these sensors with the ratio dropping as the CHL values rise. However, there is a lot more noise than usually seen with very high and very low CTD fluorescence.

##### Conversion of Full Files from Raw Data

All files were converted using 2019-069-ctd.xmlcon. The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. The T and C pairs were reasonably close during downcasts with upcasts quite noisy, as usual. Fluorescence and Dissolved Oxygen profiles looked normal. The “Green” transmissometer was generally about 8%/25cm higher than the “Red” one; the profiles had similar shapes. The altimetry looked noisy near the bottom though a reasonable value can be picked out.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 100

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 2 of deep casts to find differences between the pairs of temperature, conductivity and salinity channels. The differences are fairly small but there is evidence of drift.in conductivity and salinity. The salinity differences are consistent with those found in the bottle comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2019-005-0051 | 1000 | -0.0010 | +0.00015 | +0.0028 | High, V.Noisy |
| “ | 1900 | -0.0006 | +0.00015 | +0.0025 | “ |
| 2019-005-0159 | 1000 | -0.0009 | +0.00020 | +0.0034 | F.High, Noisy |
| “ | 1900 | -0.0007 | +0.00018 | +0.0030 | “ |
| 2019-006-0032 | 1000 | -0.0007 | +0.00025 | +0.0040 | High, XNoisy |
| “ | 1900 | -0.0004 | +0.00025 | +0.0035 | “ |
| “ | 2900 | -0.0003 | +0.00026 | +0.0035 | “ |
| 2019-006-0057 | 1000 | -0.0007 | +0.00027 | +0.0039 | High, Noisy |
| “ | 1900 | -0.0004 | +0.00026 | +0.0037 | “ |
| “ | 2900 | -0.0005 | +0.00026 | +0.0037 | “ |
| 2019-014-0002 | 1000 | -0.0008 | +0.00028 | +0.0040 | High, Noisy |
| “ | 1900 | -0.0003 | +0.00028 | +0.0035 | “ |
| “ | 2500 | -0.0002 | +0.00030 | +0.0048 | “ |
| 2019-014-0033 | 1000 | -0.0009 | +0.00034 | +0.0048 | High, Noisy |
| “ | 1900 | -0.0005 | +0.00033 | +0.0044 | “ |
| “ | 2500 | -0.0002 | +0.00032 | +0.0043 | “ |
| 2019-014-0049 | 1000 | -0.0009 | +0.00034 | +0.0050 | High, V.Noisy |
| “ | 1900 | -0.0003 | +0.00033 | +0.0045 | “ |
| “ | 2500 | -0.0004 | +0.00033 | +0.0045 | “ |
| 2019-069-0030 | 1000 | -0.0004 | +0.00041 | +0.0054 | High, Noisy |
| “ | 1900 | -0.0003 | +0.00040 | +0.0052 | “ |
| 2019-069-0051 | 1000 | -0.0009 | +0.00041 | +0.0059 | High, F Noisy |
| “ | 1900 | -0.0004 | +0.00039 | +0.0053 | “ |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run. Station names for casts 96, 121 and 158 were wrong, so were corrected. The event named as #20 was changed to #23.

The header check was run. No problems were found.

Cruise tracks were plotted. An error in the event # for cast #149 was found – that file name had been corrected earlier but the event # wasn’t; that was fixed and the cruise tracks were plotted again and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header. Some casts were found to have large check values but when the depth in the header was replaced with the depth in the log the values were small; in those cases the water depth value was changed to match the log. In one case the altimetry was very noisy and the header entry looks too high; it was replaced with an estimate based on the plot. In another case the water depth was missing in the log and clearly wrong in the file header; a historic value was entered.

The same changes were made in the SAM files for events #5 and 30.

The following information is available concerning the pressure accuracy.

- According to the log book data were logged while the CTD was coming up through the water line for a few casts. Because the pumps were not on it is hard to figure out when the CTD came out of water, but transmissivity suggests that the sensors did not record in air. Logging may have stopped while the rosette was mostly out of water, but not the sensors. The minimum pressures at the end of the casts ranged from 0.2 to 0.7db. This may suggest the pressure is reading slightly high.

- During 2019-001 the average surface pressure was 2.9db which is a little low for the Tully in open waters.

- During this cruise it was 2.2db. Many of these casts were in protected waters so that is reasonable. If we assume that the pressure is reading high by 0.5db, then for many of the casts the top of the CTD would have been out of water when acquisition began. This is unlikely

- During cast #104 there is data from 0.3db that appears to be in water.

- During cast 135, the minimum pressure is 0.96db With pumps on all values including transmissivity looks like the observations are from in water.

- So there may be a pressure error, but if so it is likely <0.5db.

There is some evidence that the pressure might be reading a little too high but it is not convincing. No recalibration will be applied.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts to assess if the settings used during 2019-005 should be used to align conductivity with temperature as judged by the effect on salinity as seen in T-S space. Those settings did look like the best choice for these data s well.

SHIFT was run twice on all SBE911 casts using -0.7 records for the primary and secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #101 which was aborted. That cast was removed from the file list and will not be processed further.

For cast #138 the SHFC1 file was put through REVERSE and then DELETE to produce a DELREV file.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature and conductivity sensors have been used on 8 cruises since they were last serviced. Two of the cruises did not have useful calibration sampling. No problems were found with pressure.

Results for salinity channels were:

- 2018-030 - the primary salinity was found to read low by 0.002 & the secondary by 0.001.

- 2018-032  - there were problems with the salinometer so the comparison was not trusted.

- 2019-003 and 2019-001 - primary salinity low by ~0.002 & secondary low by <0.001psu.

- 2019-016 – Results not trusted

- 2019-038 – Results not trusted

- 2019-005 – Primary low by 0.0023 & secondary very close to bottles

- 2019-006 – Primary low by 0.0032 & secondary low by 0.0006

- 2019-014 – Primary low by 0.0029 & secondary high by 0.0010

During 2019-008 and 2019-023 which followed this cruise:

- 2019-008 – Primary low by 0.0083 & secondary high by 0.0004psu.

- 2019-023 - Primary low by 0.0082 & secondary high by 0.0008psu.

Results for Dissolved Oxygen channel were:

The dissolved oxygen sensor was used on 7 previous cruises since the last service and there are preliminary results from the 2 cruises that followed this one. The fits against bottles have not varied much from those cruises with enough DO sampling for a good fit, with corrections having slopes that started at about 1.03 and increased to about 1.035 between May and June 2019. The slope was quite steady from June, July and August but was slightly higher for 2019-023 in September at 1.0384. The offsets vary between 0.010 and 0.028 with no temporal pattern.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Most salinity data fell within the climatology with the exception of casts in narrow inlets on Haida Gwaii where deep salinity values were low. Temperatures were frequently higher than those in the climatology at depth, including the Haida Gwaii casts and other inlet sites that were near shore in Chatham Sound and Hecate Strait. These excursions are most likely due to being compared to open waters in that general area as the climatology is not broken down into small enough blocks to represent these areas well. The open ocean casts show no evidence of a calibration problem.

Repeat Casts – The only repeat casts were too shallow for a useful check.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

There was more noise overall in the secondary channels and more unstable features in T-S space than with the primary channels.

A few casts had particular problems:

- Event #8 had bad sections in both the primary and secondary salinity channels during the downcast. There is a bad section in the primary between 12 and 21db and the primary temperature also looks poor. Just as it cleared there was a spike in fluorescence that marked the beginning of trouble in the secondary salinity that persisted throughout the downcast and likely through at least part of the upcast. The fluorometer was mounted on the primary pump and it also looks suspiciously low compared to the upcast. There were no bottle samples for this cast. The primary looks best overall but temperature, salinity and fluorescence points were removed from 12 to 21db.

- Event #67 had bad primary salinity between the surface and 12db. The secondary salinity looks ok.

- Event #168 is missing the first 132db of downcast data, so the upcast will be used. The ascent rate was fairly steady with no bottle stops so while the quality will be lower than from a downcast, it is not terrible.

The primary channels were selected for editing and eventual archival with the exception of cast #67 which had sections of poor data in the primary channels and the secondary data were better.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. There are some small unstable features that are likely real. No further editing was done.

##### Recalibration

No problem was found with the pressure calibration.

The primary salinity was found to read low by 0.0051psu and the secondary high by 0.0007psu. The primary difference is close to the last use while the secondary difference is larger. The difference between the two comparisons is consistent with the differences found between CTD channels.

File 2019-069-SBE911-recal.ccf was prepared to recalibrate by adding 0.0051psu to channel Salinity:T0:C0 and subtracting 0.0007 from Salinity:T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0357 + 0.009

This correction was first applied to the SAM and MRG files.

COMPARE was rerun for salinity to ensure that corrections were appropriate and they were. The primary salinity was found to be low by an average of 0.0002psu. See file 2019-069-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded based on residuals, the CTD DO was high by an average of 0.005mL/L. This result shows that the recalibration worked well. See file 2019-069-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When a few large outliers were removed based on differences and some others based on residuals the CTD DO was higher than the titrated samples by an average of ~0.04mL/L (standard deviation of 0.09mL/L). The largest differences between bottles and CTD are in the 3mL/L to 5mL/L range which is where the vertical gradients are generally highest. This is due to a combination of uncorrected flushing errors lowering the bottle data and slow response of the CTD DO leading to values that are too high in the high DO gradients. This is likely the best fit we can achieve.

No further recalibration was applied. See 2019-069-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #67 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

For cast #67 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------------

Conductivity, Transmissivity and Fluorescence data are nominal and unedited

except that some records were removed in editing temperature and salinity.

There were 2 WetLabs CStar transmissometers in use during this cruise:

Channel Transmissometer refers to sensor #1883DG (530nm - green)

Channel Transmissometer2 refers to sensor #1185DR (650nm - red)

For comparison with other Institute of Ocean Sciences cruises, note that the

transmissometer wavelength is 650nm unless otherwise stated.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for casts

far from shore. It is recommended that users check extracted chlorophyll values

where available.

Dissolved oxygen was calibrated using the method described in SeaBird Application

Note #64-2, June 2012 revision, except that a small offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be especially true

when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration)a rough comparison was made between

downcast SBE DO and upcast titrated samples. Some of the difference will be due to

problems with flushing of Niskin bottles and/or analysis errors, so the following

statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for are considered, very roughly, to be:

±0.50 mL/L from 0 to 50db

±0.30 mL/L from 50db to 250db

±0.06 mL/L from 250db to 1000db

±0.01 mL/L below 1000db

For event #67 only upcast data were available. Because the data are affected by passing

through the wake created by the rosette, temperature and salinity are reported with fewer

significant figures than usual.

For details on the CTD data processing see document: 2019-069_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; a few problems were found and fixed.

Profile and T-S plots were examined. There are some small unstable features in T-S plots but they are likely real.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from ~65% to 140%. Values were highly variable in the Chatham Sound area ranging from 75% to 135% and the same was true in the other nearby inlets. In Effingham Inlet they varied from 110% to 140%. The casts furthest offshore had values from about 103% - 104% with one cast at 108%. We typically see values between 103 and 105% far offshore near the surface. This suggests that the calibration correction was reasonable,  but we can conclude no more than that given the scatter.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and

One line was removed from file 2019-069-0046.ox2. Niskin #11 was not sampled and no sample # had been attached. There are other bottles fired with no analysis data available, but the records were kept as data may be added later or the CTD data may be needed.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the mission name to “Moorings and Chatham Sound”, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2019-069-bottles-final.xls. The entries were compared with the rosette log sheets. No problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files. No problems were found.

##### Thermosalinograph Data

A TSG45 was used for this cruise. It does not support having a fluorometer attached, so the TSG output has to be combined with that from a fluorometer. The intake is at about 4.5m.

There were 7 files provided but only 3 contain more than a few minutes of data. Those 3 files were renamed with format, cruise#-MMDD-HHMMSS.

The formats for positions and fluorescence are bad in the second file. Given it is very short (~25 minutes) no attempt will be made to correct these.

The files have extensions RAW but are in csv format, so the 2 files were opened in EXCEL and saved as CSV files. In opening use DELIMITED, deselect TAB, select COMMA and OTHER - *.

It was necessary to choose TEXT for the time on the 2nd page of the text import wizard.

The spreadsheet was adjusted as follows:

- 2 lines of headers were added – channel names and units.

- Replace NaN with -99. Look for any other odd pad values.

- The file was then saved in CSV format – the name is not critical at this point.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column M. Then a concentration value was calculated in column F using offset -0.04 and scale 13.1 as determined in the most recent recalibration of the fluorometer.

- Copy/Special Paste was used to save the values and then the voltage channel was removed.

- A file break column was filled with the cruise #-data/time info from the original file name. But some of the files were very large, so the file break was changed at the beginning of every second day to produce smaller files with names having format YYYY-###-YYYYMMDD-HHMMSS to show the opening date of the file.

- Time and Date formats are a problem – when converting from RAW choose TEXT but once open in EXCEL set Time Format  to HH:MM:SS and save the file again. (Any time the file is opened in EXCEL the time format may have to be set again before saving.) Using “custom” formats seems to work best.

The files were then converted to IOS format with header info added. There are 14 files, with the start time of each file given in the file name.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. (To do this tick “Add offset from Time Zero in Days” and enter Time Zero as the last day in the previous December at 00:00:00.)  A record number was also added to enable averaging (for use in comparison to CTD files).

There was a big problem with the 3rd file with some bad latitude and longitude values, ex 5.00E-04. There is no simple fix because the bad values are from spikes, with 1 low value followed by a high one. After conversion to IOS Header format a graphical editor was used to clean data by interpolation, but this showed that while the problem was deeper with all positions looking suspicious.

A few checks were made to see if the positions are reliable:

- Cast #20 CTD file positions on August 2nd at 19:39 were 50 3.58, 127 53.63

- An RMC ship file with positions was found and they were 5003.59 12753.63 – The formats are different but those 2 readings are very close.

- At the same time the TSG readings were 50.0097, -127.0104. This is a decimal format, so equivalent to  50 0.58, -127 0.62.

- Cast #169 at 22:44 on August 11th - near the end of the cruise - had much larger differences in positions with the CTD at about 49.5N/124.5W while the TSG and RMC-RAW files at about 54N/130.4W.

Clearly something went wrong with the recording of positions.

The TSG data are recorded approximately every 8.1818 seconds.

The RMC-RAW file has times and positions but contains the same errors as the TSG record.

The GPS data downloaded by the chief scientist were more coarsely and irregularly spaced than the TSG data to be used to replace the faulty positions.

GPS data sent from sea were recorded only every 15minutes.

All of these data sets have spacing that is either irregular or have intervals that are not even numbers of seconds.

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end.

a.) Plots

A track plot was produced and added to the end of this document.

Time-series plots were examined on screen.

For the first file there was no flow and no ship motion until 1827 UTC, so the first part of this file should be removed at the editing phase and the name changed to reflect the time the ship started moving.

The flow was off for part of the first and last files. No other problems were noted.

The flow rate was very steady from about 19:26 on May 22nd.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2019-069-ctd-tsg-loop-comp.xlsx. There were 90 CTD casts that overlapped with the TSG record but for the first cast the TSG loop flow was high and differences between intake and lab temperatures very high, so that was removed from the comparison. Two other casts had no temperature or salinity data available from the relevant depth, so those were removed leaving 87 points of comparison. The first TSG file did not overlap with any CTD casts.

Bin Average was used to average TSG files over 6 records (30s) and standard deviations were calculated. REMOVE was used to remove unwanted standard deviation channels.

Each file was opened in EXCEL and reduced to the just records required for comparisons with CTD casts.

The data needed for comparison with the CTD data were then exported to a spreadsheet (Time, positions, intake and lab temperatures, salinity, fluorescence and flow rates plus the standard deviations for all except time and positions.)

TSG Data were selected for the times of CTD casts and then added to file 2019-069-ctd-tsg-loop-rosette-comp.xlsx.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were 2 position differences ~0.001º while the median differences were both 0.00003º. The times were checked for the 2 largest differences and they were a good match. Most differences are very small, so the time and positions in the TSG files are reliable and the matches are good between the two systems. Small errors are caused by averaging and imperfect time matches.

c.) Comparisons

- Comparison of T and S from TSG and CTD data and Rosette samples

|  | lat diff | long diff | Tint-Tctd | Tlab-Tctd | Slab-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- | --- | --- |
| Average | -0.00003 | 0.00005 | 0.2297 | 0.7325 | 0.3785 | 0.63 |
| Median | -0.00003 | 0.00003 | 0.1921 | 0.7456 | 0.4061 | 0.46 |
| Std Dev | 0.00015 | 0.00027 | 0.2645 | 0.3337 | 0.3039 | 0.61 |
| Max | 0.00042 | 0.00107 | 1.3178 | 2.0104 | 1.2102 | 4.99 |
| Min | -0.00069 | -0.00113 | -0.3674 | -0.1530 | -0.3508 | 0.16 |

The intake temperature is higher than the CTD temperature by a median of 0.19C°, a much larger difference than we normally expect; during 2019-001 it was 0.05C° and even that was considered large. A plot of the differences against standard deviation in the TSG temperature does not suggest that noise in the data is an important factor.

The heating in the loop (Temp Lab – Temp Intake) had a median value of 0.50C° which is very close to the results of 2019-001 despite the intake waters being warmer so that we might expect less warming. Small flow rate changes might affect that. However, if the intake temperature is reading too high, then we should look at heating in the loop as being temperature in the lab minus CTD temperature. That approach suggests heating in the loop to be about 0.75C°. A plot of heating versus CTD temperature does show the usual pattern of less warming as intake temperature increases, though the temperature range is small.

The TSG salinity values were quite close to the CTD during 2019-001 (low by a median of 0.009psu and by 0.004psu from P8 to P26), but during this cruise they are high by ~0.4psu. Comparison of conductivity from a few casts show the TSG values to be higher than CTD values, but temperatures are also very different.

The ratio of TSG fluorescence to CTD fluorescence range from 0.2 to 5 with a median value of 0.46.

During 2019-001 there were serious problems with the TSG fluorescence as values fell relative to CTD values through the cruise. For this cruise no such problem is noted. The TSG is generally lower and less variable, but the pattern is similar. Note that the ranges for the 2 series are different. The TSG values are almost always lower than those from the CTD but towards the end of the cruise there are a few cases where they are higher.

- Comparisons of loop samples with rosette samples

There were just 5 cases of salinity sampling from the rosette that overlapped with loop samples. The differences ranged from the loop salinity being low by 0.0037 to high by 0.0276psu with a median showing the loop to be low by 0.0039psu.

There were 8 points of comparison for CHL and the range of the ratio Loop CHL/ROS CHL was 0.49 to 1.06 with a median value of 0.93. The lowest ratio included a loop sample flagged 36.

- Comparisons of loop samples and TSG data

There were 11 loop Salinity and 12 loop CHL samples of which 4 were taken underway.

The TSG fluorescence was lower than the loop CHL with the ratio of TSG FL/Loop CHL having a median value of ~0.4. When the ship was underway the ratio was slightly lower than when stopped. The loop CHL compares well with the rosette samples.

The TSG Salinity was higher than the loop salinity by a median value of 0.34psu; while stopped the difference was 0.32 and while moving 0.42psu. The loop salinity is higher than the rosette salinity samples by ~0.004psu, so the large difference between TSG and loop must be presumed to be primarily due to poor TSG data. Similarly the loop CHL is ~90% of the rosette CHL values so the

For details on comparisons see document 2019-069-ctd-tsg-loop-rosette-comp.xls.

d.) Calibration History

- During 2018-026 this system was used for the first time. The intake temperature sensor read higher than the CTD temperature by a median of 0.03C°. The lab temperature was higher than the intake temperature by more than expected based on previous experience with heating in the loop on the Tully. The salinity traces were quite smooth with only a few very small spikes. Salinity read lower than the CTD. The differences were not sufficiently consistent to allow recalibration. The data were archived but with a warning that there was insufficient calibration sampling to assess quality.

- The results from 2018-025 suggest that there were problems with flow in the loop.

- During 2018-034 Salinity was lower than the CTD by a median of 0.030psu. Lab temperatures were higher than CTD temperatures by a median of 0.66 C°. The intake temperatures were high by a median of 0.072 C° using all data or 0.058 C° in a “quiet section”. Fluorescence from the TSG was lower than the CTD fluorescence except in areas where CHL was low. The comparisons have a lot of scatter and recalibration was not attempted. The data were archived but with a comment about the intake temperature being further from the CTD data than expected and heating in the loop being higher than when the SBE21 was in use. Because of the limitations only 3 decimal places will be shown for temperature and salinity.

- During 2018-040 the TSG fluorescence was removed as it looked unreliable, starting with values close to the CTD but increasing until it was 60 times the CTD values. Salinity was lower than that from the CTD by a median of 0.032 and lower than loop samples by 0.034psu. The intake temperature was higher than the CTD temperature by a median of 0.0032C°.  The lab temperature was higher than the intake temperature by ~0.28 C°.

- During 2019-001 the TSG fluorescence values started unbelievably high, but the values dropped through the cruise getting closer to CTD and loop values. The TSG intake temperature was higher than the CTD by about 0.05C°. The lab temperature was higher than the intake temperature by ~0.5C°. Salinity was lower than CTD and loop by between 0.003 and 0.009psu.

e.) Conclusions

1. The TSG clock worked well.

2. There were 87 points of comparison with CTD.

3. After an initial period with no flow and then very high flow, the flow rate was very steady with values around 1.0L/min. That is close to the flows recorded during 2019-001.

4. The intake temperature is reading higher than the CTD by much more than usual, with a median difference of 0.19C°. The differences are largest near the end of the cruise. There were larger vertical temperature gradients for those casts, so a mismatch in depths can account for the increased differences. A number of casts were examined that had a differences ~0.2 and those mostly looked well mixed at the surface. So the local gradients don’t seem to support this large a difference unless the loop is drawing water from well above 4.5m.

5. TSG Salinity is reading higher than the CTD by a median of ~0.4psu except later in the cruise when the differences are lower, likely due to higher vertical salinity gradients so that drawing water from a little higher than the CTD depth would be significant. The TSG salinity is also higher than loop salinity by a median of 0.34psu. So TSG salinity is definitely reading too high. In other uses of this equipment it has always read a little too low, so this is a notable difference.

6. The lab temperature is higher than the CTD temperature by a median of 0.74C°. So heating in the loop is 0.74C° if we trust the lab temperature but not the intake temperature. If we trust both TSG temperatures heating in the loop is 0.5C° which is the amount of heating in the loop for 2019-001 when the flow rate was also ~1.

7. The history of the instrument shows the intake temperature generally reading significantly high. The only exception was for cruise 2018-040 when it was high by only 0.003C°. For that cruise the heating in the loop was also much lower.

8. We have much to learn about this new TSG. Could the intake temperature be affected by nearby ship equipment? Could there be both heating and cooling within the loop as the water passes through different environments?

9. The TSG fluorescence is reading lower than the CTD fluorescence but looks much better than for the past 2 cruises and does give a picture of variability.

10. The intake temperature will be recalibrated by subtracting 0.19C° as our best estimate. The salinity will be recalibrated by subtracting 0.4psu.

SUSPEND – Look at 2019-006 first.

g.) Editing

CTDEDIT was used to pad data in channels Temperature: Lab, Temperature: Intake and Salinity:T0:C0 from parts of the files from May 23rd? with 0 or very high flow rates. 
.

h.) Recalibration

A test was done to see if lowering the TSG intake temperature by 0.05C° and then deriving salinity using the CTD conductivity, the resulting salinity is close to the CTD salinity and it was within 0.004psu.

CALIBRATE was used to subtract 0.05 from the TSG intake temperature and to add 0.005 to the TSG salinity. These are based on rough estimates but will bring the TSG data into better line with CTD and loop data. Given the large temperature correction only 3 decimal places should be included in the intake temperature, and that is appropriate for the lab temperature and salinity too.

DERIVED QUANTITIES was run again, this time to derive sigma-T using the corrected intake temperature and salinity.

i.) Preparing Final Files

REMOVE was used to remove channel Pressure, Fluorescence:URU:Wetlabs, Temperature:Difference , and record #.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

Sampling System

---------------

The seawater is drawn from a 20-cm x 56-cm standpipe, flush with the hull

and located 4.5 metres below the water line, and pumped approximately 50

metres through a 5-cm stainless steel insulated pipe up to the main lab.

Between the intake and the pump is a 24 x 24-cm stainless steel strainer

(1/2-cm grid). The pump is a Moyno model 2L6 stainless steel progressive

cavity pump with flow and pressure control.

TSG Data Processing

-------------------

Comparisons were made between the TSG data and co-incident CTD casts, rosette

samples and loop samples. These indicate that the intake temperature is high by

roughly ~0.05C° and salinity is low by roughly 0.005psu. The small error in

salinity implies that Temperature:Lab and Conductivity are reasonably accurate.

Recalibration was applied to subtract 0.05C° from Temperature:Intake and to

add 0.005psu to Salinity.

Temperature and salinity are shown with 3 decimal places due to reflect

uncertainty in the calibration.

The only editing applied was to remove some data from a few small sections where

there was no flow in the loop. The flow rate was quite steady and the temperature

and salinity traces were free of spikes.

Fluorescence data from the TSG were removed from the files because the values were

unbelievably high early in the cruise. They gradually dropped through the cruise

even though there was little variation in extracted chlorophyll and they remained

higher than those from the CTD except for a few casts at the end of the cruise.

See report 2019-069_Processing_Report.doc for more details.

The files were saved as 2019-069-DDMM-HHMMSS.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

July 10, 2020 –updates to data

- The 2 transmissometer channels were reversed so that the Red comes before the Green. This was done so that if users choose the first channel it will correspond to the type of transmissometers used in previous cruises when Transmissivity:Green was not available.

- New CHE files were prepared for events 135, 136, 139, 144, 145. During those events there was no sampling from Niskin bottles, but there were surface samples. The CTD data in the files come from about 2m while the samples are from the surface.

Particulars

13, 14, 16, 17. Data logged while coming up through the water line.

23. File saved as event #20 – renamed as #23 to match log entry.

23. Forgot to turn off pumps. Recommunicated to turn it off while still in water.

30. Stop at ~370m to fix wire angle.

31. Touched bottom.

52. 2 bottles fired with no sample numbers – no processing required.

64. Altimeter 98 below 12m.

66. Altimeter 98 below 10m.

80. Cast aborted on upcast after bottle #3 due to hydraulic leak. No samples taken.

82. All bottles tripped to avoid hydraulic fluid entering – no samples taken.

88. Seasave crashed – restarted.

101. SAR call – cast aborted at surface– no bottles fired – no samples for IDs 336-347.

109. Lot of near-surface flocculent at CTD site CHAT2.

129. Strong current at station CH29.

135, 139, 144, 145. Bottle #1 tripped for Skeena sample at surface.

149. pause at 80m for wire angle adjustment.

152. Stn LBA in header – should be LBA1 (Corrected)

153. Stn LBA in header – should be LBA2 (Corrected)

164. Pause at 80m for wire angle adjustment.

168. Acquisition did not start until 135db of the downcast.

2019-069

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2023 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Conductivity |  |  | 1763 |  | 06Sep2017 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 5013 |  | 02Sep2017 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 3394 |  | 6Sep2017 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 21Nov2018 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 11185DR |  | 16Dec2018 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3234 |  | 16Feb2018 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3640 |  | Jan 2019 |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 2Sept2017 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 2Aug17 | Factory |  |  |
| Conductivity | 0620 | 2Aug17 | Factory |  |  |
| Temperature SBE38 | ? | ? | Factory |  |  |
| WETStar Fluorometer | 953 | 5May2017 | Factory |  |  |