## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2019-086

Agency: IOS, Ocean Sciences Division, Sidney, BC; Freshwater Institute, Central & Arctic Region, Winnipeg, MB.

Location:  Beaufort Sea

Project: Canadian Beaufort Sea Marine Ecosystem Assessment

Chief Scientist: Majewski A.; Eert J.

Platform: Frosti

Date: 31 July 2019 – 12 September 2019

Processed by: Germaine Gatien

Date of Processing: 22 June 2020  -  7 July 2020

Number of original HEX files: 70

Number of CTD files:  70	 	Number of bottle casts:   51

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (Arctic #1189) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#SCST1050DR), a SBE 43 DO sensor (#637), a SeaPoint Fluorometer (#SCF3651), a Biospherical PAR sensor (#70501), a surface PAR sensors (#20279), a SeaPoint turbidity meter (#11074), an Optode Dissolved Oxygen sensor (16) and an altimeter (#1161).

The data acquisition program was SeaSave V 7.26.7.107.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was not available. There were digital rosette sheets and Tap Log records with positions, times (BE,BO and EN), water depths and comments. There was also an electronic sampling log but it was incomplete. There was no list of equipment.

There were inconsistencies between the headers of the files and the Tap Log, mostly in station names, but the header entries are clearly correct in each case.

## There was 1 error in the configuration files used at sea. It would be helpful in processing to have the original calibration XML files from the factory either stored in a convenient place or included in the folder where the raw data are placed.

There were no event numbers included in the salinity QF file. Finding them was a tedious matter as some were entered in the Tap Log but many were not, so the digital log had to be searched. Some errors made in assigning the event numbers led to missing some salinity data in initial processing.

Deployment included a subsurface soak at around 5 to 7db. To ensure that the final files do not include the data from the initial drop to the soak depth it was necessary to remove them before running the program that removes upcast data. The # of records that should be removed from each cast was provided.

The wait at the surface after the initial soak was often very short. If it is safe to wait at least 30s that would enable the near-surface water to settle from the passage of the rosette on the way up.

There was a lot of salinity sampling in shallow water where vertical salinity gradients are high. Those bottles are not useful for calibration purposes. When only bottles below 500db were included in the fit both primary and secondary salinity were found to be close to bottles. There may be some error due to storage of samples for 3 to 4 months, but careful sealing likely minimized that.

There was no dissolved oxygen sampling. The difference in dissolved oxygen between using pre-cruise and post-cruise calibration parameters was very slight at <<0.001mL/L. While there are other sources of error due to sensor response, the surface saturation looked normal, so such errors are likely small.

An Optode dissolved oxygen sensor was used during the cruise but that channel is not included in the files to be archived. Special files including that channel were prepared for the use of the chief scientists.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

There were 2 distinct XMLCON files.

- From cast 1 to 364 all con files were the same.

- For cast #368 a new version appeared with out-of-date values entered. There was a problem during that cast with a cable twisting, so perhaps there was an associated computer crash that led to this error.

Parameters from the file used through most of the cruise were checked against the Arctic calibration summary spreadsheet, and the only problem was with one entry for the secondary conductivity. This was corrected in one file which was saved as 2019-086-ctd.xmlcon.

The rosette log sheets were obtained as well as the salinity analysis spreadsheet.

The cruise summary sheet was completed.

There was no history available for the pressure, temperature, conductivity and DO sensors since their last factory calibration and this was likely the first use for this equipment since the factory visit.

There were post-cruise calibrations for the temperature, conductivity and dissolved oxygen sensors. Configuration file 2019-086-ctd-post.xmlcon was prepared using the post-cruise parameters.

See §9 for a comparison of one file using the pre and post calibrations to estimate possible drift.

##### Conversion of Full Files from Raw Data

The deployment method was to spend a little time at the surface, then drop to between 5db and 8db with a soak at that level of about 2 minutes. The CTD was then raised to about 2db and after a short wait (sometimes only a few seconds), the full cast was run. The pumps were sometimes turned on at the beginning of the cast but more often during the initial drop.

All hex files were converted using 2019-086-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. A few features were noted:

- Turbidity and Transmissivity mirror each other.

- In low gradient errors the differences between temperature and conductivity channel pairs is small.

- Altimetry looks good.

- Descent rates were highly variable.

- A few casts were found where the pumps were still on as the conductivity dropped suddenly to near zero; the pressure at which that happened was between -0.05db and +0.2db.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2019-086-ctd.xmlcon.

The files were converted to IOS format.

CLEAN was run to create BOT files.

Temperature and salinity were plotted for all BOT files. Noisy data were investigated:

- Event #134 – Primary salinity was cleaned lightly around 340db.

- Event #205 – Secondary salinity was cleaned lightly around 60 and 75db.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged. Output was SAMAVG.

Next, the salinity analysis spreadsheet was examined to see what comments should be included in the header file. These were used to create file 2019-086-bot-hdr.txt; it may need further editing to reflect problems found during processing.

SALINITY

Salinity analysis was obtained in spreadsheet QF Cruise-2018-98 SAL v2019-12-02.xlsx. The sheet with final data for merging with CTD data was saved as 2019-086SAL.csv.. Event numbers were added based on the tap log for early casts and the rosette logs for later casts.

Comments were removed that were not needed, mostly stating there were duplicates since that is implicit in the flags.

The analysis was done in late November 2019 which was 3 to 4 months after collection.

The file was converted into SAL files.

The SAL files were merged with CST files with extensions MRG1.

The files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files, to produce MRG files.

The MRG files were put through CLEAN to produce MRGCLN2 files; 0s were entered into any empty flag channels.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was a lot of salinity sampling but most was in shallow water where the standard deviations in the CTD data were very high during the 10s window around firing time. Above 100db the CTD salinity is mostly lower than the bottle salinity. This is likely due to incomplete flushing of Niskin bottles in the presence of large vertical gradients.

Major outliers were identified:

- Sample #363  Event #245 – 20m – high std dev in CTD salinity; difference 0.75psu. There is no sample noted on the rosette log and no value in the salinity spreadsheet. This is a case of 2 bottles fired at 20m but only one sampled. This is a result of choosing to do the comparison based on pressure, not Niskin #. The second entry is not valid.

- Sample #385 or 386  Event #259 – 20m sal sample not mentioned on rosette log. There were 2 bottles fired at that level but only 1 sample. This is a case of 2 bottles fired at 20m but only one sampled. This is a result of choosing to do the comparison based on pressure, not Niskin #. The second entry is not valid.

- Sample #534 Event #353 – 30m – low std dev, difference CTD-Bottle = +2.34

& Sample #536 Event #353 – 10m – very high standard deviation in CTD data – difference -3.4

The 2 samples from event #353 were noted on the Autosal analysis log sheet as both coming from Niskin 15.  A choice was made by the analyst about which was the real Niskin 15 sample and which was really Niskin 13, though no flag or comment was entered. It looks like the assignment of order was wrong and the sample now named 534 should be 536 and 536 should be 534. If that is done the differences become -0.2 and -0.9psu.

So 2 of the outliers are of no significance. The salinity values for samples 534 and 536 were reversed in the salinity file and the merge process rerun for event 353. The differences are still large, but reasonable  for the depths concerned. After changing the values for those 2 samples and the discovery of some missing salinity samples, COMPARE was rerun.

Bottles above 200db and cases where the standard deviation in the CTD is >0.0008psu were excluded from the fit. One other bottle was removed that was a minor outlier but it was from the bottom of a cast where incomplete flushing could lead to low bottle values. The primary salinity was found to be high by an average of 0.0005psu (std dev 0.007) and the secondary salinity was high by 0.0007 (std dev 0.006). This indicates that the two CTD channels were very close while stopped, but in motion they appear to differ by about 0.0013psu as shown in section 9.

A closer look at the differences show that below 500db the differences between the two channels were larger. Including just those bottles, the primary salinity was found to be higher than bottles by an average of 0.00017psu and the secondary was low by an average of 0.00115. So the difference is 0.00132 psu. There were 4 bottles fired at the bottom and 7 well above the bottom. The slope of the fits are the same, so there is no evidence of significant pressure dependence in one channel – either both had the same dependence or there was no notable dependence. The latter is most likely. These differences agree with the differences noted between the sensors while in motion during downcasts from 500m down.

While the differences are small, the standard deviations are quite large. There are a number of potential errors in this comparison. Both delayed analysis and flushing errors contribute to the CTD looking lower than it really is except at the bottom of casts. Those errors should be fairly small below 200db as the vertical gradients are lower. At the bottom there are 2 sources of errors that both lead to the CTD salinity looking high. There are incomplete flushing errors and shed wakes carrying water from above can fill the Niskin bottle as the CTD slows down. Both errors would make the CTD salinity look high.

To get an idea of how large these errors might be, the data in the 47 bottles in the main fit for primary salinity were divided into bottles fired at the bottom and all other bottles.

- There were 26 cases of bottles fired above the bottom during the upcast. The primary salinity was low by an average of 0.0012psu with a standard deviation of 0.0069psu. Below 400db it was low by an average of 0.0001psu and standard deviation of 0.0021psu.

- There were 21 cases of bottles fired at the bottom. The primary CTD salinity was high by an average of 0.0049psu with a standard deviation of 0.0123psu.The standard deviation is quite large. A plot of differences versus depth showed small errors below 400m with CTD salinity high by an average of 0.0007psu and standard deviation of 0.0008psu.

The errors in the bottles fired at the bottom are larger and more consistent in sign than for those fired above the bottom. But for both sets the variation is small below 400db, with the CTD salinity channels within 0.002psu. However, there is one other source of error that must be considered. The storage of samples for 4 months could lead to higher bottle values so the CTD could be reading higher than this comparison suggests. That is not likely to be a large error due to careful sealing of bottles.

A plot of differences versus file pair number indicated very little temporal drift.

For full details for the COMPARE run see file 2019-086-sal-comp1.xls.

It was later discovered that salinity samples for events #51 and #290 had been labelled incorrectly, so they were missed in the COMPARE run. The bottle files were rebuilt for those 2 casts. Both casts were shallow and not likely to affect the conclusions reached above, but to be sure of that, COMPARE was run with them present. The points from just those 2 casts were plotted in red and they did not stand out as significantly different from the other data.

For details of this check on the 2 casts see file 2019-086-sal-comp2.xls.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 4 casts. Noise in the both temperature and dissolved oxygen make the tests difficult. The best setting varied from feature to feature, but a setting of +3s looked best overall.

ALIGNCTD was run using a setting of +3s to the dissolved oxygen channel.

##### CELLTM

The usual tests for the best settings for this routine are difficult to interpret because of the noisy upcast data, so the default settings was used.

CELLTM was run using (α = 0.0245, β = 9.5) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between sensor pairs were studied for 2 casts.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2019-086-0118 | 500 | +0.0002 | -0.00008 | -0.0013 | High, F.Steady |
| 2019-086-0118 | 900 | +0.0002 | -0.00013 | -0.0013 | High, F.Steady |
| 2019-086-0304 | 500 | +0.0002 | -0.00010 | -0.0012 | F High, Very Noisy |

One deep cast (#118) was put through the same steps but using the post-cruise calibrations for T, C and DO.  Using the post-cruise calibrations the difference between channels is larger especially for salinity which had primary values lower than secondary by ~0.006psu. Comparing values from the pre and post cruise calibration the secondary salinity is just slightly higher but the primary is lower by about 0.005psu. Since the comparison with bottles does not suggest significant difference between those channels and they appear to be close through the cruise, the calibration drift in the secondary conductivity must have occurred late in this cruise or in later use or in transport.

The change in dissolved oxygen calibration was small with the post-cruise values lower by <0.001mL/L.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

Conversion failed for event #47 due to an array overflow in the altimetry array; the altimetry channel was removed and conversion was then successful. The altimeter did not get within 15m of the bottom, so there is no loss of information as the value would not go into the headers anyway.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

An initial cross-reference list was examined  A few inconsistencies were found between the headers and tap log, but it was clear that the headers were correct based on positions and track plots. (Ex. Event 297)

However, the file saved as 2019-086-0338 corresponds to event #339 in both the tap log  and the rosette log. Event #338 was an MPS event. The file name was changed and the event number corrected in the profile files. The bottle file had to be recreated since the salinity sample had the correct event number for the ROS file did not. Because of this error some salinity data was missed from the 1st run of COMPARE but it was included in a rerun of COMPARE run after corrections to some other samples.

A header check was done. The only problem found is that fluorescence appears to go off-scale for cast #110 between 40.9db and 43.8db. This will be examined further after DELETE is run.

Track plots – The cruise tracks were plotted and look ok; they were added to the end of this report.

The surface check was run on the files before CLIP was applied. The average value was +0.115db with a minimum value of -0.02db. Examination of the ends of a few casts shows no evidence of significant pressure calibration drift. Pressure was checked at the end of a few casts. Often pumps were turned off after a 5db rosette firing, but in a few cases there were data from about 0.1db with pumps on. For event #323 it appeared that the primary channels were in water and the secondary very close to the surface when pressure was 0.1db. So pressures are probably within the expected accuracy.

##### CLIP

##### A list was provided of the # of scans that needed to be removed from each cast to remove the soak data;

CLIP was run to remove the soak period for each cast. Plots were made to ensure an appropriate number of scans were removed; a few casts were rerun until all looked right.

It was discovered that the bottom depth had not been entered in the headers of most files and those for which it was entered did not always look reasonable. So bottom depths were from the tap log were entered in a csv file. Water depths will be added to the files later.

The altimeter readings from the headers of the CLIP files were exported to spreadsheets. A check was made by calculating: Check Value = Max Depth Sampled + Altimetry header – Water depth. The 6 casts with the check value >4m were investigated. The altimetry was plotted to see if the entries were affected by spikes but all the traces looked good and the header entries appropriate, keeping in mind that the values are averaged over the bottom 2m so tend to be slightly high.

4 of the casts were rosette casts so the rosette logs were checked. Most have water depths recorded that are close to the header entries. But for cast #192 which had the largest check value, 22.6, the log notes that the depth started at 198m and got deeper through the cast whereas the header says 192m. So when the CTD was at the bottom the depth was likely notably higher than 198. The header entry was changed to 215m based on the maximum depth sampled plus altimetry reading. The other cases studied may have similar problems due to shoaling. Other depths may also be affected by variations through the cast, but the check values are not huge and there was no note made of depths varying through the casts, so they were not adjusted.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel in all casts using the usual advance of +24 records. Plots made after this step were hard to interpret due to noisy fluorescence and many stops for bottles but there is clearly some improvement in matching the fluorescence offset to the temperature offset.

Dissolved Oxygen

Channel Oxygen:Voltage:SBE was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel; it is difficult to judge but the alignment looked reasonable.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with the settings -0.7s for the primary and -0.4s for the secondary. SHIFT was run twice using those settings.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerns a problem during an upcast of 1 cast so will not affect downcast data.

The output files were copied to *.EDT.

Fluorescence values for cast #110 are close to the rated maximum between 40.9 and 43.2db and show little variability, so it is assumed the values were off-scale. Ultraedit was used to pad those values.

##### Other Comparisons

Previous experience with these sensors –

There was no history available for the pressure, temperature, conductivity and DO sensors since their latest factory calibration.

Historic ranges – Local climatology was not available.

Repeat Casts – The only casts that are close together are too shallow to test repeatability.

Post-Cruise Calibration – There were post-cruise calibrations in October/November 2019 for the temperature , conductivity and dissolved oxygen sensors. Temperatures were slightly higher in late 2019. The primary conductivity (and hence salinity) were higher with little difference in the secondary. The study in section 9 shows that the salinity differences were larger than observed during this cruise, so it is assumed the drift occurred either late in this cruise or later, possibly during transport. The drift in dissolved oxygen was very small.

##### DETAILED EDITING

The primary temperature and salinity channels were edited for all casts since they looked smoother overall in T-S space than the secondary channels.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and misalignment of T and C and records corrupted by shed wakes. For most casts the editing was mostly at the top and bottom of casts.

All files were edited. The output files were copied to EDT.

T-S plots were examined and the results looked ok; there are some small unstable features in a few casts. There was no obvious instrumental cause of some of these features so they may be real. In other cases the descent rate is low and T and S values reflect small variations in descent rate; it is impossible to determine which data are better.

##### Initial Recalibration

There was no Dissolved Oxygen calibration sampling so before making a decision on recalibration surface saturation of DO was derived and plotted.

CHANGE UNITS was used to derive dissolved oxygen in mass units.

DERIVED QUANTITIES was used to calculate DO surface saturation.

Plots were made of DO surface saturation and values ranged from 96% to 104% with most between 100% and 101%.

Summary of calibration review:

- Pressure checks indicate no recalibration is needed.

- The primary salinity appears to be very close to bottles, though if there was any evaporation or desorption in the samples due to delayed analysis, the CTD salinity may be reading a little high. The differences between primary and secondary in the post-cruise calibration are very different from those observed during this cruise, and that is confirmed by the bottle comparison, so drift noted likely occurred after this cruise.

- Dissolved oxygen values drifted little between calibrations. There may be errors due to other effects, but without bottle sampling no estimate is possible for that. Surface saturation values look normal, so calibration drift is likely low.

No recalibration was applied.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Secondary, Salinity:T1:C1, Oxygen:Voltage:SBE, uploy0, Altimeter, Status:Pump, Descent_Rate and Flag from all casts

REORDER was run to get the two DO channels together.

Routine “Merge CSV File to Headers” was rerun using file 2019-086-depths.csv to add depth to the CLIP files.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

---------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, Turbidity, PAR and Surface PAR data are nominal and

unedited, except that some records were removed in editing temperature and salinity.

Salinity:T0:C0 values compared well with bottle samples. A post-cruise calibration

suggested values might have drifted higher but it is likely most drift occurred

after this cruise since the difference between the 2 salinity channels was much

larger than observed during this cruise. No recalibration was applied

There was no dissolved oxygen calibration sampling but the post-cruise calibration

indicated little drift and the surface saturation looked normal.

While the CTD fluorescence data are expressed in concentration units, they do not

always compare well to extracted chlorophyll samples, particularly for casts far

from shore. It is recommended that users check extracted chlorophyll values where

available.

For details on the processing see the report: 2019-086_Processing_Report.docx.

The Standards Check routine was run and no problems were found.

A header check was run on the final files and no problems were found.

Plots of each file were examined and no further problems were found.

A second set of files were produced that included channel Oxygen:Optode:Raw for the use of the Chief  Scientists.

##### Final Bottle Files

No recalibration was required.

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run to remove the same channels as for the profile files.

A second SBE DO channel was added to the CTD DO with mass units and REORDER was run to get the 2 SBE DO channels together.

Routine “Merge CSV File to Headers” was rerun using file 2019-086-depths.csv to add depth to the  REO files.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files and a number of problems were found. The data were rerun through all steps since the turbidity channel was missing and some salinity data were missed due to errors in event numbers. The check was then rerun and no further problems were found.

##### Producing final files

Cross-reference listings were produced for the CTD and CHE files.

The sensor history was updated for the conductivity sensors.

Particulars –

68. CTD went to 50m with pumps off, returned to surface and started again.

110. Fluorescence went off scale – pad values entered.

128. No return to surface after 5m soak, so no useful data above 5m.

134. Only came up to about 5m after soak.

339. File saved at sea with wrong name – had event #338. Renamed in processing.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 1189 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 1189 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 5830 |  | 06Nov2018 |  | Factory |  | 30Oct2019 | Factory |  |
| Conductivity |  |  | 4327 |  | 6Nov2018 |  | Factory |  | 26Nov2019 | Factory |  |
| Secondary Temp. |  |  | 5831 |  | 04Nov2018 |  | Factory |  | 30Oct2019 | Factory |  |
| Secondary Cond. |  |  | 4339 |  | 06Nov2018 |  | Factory |  | 26Nov2019 | Factory |  |
| Transmissometer |  |  | 1666 |  | 14Jun2018 |  | Laurier |  | 21May2020 | IOS |  |
| SBE 43 DO sensor |  |  | 637 |  | 7Nov2018 |  | Factory |  | 21Oct2019 | Factory |  |
| Optode |  |  | 16 |  | 2018 |  | Factory |  |  |  |  |
| PAR |  |  | 70501 |  | 4Apr2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3651 |  | Feb 2108 |  | Factory |  | 21May2020 |  |  |
| Surface PAR |  |  | 20279 |  | 4Apr2016 |  | Factory |  |  |  |  |
| Pressure Sensor |  |  | 0130015 |  | 23Dec2013 |  | Factory |  |  |  |  |
| Altimeter |  |  | 1161 |  | Mar2019 |  | Factory |  |  |  |  |
| Turbidity Meter |  |  | 11074 |  | 23Mar2007 |  | Factory |  |  |  |  |