## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2019-124

Agency: PBS, Ecosystem Science Division, Nanaimo BC

Chief Scientist: King J. /  Boldt J.  		Platform: Nordic Pearl

Location: West Coast Vancouver Island		Project: IPES Synoptic Survey

Date: 15 June 2019 – 15 July 2019

Processed by: Germaine Gatien

Date of Processing: 25 May 2020 – 1 June 2019

Number of original HEX files: 58 (no Hex files found for 4 casts)

Number of CTD files: 58				Number of BOT files: 69

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0456) was used with temperature sensor #5130, conductivity sensor #3500, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1592, pH sensor and pressure sensor 0668.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with comments about problems encountered. There were also notes provided describing sampling methods. Times in the log are in PST which was clearly noted.

It was noted by the CTD operators that the CTD configuration was problematic for small ships that are subject to rough and complicated deployments. The plumbing with the upside-down V tubes can come off the oxygen sensor easily and could be lost entirely. As well as zap strapping to each spigot, the tubing should be secured so that it can’t float away if it does come off a spigot.

It is very helpful if header information is added to the raw files when they are uploaded. There is a specific format that works well:

**STATION JS03

**Latitude: 50 26.212 N

**Longitude: 126 04.403 W

**Depth: 185m`````````

During this cruise that information was added to files that had been converted, but not to the raw hex files. Since the hex files had to be converted again in processing, the header info addition was not useful. While it is useful to convert files at sea to ensure sensors are working properly, this step will most likely be repeated later with particular settings and channel choices that the processor will select. So there is no point in adding headers to those files – they need to go into the hex file.

The chief scientists provided the header information in spreadsheet format which enabled addition to the CTD files after some format manipulations.

The CTD was soaked at about 15m for long enough to ensure equilibration of dissolved oxygen and salinity. However, when the CTD was returned to the surface it was usually lowered again immediately and frequently appears to have passed through the wakes created during the upcast. That led to a considerable loss of data in the top 5m. It is recommended that there be a wait of about 30s at the surface, though in rough conditions this may not be practical.

A 3L Niskin bottle was placed 30 to 60cm above the CTD and was fired when the CTD was at roughly 3.5m. If the Niskin failed to close a surface bucket was used to gather the surface samples. Where available, upcast data from the CTD from 3m were extracted to accompany the Niskin bottle data. The raw files for 2 casts were missing but there were files for both that had been converted at sea; those could not be used to produce files for the archive since conductivity was missing and that channel is necessary to make adjustments to salinity. However, the temperature and salinity data were considered good enough to provide data to go with bottle samples that were taken during those casts. For 2 bucket samples there were CTD data available near the surface. For the bucket samples with no associated CTD cast, data from an RBR CTD that was mounted on the trawl net may be useful. Those data have not been added to the BOT files.

There was no salinity or dissolved oxygen sampling during this cruise, nor during any other cruise using this equipment since the last factory calibration. No recalibration was applied.

All extracted chlorophyll samples were taken near the surface. The fit of CTD fluorescence versus extracted chlorophyll was good. SBE Fluorescence was about 90% of CHL overall but it reads lower for high CHL values (>10ug/L.)

The file names were originally non-standard but have been corrected; 4 digits are needed for the event numbers in the file names.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were in standard format.

Notes were provided outlining 2 sampling procedures:

- Usually a Niskin bottle was fired at the surface, mounted 30 to 60cm above the CTD.

- If the Niskin failed to close a bucket was lowered over the side and surface water was retrieved.  CTD data are also available from an RBR CTD mounted on the trawl net, but will not be processed as part of this job.

2. Preliminary Steps

The Daily Log and notes on CTD deployments were obtained.

The deployment method used was as follows:

- The CTD was switched and the CTD was put in the water. It was taken down to about 15m and after about a 1-minute soak it was returned to the surface. The full cast was then started. The data acquired during the soak will be removed later. On the way up the CTD was stopped at about 3 or 4m after a wait of about 20 to 40s a Niskin bottle was fired.

- On some occasions the Niskin bottle failed to close so a bucket sample was taken.

The cruise summary sheet was completed.

The file names were non-standard; event numbers should have 4 digits. They were adjusted after conversion.

## 3. Conversion of Raw Data

The configuration file used at sea was not available. A file from a previous use of most of these sensors and an XML file for the pH sensor were found. A new con file was created from those elements with sensor channels in the same order as in the CNV files. It was saved as 2019-124-ctd.xmlcon

The parameters in the configurations were compared and all are the same as used at sea.

The HEX files were converted using configuration file 2019-124-ctd.xmlcon.

Plots show that the channels all produced reasonable values.

- The descent rates look very high and sometimes extremely noisy. The high rate does reduce the amount of data acquired so the bin-averaged data will be based on fewer observations.

- Fluorescence is very close to 0 in deep water so that instrumental noise leads to very small negative values. After bin-averaging, any negative values will be padded.

- The pressure in air appears to be between -0.3 and -0.6db at the beginning of casts and between 0 and -0.6db at the end. This reading is made complex by having some water in the plumbing even when pumps are not running. These readings are well within the specifications for this CTD type.

4. WILDEDIT

The only spikes in the data occurred at the beginning or end of the casts, so they will be removed in the normal course of editing. So WILDEDIT was not run.

5. WFILTER

Test were run to find a suitable setting to remove steps in pressure without smoothing the pressure too much. All choices helped but size 9 was best at removing unstable features in T-S and in removing small pressure reversals. WFILTER was run using a cosine filter, size 9 on the pressure, depth, temperature and conductivity channels.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look good.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. One cast was missing its latitude header; that was added and the file was reconverted.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

10. Checking Headers

A cross-reference list was produced and compared with the log book entries. There were no significant discrepancies.

The header times are in PST so ADD TIME CHANNEL was run to add 8 hours.

The next step was to remove the data collected during soaks at 15m. Plots were examined to see how many records need to be removed from each cast. CLIP was run to remove records from each cast.

Plots were made after this step and CLIP was rerun on a few casts because too many or too few data had been removed. The # of records removed was between 1150 and 2100 records.

Merge CSV files to Headers was used to add header information from file 2019-124-header-merge.csv.

HEADER CHECK was run. There were some negative values in fluorescence and pressure; those for pressure will likely be removed by DELETE. This will be checked again later. No other problems were noted.

The surface check shows an average of -0.6db before CLIP was run and most of the associated salinity values were extremely low. The pressure sensor has a 2800db range so the resolution is only 0.45db and accuracy is ±2.6db so the results are well within specifications.

##### 11. SHIFT

Conductivity

During 3 cruises in 2017 when the same sensors were used no adjustment was found to improve the alignment of conductivity with pressure. During 3 cruises in 2018 and during 2019-007 a setting of +0.5 records was found suitable to align the conductivity with pressure. Tests on the 4 casts data from 2019-124 show that there is little difference for most casts, but a shift of +0.5 records improves some.

A shift of +0.5 records was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is good.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. The alignment looks good. No further adjustment was made.

pH:SBE

This sensor has no history of use with this CTD. During 2019-050 on an SBE911 a shift of +35 records was used. Given the difference in scans per second this is equivalent to about +12 records.

Tests were run and the best results were with +10 records.

SHIFT was run on all cast to advance pH by +10 records.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

T-S plots were examined and 11 casts looked good without any editing.

For the others CTDEDIT was used to remove records corrupted by shed wakes and to clean salinity where unstable features looked likely to be caused by misalignment of T and C.

Notes of editing details were made in the headers.

Most of the editing was applied to the top 5m where it appears that the very short wait at the surface meant the CTD passed through its own wake from the upcast. A longer wait at the surface is recommended, say 30s, though in rough weather this may prove impractical.

As noted at sea there are serious problems salinity and DO for cast #179 and at the surface of #181, so those 2 files were put through REVERSE and DELETE to see if upcast data looked useful. For #179 the upcast was edited, but for #181 the downcast looked ok.

T-S plots were examined after this step and no significant problems were found; there are small unstable features but those are in areas where some instability is expected.

##### 14. Initial Bottle Data Steps

There was no rosette available for this cruise. To enable searching of bottle data, BOT files were prepared.

Samples were taken at the end of CTD casts when the CTD was between 2 and 4db, usually with the CTD moving a lot during the stop. In most cases the average depth of the stop was ~3.5db. The Niskin was about 30 to 60cm above the CTD. So choosing CTD data from about 3db looks appropriate. Upcast data were selected; downcast data are often of higher quality but for this cruise the data in the top 5m do not look very reliable. For those with sampling from a bucket there is usually no CTD data available above 1m, but for cast #70 and #181, there were data available from above 1m.

Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2019-124-bot-hdr.txt which will be updated as needed during processing. All of the spreadsheets included a precision study.

- Extracted chlorophyll and phaeo-pigment data were obtained in file QF_2019-124 CHL*.xlsx.

- Nutrient analysis was obtained in spreadsheet QF_2019-124 NUTS*.xlsx.

The merge header file was used to provide a framework for a spreadsheeet combining CTD data and chlorophyll and nutrient sample data. Extra lines were added for cases where there was no CTD cast available for comparison with bucket samples.

For cases where there was a CTD cast but the water was gathered from a surface bucket, flag 2 was added to the CHL and Nutrient data together with a comment. For one CHL sample a flag 5 was replaced with flag 2 as the 5 was clearly an error as it was only referring to a missing replicate sample.

For cases where there was no CTD cast, only a bucket sample, a comment was added.

The file was saved as 2019-124-bottle_plus_CTD.xlsx.

There was a discrepancy between the sample numbers for 2 casts with nutrients spreadsheet showing  #79 and #259 while extracted chlorophyll had #77 and #260; since the nutrients matched the log entries, those numbers were used.

For cases where there were no CTD casts and no entries in the event log provided, entries for times and positions were obtained from the log book. Positions were changed to decimal format.

Where CTD casts are available, files created just before running DELETE were put through REVERSE, DELETE, FILTER, Bin-Average (0.5m bins) and THIN. Data from 3m were then exported to a spreadsheet.

There are a number of ways in which CTD data were selected for the spreadsheet:

- For most CTD casts the data were taken from the upcast at about 3m since downcast data were often very noisy at that depth. Fortunately, the CTD was moved vertically during the “stop” so that both the ambient waters and the contents of the Niskin are likely quite well-mixed; so an accurate match in depth is not so critical.

- For the 2 casts (#70 and #181) that do have CTD casts but samples came from surface buckets,   CTD data from the top 0.5m were selected.

- For cast #120 there was no upcast data from an appropriate depth, so downcast data were selected.

- For cast #15 there was only a file converted at sea and it lacks channels needed for processing; however, temperature and salinity data were selected for the spreadsheet with fewer significant figures to reflect that data quality may be lower.

- For cast #265 there was no hex file. There was a file converted at sea, but it had no data near 3m.

A 6-line header was added to the spreadsheet and columns rearranged to match the headers.

This file was saved as 2019-124-bottle-plus_CTD-6linehdr.csv.

The file was also saved as 2019-124-bottle-plus_CTD-6linehdr.xlsx and used to do a comparison of SBE:Fluorescence and Extracted Chlorophyll.

##### 15. Compare

Fluorescence

The fit of SBE Fluorescence versus Extracted Chlorophyll looks typical of this type of fluorometer with the Fluorometer reading close to the fluorometer for CHL< ~10ug/L and then drifting somewhat lower for higher CHL. The average ratio of SBE fluorescence to CHL is 88% .

16. Other calibration checks

Sensor History – These sensors were used for 7 previous cruises since they were last calibrated. There was no calibration sampling available for any of them. No obvious problems were noted.

Historic Ranges – Where local climatology was available temperature and salinity fell within the historic values.

17. CALIBRATE

No channels were recalibrated

18. Fluorescence Filter

A median filter, size 5, was applied to the fluorescence data.

19. Bin Average and REMOVE

The files were bin averaged using 1db bins.

There are a few negative fluorescence values at depth so CLEAN was run to replace those with pad values and the channel limits were updated.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 70 to 180%. The lowest saturation rates were in Johnstone Strait which is very well mixed and the highest were from a few July casts in the southern section. The near surface data were very noisy so these values may not be indicative of real surface saturation levels. As well-near-shore casts don’t give us a very good idea about the calibration of the dissolved oxygen sensor.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add comments about processing.

Salinity format was set to F8.3 given the lack of calibration sampling over the time since the sensors were last calibrated at the factory.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

##### 20. Final BOT file preparation

File 2019-124-bottle-plus_CTD-6linehdr.csv was converted to IOS Header files for each cast.

CLEAN was run to set start and end times.

ADD TIME CHANNEL was run to add 8 hours to the header time and change the time zone entry to UTC.

Derived Quantities and REORDER were used to calculate DO in mass units and get the 2 DO channels together in the channel list.

HEADEDIT was used to add comments and to remove the END time since it is the same as START TIME.

Temperature and Salinity formats were set to F8.3 since the pressure match with bottles is approximate, the data are quite noisy during the stop and no calibration sampling is available.

Data were exported to a spreadsheet to check for problems and it was found that 0s had not been added to empty quality flag channels. CLEAN was run again to do that.

The final files have extensions BOT.

The following note was added to file 2019-124-0135.BOT: “There were 2 casts at this site. Acquisition failed on the first attempt but the bottle sample was from the 1st cast.”

The standards check was run and repeated until all errors were addressed.

The formats for temperature and salinity are non-standard but that was done deliberately to reflect the probable errors due to a mismatch between CTD and sample depth and noisy CTD data.

A cross-reference list and a header check were run on the BOT files; no problems were found.

Track plots were produced and placed at the end of the this report. They include all CTD casts and those  with only bottle sampling.

PARTICULARS – notes from logs and reports

General notes:

2 sampling methods were used:

- Usually a 3L Niskin bottle was fired at the surface, mounted 30 to 60cm above the CTD.

- If the Niskin failed to close a bucket was lowered over the side and surface water was retrieved. If no CTD data are available at the appropriate depth, data from an RBR CTD mounted on the trawl net could be used.

Conditions were rough during the first leg so descent/ascent rates of the CTD were highly variable.

Times are in PST in the log which matches the computer clock. Times in CTD and BOT files are UTC.

Particulars:

Surface buckets for sampling: 1, 2, 6, 9, 13, 17, 18, 70, 126, 181 and 209.

15. No Hex file found – a CNV file contained only shallow data logged.

70. Niskin bottle broken. Took surface sample.

92. Nutrient sample #028 was left out for 18 hours.

100. CHLb sample #033 not processed.

126. No hex file found – there is a cnv file for event 127 that should be called 126. It looks ok.

135. Repeated CTD – not turned on first time but bottle sample taken from 1st cast. Rubber ring around vent of Niskin replaced after this cast.

143. Niskin bottle dropped on July 2nd but appears to be ok.

164. Spigot on Niskin really stiff – very hard to open/close.

170. CTD turned on/off 3 times before deployment.

179. Dissolved Oxygen and salinity data bad until about 60m of downcast –judging from net tow problem may be due to large #s of doliolids. Upcast data selected for archiving.

181. Thought at sea to be bad, but looks ok except at surface and upcast data look bad at the surface too.

212. turned CTD on/off/on before cast.

209. No file – CTD must not have been turned on.

243. Noted after deployment that Y-shaped clear tubing on CTD unit was loose at one end. This had been zap-strapped to keep it in position. This was put back on and another zap-strap added to secure the upper portion of the tubing in hopes that the bottom part of the tube will hold in place.

249. Oxygen levels very high at 10m below surface but not during the soak at 15m. No sign of such a problem in Concerto CTD. DO data looks poor throughout cast.

265. No hex file found. A cnv file  shows only surface data was logged.

CRUISE SUMMARY

| Cruise ID#: 2019-124 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 17 June 2019 End: 15 July 2019 |  |  |  |  |  |  |
| Location: IPES Synoptic Survey |  |  |  |  |  |  |
| Party Chief: Young K. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 456 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0456	Cruise ID#:		2019-124

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5130 | 22Feb2017 | Factory |  |  |
| Conductivity | 3500 | 23Feb2017 | Factory |  |  |
| ECO Fluorometer | 4185 | 07Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1592 | 2Mar2017 | Factory |  |  |
| pH | 691 | 13 Dec2018 | Factory |  |  |
| Pressure | 0668 | 28Feb2017 | Factory |  |  |