## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 9 January 2019 – 12 December 2019

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 6 May 2020 – 25 May 2020

Number of original Hex files: 99 	Number of casts processed: 73

INSTRUMENT SUMMARY

Two Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1449 and 1512. Each CTD had an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

These data form part of a long-standing time-series rather than a cruise. They cover sampling by the Canadian Forces Maritime Experimental and Test Ranges during 2019 in Nanoose Bay, BC.

Calibration information for the 2 CTDs could not be confirmed. The calibrations for #1512 had been updated in 2017 but CTD #1449 has not been recalibrated since 2014.

To assist in analysis, event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project based on the date. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than 1 cast on 1 day, as frequently occurred.

The soak period as seen in the files varied greatly from 15 seconds to 3 minutes, with most lasting between 30s and 75s. It is unknown what pump delay was set or if data logging was delayed. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality. Conductivity appears to have equilibrated before most downcasts began.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, but it is believed that local time was used. All times were converted to UTC by adding 7 or 8 hours depending on whether local time was PDT or PST.

A few casts run close in time were compared in T-S space. There was good repeatability near the bottom where vertical gradients were low.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

99 hex and 100 con files were delivered. There was no hex file for con file 19032601.

The file names are non-standard and contain the letter N followed by the date in YYMMDD format and 2 digits to indicate the sequential casts on a single day. The non-standard names are in the usual format for this project and will be not be changed.

- Conversion of Raw Data

- A test conversion was done to see if the 2 voltage channels contained useful data. Voltage 0 corresponds to the altimeter channel and Voltage 1 contained only noise.

- The raw data were converted by matching the configuration file to the input file name.

- File N19032601.con had no corresponding hex file. A converted file provided showed that the same CTD was used as for file N19032602.hex. The converted file has some channels not required but is missing pressure and altimetry and has temperature in Fahrenheit degrees.

- The cast will not be processed given uncertainties in how it was derived and that there is another cast available from the same day.

- Conversion produced 99 CNV files.

- A spreadsheet “NanooseCtdStations2019.xls” was provided with positions for the casts but there were some casts for which positions were missing. In most cases those casts were shallow and on the same day as a deeper cast, so they will not be processed further.

- Those files are:

- N19032102

- N19032201

- N19032204

- N19032205

- N19032206

- N19032502

- N19032504

- N19032506

- N19032508

- N19051001

- N19051501

- N19052801

- N19061801

- N19072301

- N19072303

- N19072401

- N19081301

- N19081302

- N19082001

- N19082002

- N19082801 (a few metres deeper than N19082802 but spiky)

- N19091101

- N19091102

- N19091301

- N19112001

- The following files are the only ones for a particular day and have no position information so a position based on an average of other positions was entered (124º W 49.19.799ºN):

- N19050901

- N19102401

- The spreadsheet provided was saved as NanooseCtdStations-merge-hdr.csv. Some editing was required:

- The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this.

- A column LOC:Station was added with every row having “Nanoose” entered.

- A column of event numbers was added to the CSV file using MMDD part of the file name as event number.

- Data were acquired at the surface for between 15s to 3 minutes, with most between 30s and 75s.  The shortest soaks were for casts that will not be processed further anyway due to lack of positions. It is unknown what pump delay was set or if data logging was delayed. A 2-minute soak is generally advised for the best salinity data. For most casts the conductivity seems to have reached equilibrium before the downcast began.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom but the CTD did not always get within 15m, in which case the bottom value will not get entered into the IOS Headers.

- WILDEDIT

While no spikes were noted in the data, not all casts were examined so WILDEDIT was run on all casts to remove spikes from the pressure, depth, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 10

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

- FILTER

As usual for SBE 19 CTDs the conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.  Tests were run on a few casts to ensure this was the best choice.

The pressure was not filtered since it is quite smooth already. In previous years tests showed that filtering made things worse in sections where pressure was constant.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Temperature was advanced by 0.73s relative to pressure since that has worked well for data from this project in previous years. Two casts were checked and the results were excellent. Alignment can be fine-tuned later using SHIFT if further adjustment is found useful. Tests were also run using slightly different advancements but 0.73s looked best.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

The program MERGE CSV FILES TO HEADERS was then used to add positions, station name, event numbers and project name to the headers.

8. Checking Headers

Track Plot was used to ensure positions look normal. There are no obvious outliers though, as usual, there is scatter in locations.

Header Check was run. There are some negative conductivity values which are likely to disappear after DELETE is run.

A cross-reference list was produced. The times appear to be in local time which is the usual case for this project. The times of casts are not consistent enough to notice if clocks might have changed from standard to daylight, but there is no obvious trend towards later times in summer, so it is likely that PST and PDT were used.

ADD TIME CHANNEL was used to add 8 hours to casts up to March 9th and from November 3rd to the end of the year.

ADD TIME CHANNEL was used to add 7 hour to all other casts.

Surface check was run and values ranged from -0.5 to +3.2db with the exception of 3 casts with pad values at the surface which were associated with very low salinity.

The average surface pressure for casts that started in the top 4db was 1.14db.

Profile plots were examined. Many casts will need some editing.

Cast N19082701 had bad salinity throughout and there is another cast on the same day, so this file will not be processed further.

The need to edit will be re-examined after DELETE.

- The altimetry header information from each file was exported to a spreadsheet. 69 casts got within 15m of the bottom.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features. Three casts with features that might benefit were chosen to test. As found in the past, the results were not consistent. SHIFT was not run.

10. CLIP

CLIP was used to remove the first 50 scans for most files, but fewer for 3 casts that started down before the 50 record point.

11. DELETE

DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

All warnings pertain to pressure jumps seen in the original files. They are all near the top and bottom of casts.

12. CTDEDIT

The DEL files were copied to EDT.

CTDEDIT was used to edit 66 casts.

Salinity was smoothed in cases where unstable features appeared to be due to misalignments of T and C in the presence of temperature gradients. A few records were removed from the top and bottom of some casts.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen.

T-S plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – Profile plots of temperature and salinity data were examined with the local seasonal climatology ranges superimposed. One cast fell way outside the climatology (N19031905). When examined with other casts on the same day the data looked fine – this led to a discovery of an error in the position in the header. There were many casts that day, so the position was assumed to be an average of positions before and after this cast. After that correction all data fell within the climatology.

All casts fell within the climatology ranges with a few near-surface exceptions:

- The temperature was slightly below the maximum in the seasonal climatology in the top 25m in March and December and a little high in May and June.

- Salinity was a bit high for 2 casts on August 29th, but the values were within the climatology for September to November.

These excursions are not indicative of calibration drift; the climatology is old and is not very suitable for near-shore observations.

Repeat casts – Casts were all close in space, and on March 19th there were 6 casts on one day. On a T-S plot they look close in areas of low temperature gradient. Near the bottom they are within 0.01C° and 0.002psu along lines of constant σt. This is good correspondence.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

-----------------------------------------------------------------------------------------

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

For processing details see 2019-199_Processing_Report.doc.

-----------------------------------------------------------------------------------------

The standards check routine was run and no problems were found.

T-S plots were examined and a problem was identified in cast N19050901. The salinity values look out of line between 7.2 and 12.3db. The temperature gradient is fairly high but the change is gradual, while the salinity increases markedly and then decreases. This looks like a mismatch in alignment. It is too large a feature to justify cleaning by interpolation so the salinity values were replaced with pad values.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing and header check were produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	 4 January 2019	    End:	6 December 2019

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

CTDs

| CTD# | Make | Model | Serial# | Altimeter | Calibration data – Pressure, Temperature, Conductivity |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 19 | 1449 | 1094 | September / October 2014 |
| 2 | SEABIRD | 19 | 1512 | 945 | August 2017 |

The calibration information used in conversion can be found in the headers of the CTD.