#### 2020-008

CCGS John P Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### August 11 to September 8, 2020

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from stations  P2, P4, P12, P16, P20 and P26 for DMS.  Samples were collected from P2, P4, P12, P16, P20 and P26 for DMSPD (dissolved) & DMSPT (total).

  - DMS

Fourteen water samples from various depths (400m, 200m, 150m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at station P4, P12, P16, P20 and P26.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface). Duplicates were taken at 20m.  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in a fridge, in the dark and removed one at a time before analysis.

  - DMSP

Seven samples for both DMSPD and DMSPT were collected ( 400m, 200m, 150m, 30m, 20m, 5m, surface).  The only exception to this was P2 where there were no 400m, 200m or 150m samples, hence, only 4 samples were collected. Duplicates were taken at 20m.

- Analysis

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Helium for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasil 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a 15mL or 5mL polypropylene tube.  The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

Exactly 3.5mL of seawater was collected directly from the niskin into a 15mL or 5mL polypropylene tube. The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.  Stripping efficiency was evaluated at the beginning of the cruise and was proven to be acceptable at over 96%.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by simply treating MQ water as an actual sample.  For example, in the case of DMSPD,   3.5mL was collected from a separate funnel and for DMSPT    3.5mL  was added directly to the polypropylene tube.  Like the samples, they were then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The blank was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Data & Results

  - DMS

The dimethylsulfide system was upgraded somewhat on this cruise and it performed very well.  In February the older 5890 gas chromatograph failed and could not be repaired.  A newer 6890 gas chromatogram was subsequently configured to work with the existing components of the DMS instrument. Using newer technology resulted in better precision to the extent that the % difference in duplicates was the lowest seen in years.  The overall sensitivity of the instrument also was noticeable as was the chromatography.  In short, essentially across all parameters, the analyses improved and the resulting data for this cruise is probably better than it has ever been.

  - DMSP

The DMSP samples were collected for all stations.  Unfortunately the compressor in the fridge where they are stored, while waiting to be analysed, failed and the samples were allowed to be at room temperature for 1-3 weeks.  The integrity of the samples was subsequently compromised and all DMSP data for this cruise was lost.

- Conclusions

6.1 DMS

The failure of the gas chromatograph (GC) in February was, in some ways, a blessing because it forced some much needed updates to the overall system.  Simply using a newer GC on this cruise proved to be a great benefit to the analyses with respect to the quality of data produced.  In fact, data from this cruise is likely better than it has been in years due to the newer technology employed.  Additional improvements are now scheduled to be implemented over the next year in an effort to bring the entire DMS system up to date with current technology and innovations.

  - DMSP

The failure of the fridge back in the laboratory was unfortunate because, as of late, the DMSP analyses was working quite well with many of the problems that plagued the analyses for years being resolved.  Additional samples were actually taken on this cruise to further fine tune the analyses but they too were lost.  A new fridge was purchased and an investigation into audible alarms are being looked at so as to possibly avoid a problem like this in the future.