## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2020/12/14 17:27:57 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 0.5
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2020-008-0015.mrg | 2020-008-0015.sam | 292 | 1250.3 | 1250.36 | 0.20143 | 0.06189 | 34.4664 | 34.4614 | 0.00024659 | -0.0049591 | 34.4664 | 34.4689 | 0.0002361 | 0.0025063
2020-008-0015.mrg | 2020-008-0015.sam | 293 | 1250.38 | 1250.37 | 0.20305 | -0.013428 | 34.4655 | 34.4614 | 0.00024657 | -0.0040588 | 34.4655 | 34.4689 | 0.00023576 | 0.0034065
2020-008-0015.mrg | 2020-008-0015.sam | 241 | 999.627 | 999.627 | 0.14547 | -0.00018311 | 34.3742 | 34.3718 | 0.00022868 | -0.0024071 | 34.3742 | 34.3791 | 0.0002685 | 0.0048866
2020-008-0015.mrg | 2020-008-0015.sam | 241 | 799.713 | 799.713 | 0.13125 | 6.1035e-05 | 34.2947 | 34.2916 | 0.00058958 | -0.0030937 | 34.2947 | 34.2978 | 0.00039483 | 0.0031319
2020-008-0015.mrg | 2020-008-0015.sam | 241 | 5.005 | 5.0049 | 0.046636 | -0.00011635 | 31.897 | 31.89 | 0.00048941 | -0.0070305 | 31.897 | 31.8963 | 0.00046049 | -0.00065994
2020-008-0025.mrg | 2020-008-0025.sam | 234 | 2000.18 | 2000.19 | 0.28093 | 0.012451 | 34.5872 | 34.5825 | 0.00028374 | -0.0047302 | 34.5872 | 34.5912 | 0.00023589 | 0.0040474
2020-008-0025.mrg | 2020-008-0025.sam | 241 | 1499.96 | 1499.96 | 0.23139 | -0.00012207 | 34.5117 | 34.5056 | 0.00022655 | -0.006073 | 34.5117 | 34.5139 | 0.00023813 | 0.0022011
2020-008-0025.mrg | 2020-008-0025.sam | 241 | 1249.68 | 1249.68 | 0.16668 | -0.00024414 | 34.4653 | 34.4566 | 0.00037716 | -0.008728 | 34.4653 | 34.4648 | 0.00038171 | -0.00054932
2020-008-0025.mrg | 2020-008-0025.sam | 241 | 1000.24 | 1000.24 | 0.28414 | 0.00042725 | 34.3798 | 34.3758 | 0.00030484 | -0.0040321 | 34.3798 | 34.3839 | 0.00030497 | 0.0041504
2020-008-0025.mrg | 2020-008-0025.sam | 241 | 799.679 | 799.679 | 0.12854 | -6.1035e-05 | 34.2971 | 34.2903 | 0.00037382 | -0.0068474 | 34.2971 | 34.2981 | 0.0003125 | 0.0010414
2020-008-0025.mrg | 2020-008-0025.sam | 232 | 599.875 | 599.854 | 0.14917 | -0.020752 | 34.1141 | 34.1089 | 0.00057911 | -0.0052338 | 34.1141 | 34.1164 | 0.00045317 | 0.0022736
2020-008-0026.mrg | 2020-008-0026.sam | 241 | 5.468 | 5.4685 | 0.10313 | 0.00046492 | 32.2031 | 32.1889 | 0.00073642 | -0.01424 | 32.2031 | 32.1963 | 0.00091531 | -0.0068398
2020-008-0036.mrg | 2020-008-0036.sam | 194 | 3000.2 | 3000.29 | 0.28793 | 0.093994 | 34.6476 | 34.6407 | 0.00024948 | -0.0068817 | 34.6476 | 34.6505 | 0.00026312 | 0.0028992
2020-008-0036.mrg | 2020-008-0036.sam | 178 | 3000.14 | 3000.21 | 0.28653 | 0.072754 | 34.6468 | 34.6407 | 0.00024784 | -0.0060844 | 34.6468 | 34.6505 | 0.00027381 | 0.0036964
2020-008-0036.mrg | 2020-008-0036.sam | 135 | 2500.07 | 2500.13 | 0.27694 | 0.058105 | 34.6274 | 34.6194 | 0.00030092 | -0.0080261 | 34.6274 | 34.6293 | 0.00023604 | 0.0018616
2020-008-0036.mrg | 2020-008-0036.sam | 181 | 2000.4 | 2000.45 | 0.25964 | 0.04541 | 34.5839 | 34.5771 | 0.00029937 | -0.0068245 | 34.5839 | 34.5866 | 0.00025767 | 0.0027008
2020-008-0036.mrg | 2020-008-0036.sam | 81 | 1499.78 | 1499.81 | 0.30369 | 0.035156 | 34.512 | 34.5045 | 0.00030212 | -0.0075417 | 34.512 | 34.5137 | 0.00022772 | 0.0017471
2020-008-0036.mrg | 2020-008-0036.sam | 166 | 1249.7 | 1249.77 | 0.23107 | 0.071899 | 34.4446 | 34.4387 | 0.00059558 | -0.0059051 | 34.4446 | 34.4476 | 0.00057954 | 0.0030098
2020-008-0036.mrg | 2020-008-0036.sam | 241 | 999.52 | 999.52 | 0.25344 | 0.00018311 | 34.3554 | 34.3491 | 0.00050909 | -0.0062637 | 34.3554 | 34.3584 | 0.0004201 | 0.0030479
2020-008-0037.mrg | 2020-008-0037.sam | 182 | 5.41 | 5.3043 | 0.22746 | -0.10574 | 32.2445 | 32.2344 | 0.00033862 | -0.01012 | 32.2445 | 32.2417 | 0.00050576 | -0.0027809
2020-008-0039.mrg | 2020-008-0039.sam | 136 | 2999.89 | 2999.96 | 0.18945 | 0.068603 | 34.6518 | 34.6432 | 0.0002571 | -0.0086136 | 34.6518 | 34.6535 | 0.00023718 | 0.0017281
2020-008-0039.mrg | 2020-008-0039.sam | 94 | 2499.83 | 2499.82 | 0.28475 | -0.0056152 | 34.6314 | 34.6212 | 0.00032784 | -0.010239 | 34.6314 | 34.6314 | 0.00028192 | -4.9591e-05
2020-008-0039.mrg | 2020-008-0039.sam | 116 | 2250.09 | 2250.09 | 0.22928 | -0.00097656 | 34.6102 | 34.6021 | 0.00029237 | -0.0081139 | 34.6102 | 34.6122 | 0.00027757 | 0.0019607
2020-008-0039.mrg | 2020-008-0039.sam | 131 | 2001.1 | 2000.94 | 0.25592 | -0.15991 | 34.5853 | 34.5771 | 0.00030034 | -0.0081634 | 34.5853 | 34.587 | 0.00028738 | 0.0017166
2020-008-0039.mrg | 2020-008-0039.sam | 126 | 1750.65 | 1750.52 | 0.22683 | -0.13123 | 34.55 | 34.5426 | 0.00030858 | -0.0073509 | 34.55 | 34.5525 | 0.00025846 | 0.0024796
2020-008-0039.mrg | 2020-008-0039.sam | 88 | 1500.16 | 1500.07 | 0.30196 | -0.085205 | 34.5495 | 34.4956 | 0.00026341 | -0.053925 | 34.5495 | 34.5052 | 0.00026398 | -0.044304
2020-008-0039.mrg | 2020-008-0039.sam | 125 | 1249.12 | 1249.06 | 0.25353 | -0.061768 | 34.4399 | 34.4334 | 0.00032533 | -0.006485 | 34.4399 | 34.4428 | 0.00022036 | 0.0029259
2020-008-0039.mrg | 2020-008-0039.sam | 217 | 1000.69 | 1000.6 | 0.19772 | -0.086853 | 34.3637 | 34.3565 | 0.00029161 | -0.0071602 | 34.3637 | 34.3661 | 0.00025565 | 0.0024185
2020-008-0042.mrg | 2020-008-0042.sam | 239 | 5.349 | 5.3532 | 0.25698 | 0.0042052 | 32.1761 | 32.1664 | 0.0011043 | -0.009716 | 32.1761 | 32.1732 | 0.0015561 | -0.0028954
2020-008-0044.mrg | 2020-008-0044.sam | 314 | 3499.23 | 3499.35 | 0.16647 | 0.11475 | 34.6656 | 34.6574 | 0.00024075 | -0.0082436 | 34.6656 | 34.6676 | 0.00022761 | 0.0020256
2020-008-0044.mrg | 2020-008-0044.sam | 314 | 3499.38 | 3499.35 | 0.16647 | -0.0354 | 34.6674 | 34.6574 | 0.00024075 | -0.010044 | 34.6674 | 34.6676 | 0.00022761 | 0.00022507
2020-008-0044.mrg | 2020-008-0044.sam | 184 | 2999.81 | 2999.91 | 0.25488 | 0.10352 | 34.6556 | 34.6452 | 0.00026657 | -0.010407 | 34.6556 | 34.6555 | 0.00024229 | -0.00014114
2020-008-0044.mrg | 2020-008-0044.sam | 228 | 2499.75 | 2499.75 | 0.2953 | -0.0058594 | 34.6302 | 34.6198 | 0.00026929 | -0.010387 | 34.6302 | 34.6298 | 0.00024154 | -0.00040054
2020-008-0044.mrg | 2020-008-0044.sam | 231 | 1999.9 | 1999.93 | 0.17108 | 0.025024 | 34.5844 | 34.5733 | 0.00027212 | -0.011082 | 34.5844 | 34.583 | 0.00024225 | -0.001358
2020-008-0044.mrg | 2020-008-0044.sam | 240 | 1500.56 | 1500.56 | 0.22405 | -0.0021973 | 34.4995 | 34.4902 | 0.00027416 | -0.0093498 | 34.4995 | 34.4998 | 0.00023291 | 0.00027847
2020-008-0044.mrg | 2020-008-0044.sam | 202 | 1250.54 | 1250.65 | 0.26647 | 0.1134 | 34.4259 | 34.4165 | 0.00026316 | -0.0094109 | 34.4259 | 34.4257 | 0.00029704 | -0.00022507
2020-008-0044.mrg | 2020-008-0044.sam | 241 | 1000.6 | 1000.6 | 0.21307 | -0.00036621 | 34.3567 | 34.343 | 0.00027057 | -0.01371 | 34.3567 | 34.3524 | 0.00045785 | -0.0043373
2020-008-0050.mrg | 2020-008-0050.sam | 241 | 4.77 | 4.7699 | 0.16667 | -7.4863e-05 | 32.2318 | 32.2212 | 0.00031361 | -0.010624 | 32.2318 | 32.2292 | 0.00044036 | -0.0025902
2020-008-0054.mrg | 2020-008-0054.sam | 230 | 5.373 | 5.4024 | 0.21234 | 0.029357 | 32.1553 | 32.1451 | 0.00024311 | -0.010159 | 32.1553 | 32.1532 | 0.00036699 | -0.0021172
2020-008-0060.mrg | 2020-008-0060.sam | 119 | 3499.73 | 3499.66 | 0.2411 | -0.070068 | 34.6686 | 34.6609 | 0.00025901 | -0.0076561 | 34.6686 | 34.672 | 0.00024795 | 0.0034218
2020-008-0060.mrg | 2020-008-0060.sam | 106 | 3499.85 | 3499.75 | 0.2304 | -0.10278 | 34.6698 | 34.6609 | 0.00025512 | -0.0088501 | 34.6698 | 34.672 | 0.00024423 | 0.0022202
2020-008-0060.mrg | 2020-008-0060.sam | 191 | 2999.93 | 2999.74 | 0.23198 | -0.19019 | 34.6535 | 34.6464 | 0.00029034 | -0.0071182 | 34.6535 | 34.6573 | 0.00023957 | 0.0037613
2020-008-0060.mrg | 2020-008-0060.sam | 156 | 2500 | 2499.93 | 0.30337 | -0.064941 | 34.6283 | 34.6186 | 0.00027798 | -0.0097084 | 34.6283 | 34.6293 | 0.00026747 | 0.00095367
2020-008-0060.mrg | 2020-008-0060.sam | 140 | 2000.07 | 2000.03 | 0.28541 | -0.039795 | 34.5847 | 34.5758 | 0.0002885 | -0.0089226 | 34.5847 | 34.5863 | 0.00024972 | 0.0015488
2020-008-0060.mrg | 2020-008-0060.sam | 72 | 1500.11 | 1500.14 | 0.29747 | 0.023315 | 34.4951 | 34.4868 | 0.00027064 | -0.008255 | 34.4951 | 34.4971 | 0.000246 | 0.0020065
2020-008-0060.mrg | 2020-008-0060.sam | 194 | 1249.94 | 1250.11 | 0.2549 | 0.16711 | 34.4388 | 34.4295 | 0.00026498 | -0.0093193 | 34.4388 | 34.4398 | 0.00024341 | 0.00095749
2020-008-0060.mrg | 2020-008-0060.sam | 241 | 1000.36 | 1000.36 | 0.18849 | -0.00030518 | 34.3681 | 34.359 | 0.00030194 | -0.0090561 | 34.3681 | 34.3694 | 0.0002431 | 0.0013046
2020-008-0063.mrg | 2020-008-0063.sam | 241 | 5.18 | 5.1801 | 0.11608 | 8.297e-05 | 32.2235 | 32.2119 | 0.001954 | -0.01162 | 32.2235 | 32.2206 | 0.0016194 | -0.0029449
2020-008-0064.mrg | 2020-008-0064.sam | 78 | 3498.5 | 3498.41 | 0.295 | -0.08667 | 34.6711 | 34.6619 | 0.00022559 | -0.0091705 | 34.6711 | 34.6731 | 0.00026936 | 0.0020218
2020-008-0064.mrg | 2020-008-0064.sam | 41 | 3001.75 | 3001.8 | 0.28388 | 0.052002 | 34.6555 | 34.6464 | 0.00025221 | -0.0091438 | 34.6555 | 34.6575 | 0.00018429 | 0.0019684
2020-008-0064.mrg | 2020-008-0064.sam | 129 | 2500.47 | 2500.54 | 0.3076 | 0.077148 | 34.6309 | 34.6216 | 0.00028193 | -0.0093193 | 34.6309 | 34.6325 | 0.00025497 | 0.001648
2020-008-0064.mrg | 2020-008-0064.sam | 73 | 2001.55 | 2001.61 | 0.29108 | 0.054443 | 34.5861 | 34.5773 | 0.00028049 | -0.0088234 | 34.5861 | 34.5881 | 0.00023571 | 0.0019722
2020-008-0064.mrg | 2020-008-0064.sam | 145 | 1751.16 | 1751.17 | 0.2579 | 0.0032959 | 34.5504 | 34.5414 | 0.0003278 | -0.0089836 | 34.5504 | 34.5521 | 0.00025656 | 0.0016861
2020-008-0064.mrg | 2020-008-0064.sam | 49 | 1498.74 | 1498.67 | 0.25257 | -0.07544 | 34.4984 | 34.4892 | 0.00038622 | -0.0092277 | 34.4984 | 34.4998 | 0.00026206 | 0.0013885
2020-008-0064.mrg | 2020-008-0064.sam | 113 | 1251.28 | 1251.21 | 0.3036 | -0.076172 | 34.442 | 34.4306 | 0.00029461 | -0.011402 | 34.442 | 34.4411 | 0.00027371 | -0.0009346
2020-008-0064.mrg | 2020-008-0064.sam | 35 | 1000.85 | 1000.85 | 0.29238 | -0.00018311 | 34.3659 | 34.3577 | 0.00029694 | -0.0082283 | 34.3659 | 34.3681 | 0.00033897 | 0.0022469
2020-008-0066.mrg | 2020-008-0066.sam | 222 | 5.384 | 5.3429 | 0.29944 | -0.041081 | 32.2389 | 32.2259 | 0.00022853 | -0.01302 | 32.2389 | 32.2355 | 0.00028454 | -0.0034065
2020-008-0068.mrg | 2020-008-0068.sam | 85 | 3500.5 | 3500.4 | 0.27518 | -0.10205 | 34.6705 | 34.6614 | 0.00026366 | -0.0091515 | 34.6705 | 34.6733 | 0.00026642 | 0.0027695
2020-008-0068.mrg | 2020-008-0068.sam | 123 | 3000.88 | 3000.83 | 0.30225 | -0.052002 | 34.6546 | 34.646 | 0.0002615 | -0.0085907 | 34.6546 | 34.6578 | 0.00023569 | 0.0032082
2020-008-0068.mrg | 2020-008-0068.sam | 115 | 2499.75 | 2499.67 | 0.30619 | -0.084961 | 34.6289 | 34.6195 | 0.00028989 | -0.0094185 | 34.6289 | 34.631 | 0.00024355 | 0.0020676
2020-008-0068.mrg | 2020-008-0068.sam | 132 | 1999.93 | 1999.87 | 0.28781 | -0.060913 | 34.5806 | 34.5721 | 0.00028669 | -0.0084801 | 34.5806 | 34.5835 | 0.00027324 | 0.0028954
2020-008-0068.mrg | 2020-008-0068.sam | 142 | 1749.93 | 1749.93 | 0.31646 | 0.00085449 | 34.5417 | 34.5348 | 0.00032985 | -0.0069466 | 34.5417 | 34.5462 | 0.0002634 | 0.0045013
2020-008-0068.mrg | 2020-008-0068.sam | 102 | 1501.17 | 1501.2 | 0.304 | 0.0271 | 34.4925 | 34.4843 | 0.00034906 | -0.0081673 | 34.4925 | 34.4954 | 0.00026521 | 0.0028801
2020-008-0068.mrg | 2020-008-0068.sam | 53 | 1250.17 | 1250.05 | 0.28413 | -0.11707 | 34.4321 | 34.4227 | 0.0002612 | -0.0093498 | 34.4321 | 34.4339 | 0.00022898 | 0.001812
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 999.642 | 999.642 | 0.2142 | 0.00030518 | 34.3528 | 34.3453 | 0.00044464 | -0.0074844 | 34.3528 | 34.3562 | 0.00028994 | 0.0033722
2020-008-0068.mrg | 2020-008-0068.sam | 131 | 800.303 | 800.252 | 0.26968 | -0.051147 | 34.2796 | 34.2692 | 0.00036911 | -0.010372 | 34.2796 | 34.2803 | 0.00036131 | 0.0006752
2020-008-0068.mrg | 2020-008-0068.sam | 139 | 600.15 | 600.106 | 0.25438 | -0.043823 | 34.1633 | 34.1532 | 0.00029369 | -0.010094 | 34.1633 | 34.1639 | 0.00027824 | 0.00061035
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 400.367 | 400.367 | 0.15632 | -0.00045776 | 33.955 | 33.9451 | 0.00034766 | -0.0099449 | 33.955 | 33.9549 | 0.00069006 | -6.8665e-05
2020-008-0068.mrg | 2020-008-0068.sam | 229 | 299.703 | 299.726 | 0.2509 | 0.02301 | 33.8467 | 33.8395 | 0.00033872 | -0.0071907 | 33.8467 | 33.8503 | 0.00048147 | 0.0036316
2020-008-0068.mrg | 2020-008-0068.sam | 119 | 250.212 | 250.304 | 0.30147 | 0.091644 | 33.8002 | 33.7925 | 0.0031109 | -0.00774 | 33.8002 | 33.8017 | 0.002572 | 0.0014648
2020-008-0068.mrg | 2020-008-0068.sam | 178 | 200.158 | 200.078 | 0.28196 | -0.080353 | 33.7612 | 33.7508 | 0.00075999 | -0.010433 | 33.7612 | 33.761 | 0.0010698 | -0.00018692
2020-008-0068.mrg | 2020-008-0068.sam | 165 | 175.062 | 175.068 | 0.29758 | 0.006424 | 33.7571 | 33.7481 | 0.0020474 | -0.0090141 | 33.7571 | 33.7591 | 0.0026202 | 0.002018
2020-008-0068.mrg | 2020-008-0068.sam | 145 | 150.305 | 150.394 | 0.27201 | 0.088913 | 33.6486 | 33.6359 | 0.0015743 | -0.012688 | 33.6486 | 33.6441 | 0.00073447 | -0.0044632
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 124.669 | 124.669 | 0.23233 | 0.00039673 | 33.4055 | 33.3803 | 0.0031342 | -0.025158 | 33.4055 | 33.3875 | 0.0011417 | -0.01796
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 99.947 | 99.9468 | 0.26193 | -0.00020599 | 32.7532 | 32.7148 | 0.010908 | -0.038395 | 32.7532 | 32.7263 | 0.0052047 | -0.026863
2020-008-0068.mrg | 2020-008-0068.sam | 131 | 74.894 | 74.9348 | 0.23928 | 0.040756 | 32.4975 | 32.4834 | 0.0033667 | -0.014065 | 32.4975 | 32.4939 | 0.0040551 | -0.0035744
2020-008-0068.mrg | 2020-008-0068.sam | 240 | 49.623 | 49.6209 | 0.28298 | -0.0020943 | 32.4092 | 32.3965 | 0.0014084 | -0.012699 | 32.4092 | 32.4068 | 0.0019605 | -0.0023994
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 25.23 | 25.2305 | 0.24203 | 0.00046539 | 32.2956 | 32.2798 | 0.0021184 | -0.015831 | 32.2956 | 32.2898 | 0.0020023 | -0.0058403
2020-008-0068.mrg | 2020-008-0068.sam | 233 | 10.075 | 10.0577 | 0.23194 | -0.017301 | 32.2928 | 32.2829 | 0.0002693 | -0.0099487 | 32.2928 | 32.2922 | 0.00039167 | -0.00061035
2020-008-0068.mrg | 2020-008-0068.sam | 241 | 5.306 | 5.306 | 0.21394 | 4.53e-05 | 32.2927 | 32.2825 | 0.00025707 | -0.010197 | 32.2927 | 32.292 | 0.00029408 | -0.00072861
2020-008-0069.mrg | 2020-008-0069.sam | 241 | 5.079 | 5.0787 | 0.18878 | -0.00030708 | 32.2658 | 32.2532 | 0.0016025 | -0.012608 | 32.2658 | 32.2627 | 0.0038586 | -0.0031319
2020-008-0078.mrg | 2020-008-0078.sam | 114 | 3999.77 | 3999.88 | 0.20639 | 0.11255 | 34.6805 | 34.6691 | 0.00024898 | -0.011379 | 34.6805 | 34.682 | 0.0003076 | 0.0014687
2020-008-0078.mrg | 2020-008-0078.sam | 116 | 3499.37 | 3499.37 | 0.25852 | -0.0065918 | 34.6712 | 34.659 | 0.00023864 | -0.012188 | 34.6712 | 34.672 | 0.00025167 | 0.0007515
2020-008-0078.mrg | 2020-008-0078.sam | 196 | 3000.15 | 3000.2 | 0.27773 | 0.052246 | 34.6532 | 34.6414 | 0.00029117 | -0.011768 | 34.6532 | 34.6542 | 0.00026121 | 0.0010185
2020-008-0078.mrg | 2020-008-0078.sam | 116 | 2500.66 | 2500.63 | 0.28817 | -0.027832 | 34.6223 | 34.6128 | 0.00033143 | -0.0094795 | 34.6223 | 34.6254 | 0.0002522 | 0.003128
2020-008-0078.mrg | 2020-008-0078.sam | 116 | 1999.77 | 1999.7 | 0.28354 | -0.070801 | 34.5794 | 34.5672 | 0.00031966 | -0.012165 | 34.5794 | 34.5796 | 0.00027556 | 0.00019455
2020-008-0078.mrg | 2020-008-0078.sam | 115 | 1499.61 | 1499.49 | 0.29728 | -0.12732 | 34.4957 | 34.4826 | 0.00027113 | -0.013138 | 34.4957 | 34.4947 | 0.00021468 | -0.0010452
2020-008-0078.mrg | 2020-008-0078.sam | 194 | 1250.69 | 1250.79 | 0.25561 | 0.099243 | 34.4385 | 34.426 | 0.00029816 | -0.012497 | 34.4385 | 34.438 | 0.00029052 | -0.00045395
2020-008-0078.mrg | 2020-008-0078.sam | 175 | 1000.2 | 1000.28 | 0.24986 | 0.082397 | 34.3619 | 34.351 | 0.00032816 | -0.01086 | 34.3619 | 34.3631 | 0.00028199 | 0.001236
2020-008-0086.mrg | 2020-008-0086.sam | 283 | 1250.17 | 1250.14 | 0.23596 | -0.02478 | 34.4536 | 34.4379 | 0.00025296 | -0.015678 | 34.4536 | 34.4538 | 0.00023494 | 0.00020599
2020-008-0086.mrg | 2020-008-0086.sam | 283 | 1250.14 | 1250.14 | 0.23596 | 0.0021973 | 34.4526 | 34.4379 | 0.00025296 | -0.014675 | 34.4526 | 34.4538 | 0.00023494 | 0.0012093
2020-008-0086.mrg | 2020-008-0086.sam | 170 | 1000.09 | 1000.07 | 0.31961 | -0.020203 | 34.379 | 34.3651 | 0.00039911 | -0.013882 | 34.379 | 34.3808 | 0.000379 | 0.0018311
2020-008-0086.mrg | 2020-008-0086.sam | 241 | 799.658 | 799.658 | 0.1291 | 0.00024414 | 34.2789 | 34.2654 | 0.00037259 | -0.013515 | 34.2789 | 34.2804 | 0.00030276 | 0.0015221
2020-008-0086.mrg | 2020-008-0086.sam | 241 | 5.071 | 5.0711 | 0.21181 | 0.00012875 | 31.938 | 31.9219 | 0.00021114 | -0.016075 | 31.938 | 31.9356 | 0.0002124 | -0.0024052
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 18
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2020-008-0015.mrg | 2020-008-0015.sam | 12 | 16 | 2020-008-0015.mrg & 2020-008-0015.sam
2 | 2020-008-0025.mrg | 2020-008-0025.sam | 18 | 23 | 2020-008-0025.mrg & 2020-008-0025.sam
3 | 2020-008-0026.mrg | 2020-008-0026.sam | 25 | 25 | 2020-008-0026.mrg & 2020-008-0026.sam
4 | 2020-008-0036.mrg | 2020-008-0036.sam | 27 | 33 | 2020-008-0036.mrg & 2020-008-0036.sam
5 | 2020-008-0037.mrg | 2020-008-0037.sam | 35 | 35 | 2020-008-0037.mrg & 2020-008-0037.sam
6 | 2020-008-0039.mrg | 2020-008-0039.sam | 37 | 44 | 2020-008-0039.mrg & 2020-008-0039.sam
7 | 2020-008-0042.mrg | 2020-008-0042.sam | 46 | 46 | 2020-008-0042.mrg & 2020-008-0042.sam
8 | 2020-008-0044.mrg | 2020-008-0044.sam | 48 | 55 | 2020-008-0044.mrg & 2020-008-0044.sam
9 | 2020-008-0050.mrg | 2020-008-0050.sam | 57 | 57 | 2020-008-0050.mrg & 2020-008-0050.sam
10 | 2020-008-0054.mrg | 2020-008-0054.sam | 59 | 59 | 2020-008-0054.mrg & 2020-008-0054.sam
11 | 2020-008-0060.mrg | 2020-008-0060.sam | 61 | 68 | 2020-008-0060.mrg & 2020-008-0060.sam
12 | 2020-008-0063.mrg | 2020-008-0063.sam | 70 | 70 | 2020-008-0063.mrg & 2020-008-0063.sam
13 | 2020-008-0064.mrg | 2020-008-0064.sam | 72 | 79 | 2020-008-0064.mrg & 2020-008-0064.sam
14 | 2020-008-0066.mrg | 2020-008-0066.sam | 81 | 81 | 2020-008-0066.mrg & 2020-008-0066.sam
15 | 2020-008-0068.mrg | 2020-008-0068.sam | 83 | 105 | 2020-008-0068.mrg & 2020-008-0068.sam
16 | 2020-008-0069.mrg | 2020-008-0069.sam | 107 | 107 | 2020-008-0069.mrg & 2020-008-0069.sam
17 | 2020-008-0078.mrg | 2020-008-0078.sam | 109 | 116 | 2020-008-0078.mrg & 2020-008-0078.sam
18 | 2020-008-0086.mrg | 2020-008-0086.sam | 118 | 122 | 2020-008-0086.mrg & 2020-008-0086.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 18
2 | (Sal1 - Sal_Bot) | Pressure | 18
3 | (Sal0 - Sal_Bot) | File Pair # | 18
4 | (Sal1 - Sal_Bot) | File Pair # | 18

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.017394 |  | X | CALC Y
m= | -3.7694e-07 | 3.4612e-07 |  | Std Dev of Resid |  | 0.0026009 |  | 599.854 | -0.0082739
b= | -0.0080478 | 0.00073429 |  | Resid Sum of Squares |  | 0.00045323 |  | 3999.88 | -0.0095555
# of Values Extracted = |  |  | 94
# of Values Used in Fit = |  |  | 69
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1250.36 | 34.4614 | 0.00024659 | -0.0049591 | -0.0085191 | -0.00356 |  | 1250.36 | -0.0049591
1 | YES | 1250.37 | 34.4614 | 0.00024657 | -0.0040588 | -0.0085191 | -0.0044602 |  | 1250.37 | -0.0040588
1 | YES | 999.627 | 34.3718 | 0.00022868 | -0.0024071 | -0.0084246 | -0.0060175 |  | 999.627 | -0.0024071
1 | YES | 799.713 | 34.2916 | 0.00058958 | -0.0030937 | -0.0083492 | -0.0052555 |  | 799.713 | -0.0030937
1 | NO | 5.0049 | 31.89 | 0.00048941 | -0.0070305 | -0.0080497 | -0.0010192 |  | 2000.19 | -0.0047302
2 | YES | 2000.19 | 34.5825 | 0.00028374 | -0.0047302 | -0.0088017 | -0.0040715 |  | 1499.96 | -0.006073
2 | YES | 1499.96 | 34.5056 | 0.00022655 | -0.006073 | -0.0086132 | -0.0025402 |  | 1249.68 | -0.008728
2 | YES | 1249.68 | 34.4566 | 0.00037716 | -0.008728 | -0.0085188 | 0.00020921 |  | 1000.24 | -0.0040321
2 | YES | 1000.24 | 34.3758 | 0.00030484 | -0.0040321 | -0.0084248 | -0.0043927 |  | 799.679 | -0.0068474
2 | YES | 799.679 | 34.2903 | 0.00037382 | -0.0068474 | -0.0083492 | -0.0015018 |  | 599.854 | -0.0052338
2 | YES | 599.854 | 34.1089 | 0.00057911 | -0.0052338 | -0.0082739 | -0.0030401 |  | 3000.29 | -0.0068817
3 | NO | 5.4685 | 32.1889 | 0.00073642 | -0.01424 | -0.0080498 | 0.0061905 |  | 3000.21 | -0.0060844
4 | YES | 3000.29 | 34.6407 | 0.00024948 | -0.0068817 | -0.0091787 | -0.002297 |  | 2500.13 | -0.0080261
4 | YES | 3000.21 | 34.6407 | 0.00024784 | -0.0060844 | -0.0091787 | -0.0030942 |  | 2000.45 | -0.0068245
4 | YES | 2500.13 | 34.6194 | 0.00030092 | -0.0080261 | -0.0089902 | -0.00096405 |  | 1499.81 | -0.0075417
4 | YES | 2000.45 | 34.5771 | 0.00029937 | -0.0068245 | -0.0088018 | -0.0019773 |  | 1249.77 | -0.0059051
4 | YES | 1499.81 | 34.5045 | 0.00030212 | -0.0075417 | -0.0086131 | -0.0010714 |  | 999.52 | -0.0062637
4 | YES | 1249.77 | 34.4387 | 0.00059558 | -0.0059051 | -0.0085189 | -0.0026137 |  | 2999.96 | -0.0086136
4 | YES | 999.52 | 34.3491 | 0.00050909 | -0.0062637 | -0.0084245 | -0.0021608 |  | 2499.82 | -0.010239
5 | NO | 5.3043 | 32.2344 | 0.00033862 | -0.01012 | -0.0080498 | 0.0020706 |  | 2250.09 | -0.0081139
6 | YES | 2999.96 | 34.6432 | 0.0002571 | -0.0086136 | -0.0091786 | -0.00056499 |  | 2000.94 | -0.0081634
6 | YES | 2499.82 | 34.6212 | 0.00032784 | -0.010239 | -0.0089901 | 0.0012485 |  | 1750.52 | -0.0073509
6 | YES | 2250.09 | 34.6021 | 0.00029237 | -0.0081139 | -0.0088959 | -0.00078206 |  | 1249.06 | -0.006485
6 | YES | 2000.94 | 34.5771 | 0.00030034 | -0.0081634 | -0.008802 | -0.00063855 |  | 1000.6 | -0.0071602
6 | YES | 1750.52 | 34.5426 | 0.00030858 | -0.0073509 | -0.0087076 | -0.0013567 |  | 3499.35 | -0.0082436
6 | NO | 1500.07 | 34.4956 | 0.00026341 | -0.053925 | -0.0086132 | 0.045311 | 34.5495 | 3499.35 | -0.010044
6 | YES | 1249.06 | 34.4334 | 0.00032533 | -0.006485 | -0.0085186 | -0.0020336 | 34.4399 | 2999.91 | -0.010407
6 | YES | 1000.6 | 34.3565 | 0.00029161 | -0.0071602 | -0.0084249 | -0.0012647 | 34.3637 | 2499.75 | -0.010387
7 | NO | 5.3532 | 32.1664 | 0.0011043 | -0.009716 | -0.0080498 | 0.0016662 | 32.1761 | 1999.93 | -0.011082
8 | YES | 3499.35 | 34.6574 | 0.00024075 | -0.0082436 | -0.0093668 | -0.0011233 | 34.6656 | 1500.56 | -0.0093498
8 | YES | 3499.35 | 34.6574 | 0.00024075 | -0.010044 | -0.0093668 | 0.00067728 | 34.6674 | 1250.65 | -0.0094109
8 | YES | 2999.91 | 34.6452 | 0.00026657 | -0.010407 | -0.0091786 | 0.0012279 | 34.6556 | 3499.66 | -0.0076561
8 | YES | 2499.75 | 34.6198 | 0.00026929 | -0.010387 | -0.00899 | 0.0013974 | 34.6302 | 3499.75 | -0.0088501
8 | YES | 1999.93 | 34.5733 | 0.00027212 | -0.011082 | -0.0088016 | 0.0022801 | 34.5844 | 2999.74 | -0.0071182
8 | YES | 1500.56 | 34.4902 | 0.00027416 | -0.0093498 | -0.0086134 | 0.00073643 | 34.4995 | 2499.93 | -0.0097084
8 | YES | 1250.65 | 34.4165 | 0.00026316 | -0.0094109 | -0.0085192 | 0.00089167 | 34.4259 | 2000.03 | -0.0089226
8 | NO | 1000.6 | 34.343 | 0.00027057 | -0.01371 | -0.0084249 | 0.0052851 | 34.3567 | 1500.14 | -0.008255
9 | NO | 4.7699 | 32.2212 | 0.00031361 | -0.010624 | -0.0080496 | 0.0025743 |  | 1250.11 | -0.0093193
10 | NO | 5.4024 | 32.1451 | 0.00024311 | -0.010159 | -0.0080498 | 0.0021087 |  | 1000.36 | -0.0090561
11 | YES | 3499.66 | 34.6609 | 0.00025901 | -0.0076561 | -0.0093669 | -0.0017108 |  | 3498.41 | -0.0091705
11 | YES | 3499.75 | 34.6609 | 0.00025512 | -0.0088501 | -0.009367 | -0.00051687 |  | 3001.8 | -0.0091438
11 | YES | 2999.74 | 34.6464 | 0.00029034 | -0.0071182 | -0.0091785 | -0.0020603 |  | 2500.54 | -0.0093193
11 | YES | 2499.93 | 34.6186 | 0.00027798 | -0.0097084 | -0.0089901 | 0.00071831 |  | 2001.61 | -0.0088234
11 | YES | 2000.03 | 34.5758 | 0.0002885 | -0.0089226 | -0.0088017 | 0.00012092 |  | 1751.17 | -0.0089836
11 | YES | 1500.14 | 34.4868 | 0.00027064 | -0.008255 | -0.0086132 | -0.00035823 |  | 1498.67 | -0.0092277
11 | YES | 1250.11 | 34.4295 | 0.00026498 | -0.0093193 | -0.008519 | 0.00080033 |  | 1251.21 | -0.011402
11 | YES | 1000.36 | 34.359 | 0.00030194 | -0.0090561 | -0.0084248 | 0.00063125 |  | 1000.85 | -0.0082283
12 | NO | 5.1801 | 32.2119 | 0.001954 | -0.01162 | -0.0080497 | 0.0035699 |  | 3500.4 | -0.0091515
13 | YES | 3498.41 | 34.6619 | 0.00022559 | -0.0091705 | -0.0093665 | -0.00019593 |  | 3000.83 | -0.0085907
13 | YES | 3001.8 | 34.6464 | 0.00025221 | -0.0091438 | -0.0091793 | -3.5439e-05 |  | 2499.67 | -0.0094185
13 | YES | 2500.54 | 34.6216 | 0.00028193 | -0.0093193 | -0.0089903 | 0.00032899 |  | 1999.87 | -0.0084801
13 | YES | 2001.61 | 34.5773 | 0.00028049 | -0.0088234 | -0.0088023 | 2.1135e-05 |  | 1749.93 | -0.0069466
13 | YES | 1751.17 | 34.5414 | 0.0003278 | -0.0089836 | -0.0087079 | 0.00027576 |  | 1501.2 | -0.0081673
13 | YES | 1498.67 | 34.4892 | 0.00038622 | -0.0092277 | -0.0086127 | 0.00061507 |  | 1250.05 | -0.0093498
13 | YES | 1251.21 | 34.4306 | 0.00029461 | -0.011402 | -0.0085194 | 0.0028827 |  | 999.642 | -0.0074844
13 | YES | 1000.85 | 34.3577 | 0.00029694 | -0.0082283 | -0.008425 | -0.00019673 |  | 800.252 | -0.010372
14 | NO | 5.3429 | 32.2259 | 0.00022853 | -0.01302 | -0.0080498 | 0.0049698 |  | 600.106 | -0.010094
15 | YES | 3500.4 | 34.6614 | 0.00026366 | -0.0091515 | -0.0093672 | -0.00021575 |  | 3999.88 | -0.011379
15 | YES | 3000.83 | 34.646 | 0.0002615 | -0.0085907 | -0.0091789 | -0.0005882 |  | 3499.37 | -0.012188
15 | YES | 2499.67 | 34.6195 | 0.00028989 | -0.0094185 | -0.00899 | 0.0004285 |  | 3000.2 | -0.011768
15 | YES | 1999.87 | 34.5721 | 0.00028669 | -0.0084801 | -0.0088016 | -0.00032153 |  | 2500.63 | -0.0094795
15 | YES | 1749.93 | 34.5348 | 0.00032985 | -0.0069466 | -0.0087074 | -0.0017608 |  | 1999.7 | -0.012165
15 | YES | 1501.2 | 34.4843 | 0.00034906 | -0.0081673 | -0.0086136 | -0.00044636 |  | 1499.49 | -0.013138
15 | YES | 1250.05 | 34.4227 | 0.0002612 | -0.0093498 | -0.008519 | 0.00083086 |  | 1250.79 | -0.012497
15 | YES | 999.642 | 34.3453 | 0.00044464 | -0.0074844 | -0.0084246 | -0.00094013 |  | 1000.28 | -0.01086
15 | YES | 800.252 | 34.2692 | 0.00036911 | -0.010372 | -0.0083494 | 0.0020228 |  | 1250.14 | -0.015678
15 | YES | 600.106 | 34.1532 | 0.00029369 | -0.010094 | -0.008274 | 0.0018197 |  | 1250.14 | -0.014675
15 | NO | 400.367 | 33.9451 | 0.00034766 | -0.0099449 | -0.0081987 | 0.0017462 |  | 1000.07 | -0.013882
15 | NO | 299.726 | 33.8395 | 0.00033872 | -0.0071907 | -0.0081607 | -0.00097004 |  | 799.658 | -0.013515 | -0.0087711 | 0.0026045
15 | NO | 250.304 | 33.7925 | 0.0031109 | -0.00774 | -0.0081421 | -0.0004021
15 | NO | 200.078 | 33.7508 | 0.00075999 | -0.010433 | -0.0081232 | 0.00231
15 | NO | 175.068 | 33.7481 | 0.0020474 | -0.0090141 | -0.0081138 | 0.00090037
15 | NO | 150.394 | 33.6359 | 0.0015743 | -0.012688 | -0.0081045 | 0.0045832
15 | NO | 124.669 | 33.3803 | 0.0031342 | -0.025158 | -0.0080948 | 0.017063
15 | NO | 99.9468 | 32.7148 | 0.010908 | -0.038395 | -0.0080854 | 0.030309
15 | NO | 74.9348 | 32.4834 | 0.0033667 | -0.014065 | -0.008076 | 0.0059888
15 | NO | 49.6209 | 32.3965 | 0.0014084 | -0.012699 | -0.0080665 | 0.0046326
15 | NO | 25.2305 | 32.2798 | 0.0021184 | -0.015831 | -0.0080573 | 0.0077737
15 | NO | 10.0577 | 32.2829 | 0.0002693 | -0.0099487 | -0.0080516 | 0.0018972
15 | NO | 5.306 | 32.2825 | 0.00025707 | -0.010197 | -0.0080498 | 0.0021469
16 | NO | 5.0787 | 32.2532 | 0.0016025 | -0.012608 | -0.0080497 | 0.0045579
17 | YES | 3999.88 | 34.6691 | 0.00024898 | -0.011379 | -0.0095555 | 0.0018237
17 | YES | 3499.37 | 34.659 | 0.00023864 | -0.012188 | -0.0093668 | 0.0028212
17 | YES | 3000.2 | 34.6414 | 0.00029117 | -0.011768 | -0.0091787 | 0.0025896
17 | YES | 2500.63 | 34.6128 | 0.00033143 | -0.0094795 | -0.0089904 | 0.00048916
17 | YES | 1999.7 | 34.5672 | 0.00031966 | -0.012165 | -0.0088015 | 0.0033636
17 | YES | 1499.49 | 34.4826 | 0.00027113 | -0.013138 | -0.008613 | 0.0045248
17 | YES | 1250.79 | 34.426 | 0.00029816 | -0.012497 | -0.0085192 | 0.0039777
17 | YES | 1000.28 | 34.351 | 0.00032816 | -0.01086 | -0.0084248 | 0.0024356
18 | YES | 1250.14 | 34.4379 | 0.00025296 | -0.015678 | -0.008519 | 0.0071594
18 | YES | 1250.14 | 34.4379 | 0.00025296 | -0.014675 | -0.008519 | 0.0061561
18 | YES | 1000.07 | 34.3651 | 0.00039911 | -0.013882 | -0.0084247 | 0.005457
18 | YES | 799.658 | 34.2654 | 0.00037259 | -0.013515 | -0.0083492 | 0.0051663
18 | NO | 5.0711 | 31.9219 | 0.00021114 | -0.016075 | -0.0080497 | 0.0080254
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0011351 |  | X | CALC Y
m= | 5.1126e-08 | 1.8529e-07 |  | Std Dev of Resid |  | 0.0013923 |  | 599.854 | 0.0017249
b= | 0.0016942 | 0.00039308 |  | Resid Sum of Squares |  | 0.00012988 |  | 3999.88 | 0.0018987
# of Values Extracted = |  |  | 94
# of Values Used in Fit = |  |  | 69
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1250.36 | 34.4689 | 0.0002361 | 0.0025063 | 0.0017582 | -0.00074809 |  | 1250.36 | 0.0025063
1 | YES | 1250.37 | 34.4689 | 0.00023576 | 0.0034065 | 0.0017582 | -0.0016483 |  | 1250.37 | 0.0034065
1 | YES | 999.627 | 34.3791 | 0.0002685 | 0.0048866 | 0.0017454 | -0.0031413 |  | 999.627 | 0.0048866
1 | YES | 799.713 | 34.2978 | 0.00039483 | 0.0031319 | 0.0017351 | -0.0013967 |  | 799.713 | 0.0031319
1 | NO | 5.0049 | 31.8963 | 0.00046049 | -0.00065994 | 0.0016945 | 0.0023544 |  | 2000.19 | 0.0040474
2 | YES | 2000.19 | 34.5912 | 0.00023589 | 0.0040474 | 0.0017965 | -0.0022509 |  | 1499.96 | 0.0022011
2 | YES | 1499.96 | 34.5139 | 0.00023813 | 0.0022011 | 0.0017709 | -0.00043015 |  | 1249.68 | -0.00054932
2 | YES | 1249.68 | 34.4648 | 0.00038171 | -0.00054932 | 0.0017581 | 0.0023075 |  | 1000.24 | 0.0041504
2 | YES | 1000.24 | 34.3839 | 0.00030497 | 0.0041504 | 0.0017454 | -0.002405 |  | 799.679 | 0.0010414
2 | YES | 799.679 | 34.2981 | 0.0003125 | 0.0010414 | 0.0017351 | 0.00069372 |  | 599.854 | 0.0022736
2 | YES | 599.854 | 34.1164 | 0.00045317 | 0.0022736 | 0.0017249 | -0.00054865 |  | 3000.29 | 0.0028992
3 | NO | 5.4685 | 32.1963 | 0.00091531 | -0.0068398 | 0.0016945 | 0.0085343 |  | 3000.21 | 0.0036964
4 | YES | 3000.29 | 34.6505 | 0.00026312 | 0.0028992 | 0.0018476 | -0.0010515 |  | 2500.13 | 0.0018616
4 | YES | 3000.21 | 34.6505 | 0.00027381 | 0.0036964 | 0.0018476 | -0.0018488 |  | 2000.45 | 0.0027008
4 | YES | 2500.13 | 34.6293 | 0.00023604 | 0.0018616 | 0.0018221 | -3.9501e-05 |  | 1499.81 | 0.0017471
4 | YES | 2000.45 | 34.5866 | 0.00025767 | 0.0027008 | 0.0017965 | -0.00090429 |  | 1249.77 | 0.0030098
4 | YES | 1499.81 | 34.5137 | 0.00022772 | 0.0017471 | 0.0017709 | 2.3796e-05 |  | 999.52 | 0.0030479
4 | YES | 1249.77 | 34.4476 | 0.00057954 | 0.0030098 | 0.0017581 | -0.0012517 |  | 2999.96 | 0.0017281
4 | YES | 999.52 | 34.3584 | 0.0004201 | 0.0030479 | 0.0017453 | -0.0013026 |  | 2499.82 | -4.9591e-05
5 | NO | 5.3043 | 32.2417 | 0.00050576 | -0.0027809 | 0.0016945 | 0.0044754 |  | 2250.09 | 0.0019607
6 | YES | 2999.96 | 34.6535 | 0.00023718 | 0.0017281 | 0.0018476 | 0.00011956 |  | 2000.94 | 0.0017166
6 | YES | 2499.82 | 34.6314 | 0.00028192 | -4.9591e-05 | 0.0018221 | 0.0018716 |  | 1750.52 | 0.0024796
6 | YES | 2250.09 | 34.6122 | 0.00027757 | 0.0019607 | 0.0018093 | -0.00015146 |  | 1249.06 | 0.0029259
6 | YES | 2000.94 | 34.587 | 0.00028738 | 0.0017166 | 0.0017965 | 7.9937e-05 | 34.5853 | 1000.6 | 0.0024185
6 | YES | 1750.52 | 34.5525 | 0.00025846 | 0.0024796 | 0.0017837 | -0.00069581 | 34.55 | 3499.35 | 0.0020256
6 | NO | 1500.07 | 34.5052 | 0.00026398 | -0.044304 | 0.0017709 | 0.046075 | 34.5495 | 3499.35 | 0.00022507
6 | YES | 1249.06 | 34.4428 | 0.00022036 | 0.0029259 | 0.0017581 | -0.0011678 | 34.4399 | 2999.91 | -0.00014114
6 | YES | 1000.6 | 34.3661 | 0.00025565 | 0.0024185 | 0.0017454 | -0.00067312 | 34.3637 | 2499.75 | -0.00040054
7 | NO | 5.3532 | 32.1732 | 0.0015561 | -0.0028954 | 0.0016945 | 0.0045899 | 32.1761 | 1999.93 | -0.001358
8 | YES | 3499.35 | 34.6676 | 0.00022761 | 0.0020256 | 0.0018732 | -0.00015245 | 34.6656 | 1500.56 | 0.00027847
8 | YES | 3499.35 | 34.6676 | 0.00022761 | 0.00022507 | 0.0018732 | 0.0016481 | 34.6674 | 1250.65 | -0.00022507
8 | YES | 2999.91 | 34.6555 | 0.00024229 | -0.00014114 | 0.0018476 | 0.0019888 | 34.6556 | 3499.66 | 0.0034218
8 | YES | 2499.75 | 34.6298 | 0.00024154 | -0.00040054 | 0.001822 | 0.0022226 | 34.6302 | 3499.75 | 0.0022202
8 | YES | 1999.93 | 34.583 | 0.00024225 | -0.001358 | 0.0017965 | 0.0031545 | 34.5844 | 2999.74 | 0.0037613
8 | YES | 1500.56 | 34.4998 | 0.00023291 | 0.00027847 | 0.001771 | 0.0014925 | 34.4995 | 2499.93 | 0.00095367
8 | YES | 1250.65 | 34.4257 | 0.00029704 | -0.00022507 | 0.0017582 | 0.0019833 | 34.4259 | 2000.03 | 0.0015488
8 | NO | 1000.6 | 34.3524 | 0.00045785 | -0.0043373 | 0.0017454 | 0.0060827 | 34.3567 | 1500.14 | 0.0020065
9 | NO | 4.7699 | 32.2292 | 0.00044036 | -0.0025902 | 0.0016945 | 0.0042847 |  | 1250.11 | 0.00095749
10 | NO | 5.4024 | 32.1532 | 0.00036699 | -0.0021172 | 0.0016945 | 0.0038117 |  | 1000.36 | 0.0013046
11 | YES | 3499.66 | 34.672 | 0.00024795 | 0.0034218 | 0.0018732 | -0.0015486 |  | 3498.41 | 0.0020218
11 | YES | 3499.75 | 34.672 | 0.00024423 | 0.0022202 | 0.0018732 | -0.00034697 |  | 3001.8 | 0.0019684
11 | YES | 2999.74 | 34.6573 | 0.00023957 | 0.0037613 | 0.0018476 | -0.0019137 |  | 2500.54 | 0.001648
11 | YES | 2499.93 | 34.6293 | 0.00026747 | 0.00095367 | 0.0018221 | 0.00086838 |  | 2001.61 | 0.0019722
11 | YES | 2000.03 | 34.5863 | 0.00024972 | 0.0015488 | 0.0017965 | 0.00024773 |  | 1751.17 | 0.0016861
11 | YES | 1500.14 | 34.4971 | 0.000246 | 0.0020065 | 0.0017709 | -0.00023559 |  | 1498.67 | 0.0013885
11 | YES | 1250.11 | 34.4398 | 0.00024341 | 0.00095749 | 0.0017582 | 0.00080067 |  | 1251.21 | -0.0009346
11 | YES | 1000.36 | 34.3694 | 0.0002431 | 0.0013046 | 0.0017454 | 0.00044076 |  | 1000.85 | 0.0022469
12 | NO | 5.1801 | 32.2206 | 0.0016194 | -0.0029449 | 0.0016945 | 0.0046395 |  | 3500.4 | 0.0027695
13 | YES | 3498.41 | 34.6731 | 0.00026936 | 0.0020218 | 0.0018731 | -0.00014868 |  | 3000.83 | 0.0032082
13 | YES | 3001.8 | 34.6575 | 0.00018429 | 0.0019684 | 0.0018477 | -0.00012066 |  | 2499.67 | 0.0020676
13 | YES | 2500.54 | 34.6325 | 0.00025497 | 0.001648 | 0.0018221 | 0.00017414 |  | 1999.87 | 0.0028954
13 | YES | 2001.61 | 34.5881 | 0.00023571 | 0.0019722 | 0.0017966 | -0.00017562 |  | 1749.93 | 0.0045013
13 | YES | 1751.17 | 34.5521 | 0.00025656 | 0.0016861 | 0.0017838 | 9.7677e-05 |  | 1501.2 | 0.0028801
13 | YES | 1498.67 | 34.4998 | 0.00026206 | 0.0013885 | 0.0017709 | 0.00038232 |  | 1250.05 | 0.001812
13 | YES | 1251.21 | 34.4411 | 0.00027371 | -0.0009346 | 0.0017582 | 0.0026928 |  | 999.642 | 0.0033722
13 | YES | 1000.85 | 34.3681 | 0.00033897 | 0.0022469 | 0.0017454 | -0.00050144 |  | 800.252 | 0.0006752
14 | NO | 5.3429 | 32.2355 | 0.00028454 | -0.0034065 | 0.0016945 | 0.005101 |  | 600.106 | 0.00061035
15 | YES | 3500.4 | 34.6733 | 0.00026642 | 0.0027695 | 0.0018732 | -0.00089626 |  | 3999.88 | 0.0014687
15 | YES | 3000.83 | 34.6578 | 0.00023569 | 0.0032082 | 0.0018477 | -0.0013605 |  | 3499.37 | 0.0007515
15 | YES | 2499.67 | 34.631 | 0.00024355 | 0.0020676 | 0.001822 | -0.00024552 |  | 3000.2 | 0.0010185
15 | YES | 1999.87 | 34.5835 | 0.00027324 | 0.0028954 | 0.0017965 | -0.0010989 |  | 2500.63 | 0.003128
15 | YES | 1749.93 | 34.5462 | 0.0002634 | 0.0045013 | 0.0017837 | -0.0027176 |  | 1999.7 | 0.00019455
15 | YES | 1501.2 | 34.4954 | 0.00026521 | 0.0028801 | 0.001771 | -0.0011091 |  | 1499.49 | -0.0010452
15 | YES | 1250.05 | 34.4339 | 0.00022898 | 0.001812 | 0.0017582 | -5.3823e-05 |  | 1250.79 | -0.00045395
15 | YES | 999.642 | 34.3562 | 0.00028994 | 0.0033722 | 0.0017454 | -0.0016268 |  | 1000.28 | 0.001236
15 | YES | 800.252 | 34.2803 | 0.00036131 | 0.0006752 | 0.0017352 | 0.00106 |  | 1250.14 | 0.00020599
15 | YES | 600.106 | 34.1639 | 0.00027824 | 0.00061035 | 0.0017249 | 0.0011146 |  | 1250.14 | 0.0012093
15 | NO | 400.367 | 33.9549 | 0.00069006 | -6.8665e-05 | 0.0017147 | 0.0017834 |  | 1000.07 | 0.0018311
15 | NO | 299.726 | 33.8503 | 0.00048147 | 0.0036316 | 0.0017096 | -0.001922 |  | 799.658 | 0.0015221 | 0.0017924 | 0.0013828
15 | NO | 250.304 | 33.8017 | 0.002572 | 0.0014648 | 0.001707 | 0.0002422
15 | NO | 200.078 | 33.761 | 0.0010698 | -0.00018692 | 0.0017045 | 0.0018914
15 | NO | 175.068 | 33.7591 | 0.0026202 | 0.002018 | 0.0017032 | -0.00031477
15 | NO | 150.394 | 33.6441 | 0.00073447 | -0.0044632 | 0.0017019 | 0.0061651
15 | NO | 124.669 | 33.3875 | 0.0011417 | -0.01796 | 0.0017006 | 0.01966
15 | NO | 99.9468 | 32.7263 | 0.0052047 | -0.026863 | 0.0016994 | 0.028562
15 | NO | 74.9348 | 32.4939 | 0.0040551 | -0.0035744 | 0.0016981 | 0.0052724
15 | NO | 49.6209 | 32.4068 | 0.0019605 | -0.0023994 | 0.0016968 | 0.0040962
15 | NO | 25.2305 | 32.2898 | 0.0020023 | -0.0058403 | 0.0016955 | 0.0075358
15 | NO | 10.0577 | 32.2922 | 0.00039167 | -0.00061035 | 0.0016948 | 0.0023051
15 | NO | 5.306 | 32.292 | 0.00029408 | -0.00072861 | 0.0016945 | 0.0024231
16 | NO | 5.0787 | 32.2627 | 0.0038586 | -0.0031319 | 0.0016945 | 0.0048264
17 | YES | 3999.88 | 34.682 | 0.0003076 | 0.0014687 | 0.0018987 | 0.00043009
17 | YES | 3499.37 | 34.672 | 0.00025167 | 0.0007515 | 0.0018732 | 0.0011217
17 | YES | 3000.2 | 34.6542 | 0.00026121 | 0.0010185 | 0.0018476 | 0.00082912
17 | YES | 2500.63 | 34.6254 | 0.0002522 | 0.003128 | 0.0018221 | -0.001306
17 | YES | 1999.7 | 34.5796 | 0.00027556 | 0.00019455 | 0.0017965 | 0.0016019
17 | YES | 1499.49 | 34.4947 | 0.00021468 | -0.0010452 | 0.0017709 | 0.0028161
17 | YES | 1250.79 | 34.438 | 0.00029052 | -0.00045395 | 0.0017582 | 0.0022121
17 | YES | 1000.28 | 34.3631 | 0.00028199 | 0.001236 | 0.0017454 | 0.00050943
18 | YES | 1250.14 | 34.4538 | 0.00023494 | 0.00020599 | 0.0017582 | 0.0015522
18 | YES | 1250.14 | 34.4538 | 0.00023494 | 0.0012093 | 0.0017582 | 0.0005489
18 | YES | 1000.07 | 34.3808 | 0.000379 | 0.0018311 | 0.0017454 | -8.5674e-05
18 | YES | 799.658 | 34.2804 | 0.00030276 | 0.0015221 | 0.0017351 | 0.00021307
18 | NO | 5.0711 | 31.9356 | 0.0002124 | -0.0024052 | 0.0016945 | 0.0040997
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.6146 |  | X | CALC Y
m= | -0.00036936 | 3.5733e-05 |  | Std Dev of Resid |  | 0.0016289 |  | 1 | -0.0054362
b= | -0.0050668 | 0.00040851 |  | Resid Sum of Squares |  | 0.00017777 |  | 18 | -0.011715
# of Values Extracted = |  |  | 94
# of Values Used in Fit = |  |  | 69
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1250.36 | 34.4614 | 0.00024659 | -0.0049591 | -0.0054362 | -0.00047704 |  | 1 | -0.0049591
1 | YES | 1250.37 | 34.4614 | 0.00024657 | -0.0040588 | -0.0054362 | -0.0013773 |  | 1 | -0.0040588
1 | YES | 999.627 | 34.3718 | 0.00022868 | -0.0024071 | -0.0054362 | -0.0030291 |  | 1 | -0.0024071
1 | YES | 799.713 | 34.2916 | 0.00058958 | -0.0030937 | -0.0054362 | -0.0023424 |  | 1 | -0.0030937
1 | NO | 5.0049 | 31.89 | 0.00048941 | -0.0070305 | -0.0054362 | 0.0015943 |  | 2 | -0.0047302
2 | YES | 2000.19 | 34.5825 | 0.00028374 | -0.0047302 | -0.0058055 | -0.0010753 |  | 2 | -0.006073
2 | YES | 1499.96 | 34.5056 | 0.00022655 | -0.006073 | -0.0058055 | 0.00026749 |  | 2 | -0.008728
2 | YES | 1249.68 | 34.4566 | 0.00037716 | -0.008728 | -0.0058055 | 0.0029225 |  | 2 | -0.0040321
2 | YES | 1000.24 | 34.3758 | 0.00030484 | -0.0040321 | -0.0058055 | -0.0017734 |  | 2 | -0.0068474
2 | YES | 799.679 | 34.2903 | 0.00037382 | -0.0068474 | -0.0058055 | 0.0010419 |  | 2 | -0.0052338
2 | YES | 599.854 | 34.1089 | 0.00057911 | -0.0052338 | -0.0058055 | -0.00057175 |  | 4 | -0.0068817
3 | NO | 5.4685 | 32.1889 | 0.00073642 | -0.01424 | -0.0061749 | 0.0080654 |  | 4 | -0.0060844
4 | YES | 3000.29 | 34.6407 | 0.00024948 | -0.0068817 | -0.0065442 | 0.00033748 |  | 4 | -0.0080261
4 | YES | 3000.21 | 34.6407 | 0.00024784 | -0.0060844 | -0.0065442 | -0.00045979 |  | 4 | -0.0068245
4 | YES | 2500.13 | 34.6194 | 0.00030092 | -0.0080261 | -0.0065442 | 0.0014819 |  | 4 | -0.0075417
4 | YES | 2000.45 | 34.5771 | 0.00029937 | -0.0068245 | -0.0065442 | 0.00028026 |  | 4 | -0.0059051
4 | YES | 1499.81 | 34.5045 | 0.00030212 | -0.0075417 | -0.0065442 | 0.00099743 |  | 4 | -0.0062637
4 | YES | 1249.77 | 34.4387 | 0.00059558 | -0.0059051 | -0.0065442 | -0.00063908 |  | 6 | -0.0086136
4 | YES | 999.52 | 34.3491 | 0.00050909 | -0.0062637 | -0.0065442 | -0.0002805 |  | 6 | -0.010239
5 | NO | 5.3043 | 32.2344 | 0.00033862 | -0.01012 | -0.0069136 | 0.0032068 |  | 6 | -0.0081139
6 | YES | 2999.96 | 34.6432 | 0.0002571 | -0.0086136 | -0.0072829 | 0.0013306 |  | 6 | -0.0081634
6 | YES | 2499.82 | 34.6212 | 0.00032784 | -0.010239 | -0.0072829 | 0.0029557 |  | 6 | -0.0073509
6 | YES | 2250.09 | 34.6021 | 0.00029237 | -0.0081139 | -0.0072829 | 0.00083091 |  | 6 | -0.006485
6 | YES | 2000.94 | 34.5771 | 0.00030034 | -0.0081634 | -0.0072829 | 0.0008805 |  | 6 | -0.0071602
6 | YES | 1750.52 | 34.5426 | 0.00030858 | -0.0073509 | -0.0072829 | 6.797e-05 |  | 8 | -0.0082436
6 | NO | 1500.07 | 34.4956 | 0.00026341 | -0.053925 | -0.0072829 | 0.046642 |  | 8 | -0.010044
6 | YES | 1249.06 | 34.4334 | 0.00032533 | -0.006485 | -0.0072829 | -0.00079796 |  | 8 | -0.010407
6 | YES | 1000.6 | 34.3565 | 0.00029161 | -0.0071602 | -0.0072829 | -0.00012276 |  | 8 | -0.010387
7 | NO | 5.3532 | 32.1664 | 0.0011043 | -0.009716 | -0.0076523 | 0.0020637 |  | 8 | -0.011082
8 | YES | 3499.35 | 34.6574 | 0.00024075 | -0.0082436 | -0.0080217 | 0.00022189 |  | 8 | -0.0093498
8 | YES | 3499.35 | 34.6574 | 0.00024075 | -0.010044 | -0.0080217 | 0.0020224 |  | 8 | -0.0094109
8 | YES | 2999.91 | 34.6452 | 0.00026657 | -0.010407 | -0.0080217 | 0.0023848 |  | 11 | -0.0076561
8 | YES | 2499.75 | 34.6198 | 0.00026929 | -0.010387 | -0.0080217 | 0.0023657 |  | 11 | -0.0088501
8 | YES | 1999.93 | 34.5733 | 0.00027212 | -0.011082 | -0.0080217 | 0.00306 |  | 11 | -0.0071182
8 | YES | 1500.56 | 34.4902 | 0.00027416 | -0.0093498 | -0.0080217 | 0.0013282 |  | 11 | -0.0097084
8 | YES | 1250.65 | 34.4165 | 0.00026316 | -0.0094109 | -0.0080217 | 0.0013892 |  | 11 | -0.0089226
8 | NO | 1000.6 | 34.343 | 0.00027057 | -0.01371 | -0.0080217 | 0.0056883 |  | 11 | -0.008255
9 | NO | 4.7699 | 32.2212 | 0.00031361 | -0.010624 | -0.008391 | 0.0022329 |  | 11 | -0.0093193
10 | NO | 5.4024 | 32.1451 | 0.00024311 | -0.010159 | -0.0087604 | 0.0013981 |  | 11 | -0.0090561
11 | YES | 3499.66 | 34.6609 | 0.00025901 | -0.0076561 | -0.0091297 | -0.0014736 |  | 13 | -0.0091705
11 | YES | 3499.75 | 34.6609 | 0.00025512 | -0.0088501 | -0.0091297 | -0.00027965 |  | 13 | -0.0091438
11 | YES | 2999.74 | 34.6464 | 0.00029034 | -0.0071182 | -0.0091297 | -0.0020115 |  | 13 | -0.0093193
11 | YES | 2499.93 | 34.6186 | 0.00027798 | -0.0097084 | -0.0091297 | 0.00057865 |  | 13 | -0.0088234
11 | YES | 2000.03 | 34.5758 | 0.0002885 | -0.0089226 | -0.0091297 | -0.00020717 |  | 13 | -0.0089836
11 | YES | 1500.14 | 34.4868 | 0.00027064 | -0.008255 | -0.0091297 | -0.00087475 |  | 13 | -0.0092277
11 | YES | 1250.11 | 34.4295 | 0.00026498 | -0.0093193 | -0.0091297 | 0.00018956 |  | 13 | -0.011402
11 | YES | 1000.36 | 34.359 | 0.00030194 | -0.0090561 | -0.0091297 | -7.3655e-05 |  | 13 | -0.0082283
12 | NO | 5.1801 | 32.2119 | 0.001954 | -0.01162 | -0.0094991 | 0.0021205 |  | 15 | -0.0091515
13 | YES | 3498.41 | 34.6619 | 0.00022559 | -0.0091705 | -0.0098685 | -0.00069793 |  | 15 | -0.0085907
13 | YES | 3001.8 | 34.6464 | 0.00025221 | -0.0091438 | -0.0098685 | -0.00072463 |  | 15 | -0.0094185
13 | YES | 2500.54 | 34.6216 | 0.00028193 | -0.0093193 | -0.0098685 | -0.00054915 |  | 15 | -0.0084801
13 | YES | 2001.61 | 34.5773 | 0.00028049 | -0.0088234 | -0.0098685 | -0.0010451 |  | 15 | -0.0069466
13 | YES | 1751.17 | 34.5414 | 0.0003278 | -0.0089836 | -0.0098685 | -0.00088485 |  | 15 | -0.0081673
13 | YES | 1498.67 | 34.4892 | 0.00038622 | -0.0092277 | -0.0098685 | -0.00064071 |  | 15 | -0.0093498
13 | YES | 1251.21 | 34.4306 | 0.00029461 | -0.011402 | -0.0098685 | 0.0015336 |  | 15 | -0.0074844
13 | YES | 1000.85 | 34.3577 | 0.00029694 | -0.0082283 | -0.0098685 | -0.0016402 |  | 15 | -0.010372
14 | NO | 5.3429 | 32.2259 | 0.00022853 | -0.01302 | -0.010238 | 0.0027818 |  | 15 | -0.010094
15 | YES | 3500.4 | 34.6614 | 0.00026366 | -0.0091515 | -0.010607 | -0.0014557 |  | 17 | -0.011379
15 | YES | 3000.83 | 34.646 | 0.0002615 | -0.0085907 | -0.010607 | -0.0020165 |  | 17 | -0.012188
15 | YES | 2499.67 | 34.6195 | 0.00028989 | -0.0094185 | -0.010607 | -0.0011887 |  | 17 | -0.011768
15 | YES | 1999.87 | 34.5721 | 0.00028669 | -0.0084801 | -0.010607 | -0.0021271 |  | 17 | -0.0094795
15 | YES | 1749.93 | 34.5348 | 0.00032985 | -0.0069466 | -0.010607 | -0.0036606 |  | 17 | -0.012165
15 | YES | 1501.2 | 34.4843 | 0.00034906 | -0.0081673 | -0.010607 | -0.0024399 |  | 17 | -0.013138
15 | YES | 1250.05 | 34.4227 | 0.0002612 | -0.0093498 | -0.010607 | -0.0012574 |  | 17 | -0.012497
15 | YES | 999.642 | 34.3453 | 0.00044464 | -0.0074844 | -0.010607 | -0.0031227 |  | 17 | -0.01086
15 | YES | 800.252 | 34.2692 | 0.00036911 | -0.010372 | -0.010607 | -0.00023498 |  | 18 | -0.015678
15 | YES | 600.106 | 34.1532 | 0.00029369 | -0.010094 | -0.010607 | -0.00051348 |  | 18 | -0.014675
15 | NO | 400.367 | 33.9451 | 0.00034766 | -0.0099449 | -0.010607 | -0.00066226 |  | 18 | -0.013882
15 | NO | 299.726 | 33.8395 | 0.00033872 | -0.0071907 | -0.010607 | -0.0034165 |  | 18 | -0.013515
15 | NO | 250.304 | 33.7925 | 0.0031109 | -0.00774 | -0.010607 | -0.0028672
15 | NO | 200.078 | 33.7508 | 0.00075999 | -0.010433 | -0.010607 | -0.00017398
15 | NO | 175.068 | 33.7481 | 0.0020474 | -0.0090141 | -0.010607 | -0.0015931
15 | NO | 150.394 | 33.6359 | 0.0015743 | -0.012688 | -0.010607 | 0.0020805
15 | NO | 124.669 | 33.3803 | 0.0031342 | -0.025158 | -0.010607 | 0.014551
15 | NO | 99.9468 | 32.7148 | 0.010908 | -0.038395 | -0.010607 | 0.027788
15 | NO | 74.9348 | 32.4834 | 0.0033667 | -0.014065 | -0.010607 | 0.0034576
15 | NO | 49.6209 | 32.3965 | 0.0014084 | -0.012699 | -0.010607 | 0.0020919
15 | NO | 25.2305 | 32.2798 | 0.0021184 | -0.015831 | -0.010607 | 0.0052238
15 | NO | 10.0577 | 32.2829 | 0.0002693 | -0.0099487 | -0.010607 | -0.00065845
15 | NO | 5.306 | 32.2825 | 0.00025707 | -0.010197 | -0.010607 | -0.00041048
16 | NO | 5.0787 | 32.2532 | 0.0016025 | -0.012608 | -0.010977 | 0.0016311
17 | YES | 3999.88 | 34.6691 | 0.00024898 | -0.011379 | -0.011346 | 3.33e-05
17 | YES | 3499.37 | 34.659 | 0.00023864 | -0.012188 | -0.011346 | 0.0008421
17 | YES | 3000.2 | 34.6414 | 0.00029117 | -0.011768 | -0.011346 | 0.0004224
17 | YES | 2500.63 | 34.6128 | 0.00033143 | -0.0094795 | -0.011346 | -0.0018664
17 | YES | 1999.7 | 34.5672 | 0.00031966 | -0.012165 | -0.011346 | 0.0008192
17 | YES | 1499.49 | 34.4826 | 0.00027113 | -0.013138 | -0.011346 | 0.0017919
17 | YES | 1250.79 | 34.426 | 0.00029816 | -0.012497 | -0.011346 | 0.001151
17 | YES | 1000.28 | 34.351 | 0.00032816 | -0.01086 | -0.011346 | -0.0004855
18 | YES | 1250.14 | 34.4379 | 0.00025296 | -0.015678 | -0.011715 | 0.0039631
18 | YES | 1250.14 | 34.4379 | 0.00025296 | -0.014675 | -0.011715 | 0.0029598
18 | YES | 1000.07 | 34.3651 | 0.00039911 | -0.013882 | -0.011715 | 0.0021664
18 | YES | 799.658 | 34.2654 | 0.00037259 | -0.013515 | -0.011715 | 0.0018002
18 | NO | 5.0711 | 31.9219 | 0.00021114 | -0.016075 | -0.011715 | 0.0043598
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.075925 |  | X | CALC Y
m= | -6.8928e-05 | 2.9378e-05 |  | Std Dev of Resid |  | 0.0013392 |  | 1 | 0.0024147
b= | 0.0024836 | 0.00033586 |  | Resid Sum of Squares |  | 0.00012016 |  | 18 | 0.0012429
# of Values Extracted = |  |  | 94
# of Values Used in Fit = |  |  | 69
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1250.36 | 34.4689 | 0.0002361 | 0.0025063 | 0.0024147 | -9.1551e-05 |  | 1 | 0.0025063
1 | YES | 1250.37 | 34.4689 | 0.00023576 | 0.0034065 | 0.0024147 | -0.00099181 |  | 1 | 0.0034065
1 | YES | 999.627 | 34.3791 | 0.0002685 | 0.0048866 | 0.0024147 | -0.0024719 |  | 1 | 0.0048866
1 | YES | 799.713 | 34.2978 | 0.00039483 | 0.0031319 | 0.0024147 | -0.00071716 |  | 1 | 0.0031319
1 | NO | 5.0049 | 31.8963 | 0.00046049 | -0.00065994 | 0.0024147 | 0.0030747 |  | 2 | 0.0040474
2 | YES | 2000.19 | 34.5912 | 0.00023589 | 0.0040474 | 0.0023458 | -0.0017016 |  | 2 | 0.0022011
2 | YES | 1499.96 | 34.5139 | 0.00023813 | 0.0022011 | 0.0023458 | 0.0001447 |  | 2 | -0.00054932
2 | YES | 1249.68 | 34.4648 | 0.00038171 | -0.00054932 | 0.0023458 | 0.0028951 |  | 2 | 0.0041504
2 | YES | 1000.24 | 34.3839 | 0.00030497 | 0.0041504 | 0.0023458 | -0.0018046 |  | 2 | 0.0010414
2 | YES | 799.679 | 34.2981 | 0.0003125 | 0.0010414 | 0.0023458 | 0.0013044 |  | 2 | 0.0022736
2 | YES | 599.854 | 34.1164 | 0.00045317 | 0.0022736 | 0.0023458 | 7.222e-05 |  | 4 | 0.0028992
3 | NO | 5.4685 | 32.1963 | 0.00091531 | -0.0068398 | 0.0022769 | 0.0091166 |  | 4 | 0.0036964
4 | YES | 3000.29 | 34.6505 | 0.00026312 | 0.0028992 | 0.0022079 | -0.00069125 |  | 4 | 0.0018616
4 | YES | 3000.21 | 34.6505 | 0.00027381 | 0.0036964 | 0.0022079 | -0.0014885 |  | 4 | 0.0027008
4 | YES | 2500.13 | 34.6293 | 0.00023604 | 0.0018616 | 0.0022079 | 0.00034635 |  | 4 | 0.0017471
4 | YES | 2000.45 | 34.5866 | 0.00025767 | 0.0027008 | 0.0022079 | -0.00049289 |  | 4 | 0.0030098
4 | YES | 1499.81 | 34.5137 | 0.00022772 | 0.0017471 | 0.0022079 | 0.00046079 |  | 4 | 0.0030479
4 | YES | 1249.77 | 34.4476 | 0.00057954 | 0.0030098 | 0.0022079 | -0.00080188 |  | 6 | 0.0017281
4 | YES | 999.52 | 34.3584 | 0.0004201 | 0.0030479 | 0.0022079 | -0.00084002 |  | 6 | -4.9591e-05
5 | NO | 5.3043 | 32.2417 | 0.00050576 | -0.0027809 | 0.002139 | 0.0049199 |  | 6 | 0.0019607
6 | YES | 2999.96 | 34.6535 | 0.00023718 | 0.0017281 | 0.0020701 | 0.00034201 |  | 6 | 0.0017166
6 | YES | 2499.82 | 34.6314 | 0.00028192 | -4.9591e-05 | 0.0020701 | 0.0021197 |  | 6 | 0.0024796
6 | YES | 2250.09 | 34.6122 | 0.00027757 | 0.0019607 | 0.0020701 | 0.00010932 |  | 6 | 0.0029259
6 | YES | 2000.94 | 34.587 | 0.00028738 | 0.0017166 | 0.0020701 | 0.00035346 |  | 6 | 0.0024185
6 | YES | 1750.52 | 34.5525 | 0.00025846 | 0.0024796 | 0.0020701 | -0.00040948 |  | 8 | 0.0020256
6 | NO | 1500.07 | 34.5052 | 0.00026398 | -0.044304 | 0.0020701 | 0.046374 |  | 8 | 0.00022507
6 | YES | 1249.06 | 34.4428 | 0.00022036 | 0.0029259 | 0.0020701 | -0.0008558 |  | 8 | -0.00014114
6 | YES | 1000.6 | 34.3661 | 0.00025565 | 0.0024185 | 0.0020701 | -0.00034845 |  | 8 | -0.00040054
7 | NO | 5.3532 | 32.1732 | 0.0015561 | -0.0028954 | 0.0020011 | 0.0048965 |  | 8 | -0.001358
8 | YES | 3499.35 | 34.6676 | 0.00022761 | 0.0020256 | 0.0019322 | -9.3391e-05 |  | 8 | 0.00027847
8 | YES | 3499.35 | 34.6676 | 0.00022761 | 0.00022507 | 0.0019322 | 0.0017071 |  | 8 | -0.00022507
8 | YES | 2999.91 | 34.6555 | 0.00024229 | -0.00014114 | 0.0019322 | 0.0020734 |  | 11 | 0.0034218
8 | YES | 2499.75 | 34.6298 | 0.00024154 | -0.00040054 | 0.0019322 | 0.0023328 |  | 11 | 0.0022202
8 | YES | 1999.93 | 34.583 | 0.00024225 | -0.001358 | 0.0019322 | 0.0032902 |  | 11 | 0.0037613
8 | YES | 1500.56 | 34.4998 | 0.00023291 | 0.00027847 | 0.0019322 | 0.0016537 |  | 11 | 0.00095367
8 | YES | 1250.65 | 34.4257 | 0.00029704 | -0.00022507 | 0.0019322 | 0.0021573 |  | 11 | 0.0015488
8 | NO | 1000.6 | 34.3524 | 0.00045785 | -0.0043373 | 0.0019322 | 0.0062695 |  | 11 | 0.0020065
9 | NO | 4.7699 | 32.2292 | 0.00044036 | -0.0025902 | 0.0018633 | 0.0044535 |  | 11 | 0.00095749
10 | NO | 5.4024 | 32.1532 | 0.00036699 | -0.0021172 | 0.0017944 | 0.0039115 |  | 11 | 0.0013046
11 | YES | 3499.66 | 34.672 | 0.00024795 | 0.0034218 | 0.0017254 | -0.0016964 |  | 13 | 0.0020218
11 | YES | 3499.75 | 34.672 | 0.00024423 | 0.0022202 | 0.0017254 | -0.00049473 |  | 13 | 0.0019684
11 | YES | 2999.74 | 34.6573 | 0.00023957 | 0.0037613 | 0.0017254 | -0.0020359 |  | 13 | 0.001648
11 | YES | 2499.93 | 34.6293 | 0.00026747 | 0.00095367 | 0.0017254 | 0.00077175 |  | 13 | 0.0019722
11 | YES | 2000.03 | 34.5863 | 0.00024972 | 0.0015488 | 0.0017254 | 0.00017665 |  | 13 | 0.0016861
11 | YES | 1500.14 | 34.4971 | 0.000246 | 0.0020065 | 0.0017254 | -0.00028111 |  | 13 | 0.0013885
11 | YES | 1250.11 | 34.4398 | 0.00024341 | 0.00095749 | 0.0017254 | 0.00076794 |  | 13 | -0.0009346
11 | YES | 1000.36 | 34.3694 | 0.0002431 | 0.0013046 | 0.0017254 | 0.00042079 |  | 13 | 0.0022469
12 | NO | 5.1801 | 32.2206 | 0.0016194 | -0.0029449 | 0.0016565 | 0.0046014 |  | 15 | 0.0027695
13 | YES | 3498.41 | 34.6731 | 0.00026936 | 0.0020218 | 0.0015876 | -0.00043422 |  | 15 | 0.0032082
13 | YES | 3001.8 | 34.6575 | 0.00018429 | 0.0019684 | 0.0015876 | -0.00038081 |  | 15 | 0.0020676
13 | YES | 2500.54 | 34.6325 | 0.00025497 | 0.001648 | 0.0015876 | -6.0383e-05 |  | 15 | 0.0028954
13 | YES | 2001.61 | 34.5881 | 0.00023571 | 0.0019722 | 0.0015876 | -0.00038463 |  | 15 | 0.0045013
13 | YES | 1751.17 | 34.5521 | 0.00025656 | 0.0016861 | 0.0015876 | -9.8533e-05 |  | 15 | 0.0028801
13 | YES | 1498.67 | 34.4998 | 0.00026206 | 0.0013885 | 0.0015876 | 0.00019902 |  | 15 | 0.001812
13 | YES | 1251.21 | 34.4411 | 0.00027371 | -0.0009346 | 0.0015876 | 0.0025222 |  | 15 | 0.0033722
13 | YES | 1000.85 | 34.3681 | 0.00033897 | 0.0022469 | 0.0015876 | -0.00065929 |  | 15 | 0.0006752
14 | NO | 5.3429 | 32.2355 | 0.00028454 | -0.0034065 | 0.0015186 | 0.0049252 |  | 15 | 0.00061035
15 | YES | 3500.4 | 34.6733 | 0.00026642 | 0.0027695 | 0.0014497 | -0.0013198 |  | 17 | 0.0014687
15 | YES | 3000.83 | 34.6578 | 0.00023569 | 0.0032082 | 0.0014497 | -0.0017584 |  | 17 | 0.0007515
15 | YES | 2499.67 | 34.631 | 0.00024355 | 0.0020676 | 0.0014497 | -0.00061786 |  | 17 | 0.0010185
15 | YES | 1999.87 | 34.5835 | 0.00027324 | 0.0028954 | 0.0014497 | -0.0014456 |  | 17 | 0.003128
15 | YES | 1749.93 | 34.5462 | 0.0002634 | 0.0045013 | 0.0014497 | -0.0030516 |  | 17 | 0.00019455
15 | YES | 1501.2 | 34.4954 | 0.00026521 | 0.0028801 | 0.0014497 | -0.0014304 |  | 17 | -0.0010452
15 | YES | 1250.05 | 34.4339 | 0.00022898 | 0.001812 | 0.0014497 | -0.00036227 |  | 17 | -0.00045395
15 | YES | 999.642 | 34.3562 | 0.00028994 | 0.0033722 | 0.0014497 | -0.0019225 |  | 17 | 0.001236
15 | YES | 800.252 | 34.2803 | 0.00036131 | 0.0006752 | 0.0014497 | 0.00077451 |  | 18 | 0.00020599
15 | YES | 600.106 | 34.1639 | 0.00027824 | 0.00061035 | 0.0014497 | 0.00083936 |  | 18 | 0.0012093
15 | NO | 400.367 | 33.9549 | 0.00069006 | -6.8665e-05 | 0.0014497 | 0.0015184 |  | 18 | 0.0018311
15 | NO | 299.726 | 33.8503 | 0.00048147 | 0.0036316 | 0.0014497 | -0.0021819 |  | 18 | 0.0015221
15 | NO | 250.304 | 33.8017 | 0.002572 | 0.0014648 | 0.0014497 | -1.513e-05
15 | NO | 200.078 | 33.761 | 0.0010698 | -0.00018692 | 0.0014497 | 0.0016366
15 | NO | 175.068 | 33.7591 | 0.0026202 | 0.002018 | 0.0014497 | -0.00056826
15 | NO | 150.394 | 33.6441 | 0.00073447 | -0.0044632 | 0.0014497 | 0.0059129
15 | NO | 124.669 | 33.3875 | 0.0011417 | -0.01796 | 0.0014497 | 0.019409
15 | NO | 99.9468 | 32.7263 | 0.0052047 | -0.026863 | 0.0014497 | 0.028313
15 | NO | 74.9348 | 32.4939 | 0.0040551 | -0.0035744 | 0.0014497 | 0.0050241
15 | NO | 49.6209 | 32.4068 | 0.0019605 | -0.0023994 | 0.0014497 | 0.0038492
15 | NO | 25.2305 | 32.2898 | 0.0020023 | -0.0058403 | 0.0014497 | 0.00729
15 | NO | 10.0577 | 32.2922 | 0.00039167 | -0.00061035 | 0.0014497 | 0.0020601
15 | NO | 5.306 | 32.292 | 0.00029408 | -0.00072861 | 0.0014497 | 0.0021783
16 | NO | 5.0787 | 32.2627 | 0.0038586 | -0.0031319 | 0.0013808 | 0.0045127
17 | YES | 3999.88 | 34.682 | 0.0003076 | 0.0014687 | 0.0013119 | -0.00015681
17 | YES | 3499.37 | 34.672 | 0.00025167 | 0.0007515 | 0.0013119 | 0.00056036
17 | YES | 3000.2 | 34.6542 | 0.00026121 | 0.0010185 | 0.0013119 | 0.00029333
17 | YES | 2500.63 | 34.6254 | 0.0002522 | 0.003128 | 0.0013119 | -0.0018162
17 | YES | 1999.7 | 34.5796 | 0.00027556 | 0.00019455 | 0.0013119 | 0.0011173
17 | YES | 1499.49 | 34.4947 | 0.00021468 | -0.0010452 | 0.0013119 | 0.0023571
17 | YES | 1250.79 | 34.438 | 0.00029052 | -0.00045395 | 0.0013119 | 0.0017658
17 | YES | 1000.28 | 34.3631 | 0.00028199 | 0.001236 | 0.0013119 | 7.5893e-05
18 | YES | 1250.14 | 34.4538 | 0.00023494 | 0.00020599 | 0.0012429 | 0.0010369
18 | YES | 1250.14 | 34.4538 | 0.00023494 | 0.0012093 | 0.0012429 | 3.3665e-05
18 | YES | 1000.07 | 34.3808 | 0.000379 | 0.0018311 | 0.0012429 | -0.00058813
18 | YES | 799.658 | 34.2804 | 0.00030276 | 0.0015221 | 0.0012429 | -0.00027914
18 | NO | 5.0711 | 31.9356 | 0.0002124 | -0.0024052 | 0.0012429 | 0.0036481
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE