## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2021/02/24 08:55:38 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 0.5
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2020-009-0022.mrg | 2020-009-0022.sam | 241 | 5.503 | 5.5033 | 0.18477 | 0.00024939 | 31.8825 | 31.8902 | 0.0064545 | 0.0077324 | 31.8825 | 31.8993 | 0.006048 | 0.016819
2020-009-0023.mrg | 2020-009-0023.sam | 241 | 5.401 | 5.4007 | 0.08729 | -0.00025749 | 31.2645 | 31.2491 | 0.013314 | -0.01536 | 31.2645 | 31.2652 | 0.020152 | 0.00069237
2020-009-0027.mrg | 2020-009-0027.sam | 239 | 400.684 | 400.684 | 0.31356 | 0.00012207 | 34.0365 | 34.0217 | 0.00025242 | -0.014763 | 34.0365 | 34.0375 | 0.00024373 | 0.00098038
2020-009-0029.mrg | 2020-009-0029.sam | 170 | 999.865 | 999.769 | 0.3163 | -0.096191 | 34.3793 | 34.3647 | 0.00031769 | -0.014614 | 34.3793 | 34.3802 | 0.00032798 | 0.00089264
2020-009-0029.mrg | 2020-009-0029.sam | 241 | 749.78 | 749.78 | 0.2304 | 0.00030518 | 34.2469 | 34.2324 | 0.00068821 | -0.014465 | 34.2469 | 34.2478 | 0.00078512 | 0.00085449
2020-009-0029.mrg | 2020-009-0029.sam | 240 | 499.919 | 499.917 | 0.19544 | -0.0021057 | 34.0562 | 34.0443 | 0.00059345 | -0.011948 | 34.0562 | 34.0597 | 0.00074487 | 0.0034943
2020-009-0029.mrg | 2020-009-0029.sam | 196 | 399.966 | 399.818 | 0.29036 | -0.1481 | 34.0432 | 34.0009 | 0.0015504 | -0.042351 | 34.0432 | 34.0174 | 0.001251 | -0.025841
2020-009-0029.mrg | 2020-009-0029.sam | 241 | 5.003 | 5.003 | 0.12577 | -1.2398e-05 | 31.8842 | 31.8691 | 0.00023901 | -0.015137 | 31.8842 | 31.8831 | 0.00019925 | -0.00107
2020-009-0033.mrg | 2020-009-0033.sam | 212 | 1600.76 | 1600.79 | 0.284 | 0.030029 | 34.5324 | 34.5195 | 0.0003023 | -0.012897 | 34.5324 | 34.536 | 0.00026636 | 0.0036469
2020-009-0033.mrg | 2020-009-0033.sam | 240 | 1499.99 | 1499.99 | 0.19701 | -0.0019531 | 34.5173 | 34.5022 | 0.00025571 | -0.015114 | 34.5173 | 34.5186 | 0.00027346 | 0.0013466
2020-009-0033.mrg | 2020-009-0033.sam | 239 | 1000.35 | 1000.35 | 0.27828 | -0.0049438 | 34.3829 | 34.3694 | 0.00026899 | -0.013493 | 34.3829 | 34.3853 | 0.000337 | 0.0024338
2020-009-0033.mrg | 2020-009-0033.sam | 241 | 5.038 | 5.0381 | 0.20288 | 0.00010347 | 31.9062 | 31.8884 | 0.00053255 | -0.017769 | 31.9062 | 31.9029 | 0.00077753 | -0.0032654
2020-009-0035.mrg | 2020-009-0035.sam | 223 | 2004.91 | 2004.87 | 0.2688 | -0.038086 | 34.5932 | 34.5799 | 0.00026907 | -0.013332 | 34.5932 | 34.5967 | 0.00024344 | 0.0034866
2020-009-0035.mrg | 2020-009-0035.sam | 241 | 1499.69 | 1499.69 | 0.17402 | -0.00012207 | 34.5162 | 34.5022 | 0.00026082 | -0.01403 | 34.5162 | 34.5184 | 0.00025723 | 0.002243
2020-009-0035.mrg | 2020-009-0035.sam | 241 | 1000.7 | 1000.7 | 0.13116 | 0.00036621 | 34.3769 | 34.363 | 0.00023335 | -0.013901 | 34.3769 | 34.3789 | 0.00023092 | 0.0020256
2020-009-0039.mrg | 2020-009-0039.sam | 219 | 1250.37 | 1250.42 | 0.25559 | 0.047974 | 34.4514 | 34.4376 | 0.0002446 | -0.01384 | 34.4514 | 34.4539 | 0.00024976 | 0.0024834
2020-009-0039.mrg | 2020-009-0039.sam | 241 | 1000.45 | 1000.45 | 0.25176 | -0.00024414 | 34.386 | 34.3714 | 0.00050775 | -0.01458 | 34.386 | 34.3876 | 0.00037061 | 0.0016213
2020-009-0039.mrg | 2020-009-0039.sam | 241 | 750.215 | 750.215 | 0.14833 | -0.00018311 | 34.2733 | 34.2573 | 0.00024963 | -0.016006 | 34.2733 | 34.2732 | 0.0002485 | -0.00013351
2020-009-0039.mrg | 2020-009-0039.sam | 241 | 299.815 | 299.815 | 0.23382 | 0.00042725 | 33.9346 | 33.9196 | 0.00026635 | -0.014969 | 33.9346 | 33.9354 | 0.00029319 | 0.00079346
2020-009-0039.mrg | 2020-009-0039.sam | 241 | 200.472 | 200.472 | 0.20522 | -0.00039673 | 33.8955 | 33.8798 | 0.0032208 | -0.015739 | 33.8955 | 33.8959 | 0.00031884 | 0.0004425
2020-009-0039.mrg | 2020-009-0039.sam | 241 | 4.902 | 4.9019 | 0.16442 | -8.7261e-05 | 31.9559 | 31.9393 | 0.0002229 | -0.016613 | 31.9559 | 31.954 | 0.00025289 | -0.0019417
2020-009-0043.mrg | 2020-009-0043.sam | 241 | 500.871 | 500.871 | 0.15005 | -0.00027466 | 34.0871 | 34.0732 | 0.00026028 | -0.013901 | 34.0871 | 34.0887 | 0.00029972 | 0.0016403
2020-009-0043.mrg | 2020-009-0043.sam | 241 | 400.516 | 400.516 | 0.2126 | -0.00036621 | 34.042 | 34.0269 | 0.00024055 | -0.015083 | 34.042 | 34.0425 | 0.00022015 | 0.00054169
2020-009-0057.mrg | 2020-009-0057.sam | 241 | 4.84 | 4.84 | 0.15892 | 2.0504e-05 | 30.2683 | 30.2488 | 0.0048652 | -0.019524 | 30.2683 | 30.2616 | 0.0043393 | -0.0066834
2020-009-0069.mrg | 2020-009-0069.sam | 241 | 500.44 | 500.44 | 0.16817 | -9.1553e-05 | 34.0653 | 34.0507 | 0.00041012 | -0.014561 | 34.0653 | 34.0667 | 0.00074509 | 0.0014114
2020-009-0077.mrg | 2020-009-0077.sam | 241 | 1249.9 | 1249.9 | 0.12596 | 0 | 34.4656 | 34.4502 | 0.0014198 | -0.01535 | 34.4656 | 34.4667 | 0.00066694 | 0.0010757
2020-009-0077.mrg | 2020-009-0077.sam | 241 | 1000.15 | 1000.15 | 0.11939 | -6.1035e-05 | 34.386 | 34.3722 | 0.00021775 | -0.013794 | 34.386 | 34.3878 | 0.00025064 | 0.0017891
2020-009-0077.mrg | 2020-009-0077.sam | 241 | 749.87 | 749.87 | 0.090114 | 0.00024414 | 34.2547 | 34.2412 | 0.00087882 | -0.013508 | 34.2547 | 34.2566 | 0.00062783 | 0.0019455
2020-009-0077.mrg | 2020-009-0077.sam | 241 | 499.977 | 499.977 | 0.098468 | 0.00012207 | 34.0933 | 34.0784 | 0.00028814 | -0.014938 | 34.0933 | 34.0937 | 0.00024801 | 0.00035858
2020-009-0078.mrg | 2020-009-0078.sam | 241 | 1999.98 | 1999.98 | 0.20858 | -0.00012207 | 34.5967 | 34.5829 | 0.00029701 | -0.013817 | 34.5967 | 34.5995 | 0.00026104 | 0.0028076
2020-009-0078.mrg | 2020-009-0078.sam | 241 | 1500.14 | 1500.14 | 0.11216 | 0.00012207 | 34.5073 | 34.4927 | 0.0002911 | -0.014572 | 34.5073 | 34.5089 | 0.00029375 | 0.0016174
2020-009-0078.mrg | 2020-009-0078.sam | 241 | 1249.94 | 1249.94 | 0.1624 | 0.00024414 | 34.462 | 34.4495 | 0.00025847 | -0.012516 | 34.462 | 34.4656 | 0.00023307 | 0.0036163
2020-009-0078.mrg | 2020-009-0078.sam | 241 | 1000.25 | 1000.25 | 0.10238 | -0.00030518 | 34.3855 | 34.3696 | 0.00025999 | -0.015926 | 34.3855 | 34.3853 | 0.00024216 | -0.00016403
2020-009-0078.mrg | 2020-009-0078.sam | 241 | 4.443 | 4.4433 | 0.13071 | 0.00035286 | 32.0988 | 32.083 | 0.00041769 | -0.015797 | 32.0988 | 32.0971 | 0.00051837 | -0.0016556
2020-009-0085.mrg | 2020-009-0085.sam | 325 | 750.376 | 750.357 | 0.24504 | -0.019226 | 34.2414 | 34.2284 | 0.00048026 | -0.012958 | 34.2414 | 34.2437 | 0.00035157 | 0.0023422
2020-009-0085.mrg | 2020-009-0085.sam | 241 | 500.63 | 500.63 | 0.074591 | -0.00018311 | 34.12 | 34.1042 | 0.00023353 | -0.01577 | 34.12 | 34.1197 | 0.00023104 | -0.00032806
2020-009-0085.mrg | 2020-009-0085.sam | 241 | 399.978 | 399.978 | 0.20881 | -0.00039673 | 34.0434 | 34.0269 | 0.0003111 | -0.016491 | 34.0434 | 34.0423 | 0.00031916 | -0.0011253
2020-009-0085.mrg | 2020-009-0085.sam | 241 | 4.937 | 4.9374 | 0.1207 | 0.00039005 | 31.1977 | 31.1812 | 0.0012649 | -0.016546 | 31.1977 | 31.1932 | 0.0010302 | -0.004528
2020-009-0096.mrg | 2020-009-0096.sam | 241 | 4.887 | 4.887 | 0.14397 | 2.4796e-05 | 31.5141 | 31.6291 | 0.017306 | 0.11501 | 31.5141 | 31.6379 | 0.010963 | 0.12385
2020-009-0103.mrg | 2020-009-0103.sam | 241 | 4.83 | 4.8299 | 0.23446 | -5.8174e-05 | 31.6254 | 31.609 | 0.00043197 | -0.016439 | 31.6254 | 31.6232 | 0.0004461 | -0.0021973
2020-009-0107.mrg | 2020-009-0107.sam | 154 | 2005.43 | 2005.5 | 0.22518 | 0.074707 | 34.5869 | 34.575 | 0.0002952 | -0.011913 | 34.5869 | 34.5913 | 0.00022572 | 0.0044365
2020-009-0107.mrg | 2020-009-0107.sam | 178 | 1500.7 | 1500.67 | 0.26719 | -0.031006 | 34.5072 | 34.4926 | 0.00028874 | -0.014587 | 34.5072 | 34.5085 | 0.00025337 | 0.0013313
2020-009-0107.mrg | 2020-009-0107.sam | 185 | 1250.31 | 1250.37 | 0.26801 | 0.063232 | 34.4533 | 34.4381 | 0.00030415 | -0.015213 | 34.4533 | 34.4536 | 0.00028762 | 0.00027466
2020-009-0107.mrg | 2020-009-0107.sam | 120 | 1000.15 | 1000.06 | 0.22033 | -0.098511 | 34.371 | 34.3571 | 0.00029141 | -0.013882 | 34.371 | 34.3725 | 0.00031473 | 0.0014572
2020-009-0115.mrg | 2020-009-0115.sam | 175 | 2001.1 | 2001.16 | 0.28262 | 0.053589 | 34.594 | 34.5805 | 0.00025166 | -0.01347 | 34.594 | 34.5969 | 0.00023106 | 0.002903
2020-009-0115.mrg | 2020-009-0115.sam | 482 | 1500.49 | 1500.51 | 0.20503 | 0.017334 | 34.5079 | 34.4931 | 0.00031121 | -0.014763 | 34.5079 | 34.5086 | 0.0005391 | 0.00072479
2020-009-0115.mrg | 2020-009-0115.sam | 149 | 1250.75 | 1250.89 | 0.22177 | 0.13501 | 34.4547 | 34.4405 | 0.00027121 | -0.01424 | 34.4547 | 34.456 | 0.00023702 | 0.0012779
2020-009-0115.mrg | 2020-009-0115.sam | 49 | 1001.16 | 1001.2 | 0.29729 | 0.047669 | 34.381 | 34.3674 | 0.00030601 | -0.013638 | 34.381 | 34.3827 | 0.00032176 | 0.0017204
2020-009-0117.mrg | 2020-009-0117.sam | 217 | 1249.94 | 1249.88 | 0.27282 | -0.060303 | 34.4338 | 34.4185 | 0.00024692 | -0.015251 | 34.4338 | 34.4343 | 0.00023681 | 0.00053024
2020-009-0117.mrg | 2020-009-0117.sam | 206 | 999.529 | 999.458 | 0.28851 | -0.070801 | 34.3609 | 34.3466 | 0.00021032 | -0.014343 | 34.3609 | 34.3622 | 0.00024156 | 0.0012856
2020-009-0117.mrg | 2020-009-0117.sam | 111 | 749.049 | 749.064 | 0.29521 | 0.015259 | 34.2251 | 34.2108 | 0.00022792 | -0.014259 | 34.2251 | 34.2261 | 0.0002192 | 0.0010414
2020-009-0117.mrg | 2020-009-0117.sam | 118 | 500.476 | 500.611 | 0.26694 | 0.13544 | 34.0624 | 34.0494 | 0.00054046 | -0.012966 | 34.0624 | 34.0649 | 0.00040865 | 0.002533
2020-009-0122.mrg | 2020-009-0122.sam | 147 | 87.32 | 87.4554 | 0.31127 | 0.13539 | 33.7046 | 33.6795 | 0.0024446 | -0.025131 | 33.7046 | 33.7063 | 0.00044582 | 0.0016861
2020-009-0133.mrg | 2020-009-0133.sam | 162 | 1999.88 | 1999.81 | 0.25983 | -0.071411 | 34.5875 | 34.5761 | 0.00028433 | -0.01144 | 34.5875 | 34.5917 | 0.00025465 | 0.004158
2020-009-0133.mrg | 2020-009-0133.sam | 127 | 1499.96 | 1499.95 | 0.28348 | -0.0081787 | 34.5068 | 34.4928 | 0.00030234 | -0.014046 | 34.5068 | 34.5077 | 0.00027991 | 0.00092697
2020-009-0133.mrg | 2020-009-0133.sam | 212 | 1249.57 | 1249.5 | 0.23822 | -0.071289 | 34.4529 | 34.4407 | 0.00031627 | -0.012215 | 34.4529 | 34.4555 | 0.00040275 | 0.0026207
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 999.526 | 999.526 | 0.12182 | 0.00042725 | 34.3718 | 34.3588 | 0.0009639 | -0.013008 | 34.3718 | 34.3742 | 0.00052184 | 0.0024262
2020-009-0133.mrg | 2020-009-0133.sam | 215 | 749.842 | 749.773 | 0.24659 | -0.068726 | 34.2649 | 34.2521 | 0.0007455 | -0.012764 | 34.2649 | 34.2675 | 0.00064808 | 0.002552
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 499.713 | 499.713 | 0.12962 | -3.0518e-05 | 34.0551 | 34.0385 | 0.00041172 | -0.016563 | 34.0551 | 34.0527 | 0.00037488 | -0.0023766
2020-009-0133.mrg | 2020-009-0133.sam | 229 | 400.037 | 400.008 | 0.24261 | -0.028503 | 33.9829 | 33.9698 | 0.00027041 | -0.013092 | 33.9829 | 33.9842 | 0.00031366 | 0.0013084
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 300.351 | 300.351 | 0.19489 | -3.0518e-05 | 33.9378 | 33.9249 | 0.00045248 | -0.012928 | 33.9378 | 33.9393 | 0.00059728 | 0.0015297
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 249.649 | 249.649 | 0.23752 | 0.00018311 | 33.9118 | 33.898 | 0.00025133 | -0.01379 | 33.9118 | 33.9123 | 0.00026877 | 0.00054932
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 199.798 | 199.798 | 0.1975 | -0.00016785 | 33.8938 | 33.876 | 0.00028964 | -0.017773 | 33.8938 | 33.8903 | 0.00038065 | -0.0035286
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 175.224 | 175.224 | 0.22069 | -0.00033569 | 33.8461 | 33.8298 | 0.0011428 | -0.016277 | 33.8461 | 33.8392 | 0.0025712 | -0.006916
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 149.667 | 149.667 | 0.17581 | 6.1035e-05 | 33.7358 | 33.7162 | 0.0023905 | -0.019627 | 33.7358 | 33.7287 | 0.0022898 | -0.007103
2020-009-0133.mrg | 2020-009-0133.sam | 234 | 125.086 | 125.07 | 0.27024 | -0.015915 | 33.4781 | 33.4654 | 0.0064108 | -0.012691 | 33.4781 | 33.4802 | 0.0039028 | 0.0020828
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 99.525 | 99.5247 | 0.20023 | -0.0003128 | 33.2965 | 33.2999 | 0.0021689 | 0.003437 | 33.2965 | 33.3147 | 0.0021777 | 0.018169
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 74.787 | 74.7872 | 0.16032 | 0.00021362 | 32.8056 | 32.8025 | 0.0093709 | -0.003067 | 32.8056 | 32.818 | 0.0090986 | 0.012386
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 49.831 | 49.831 | 0.13032 | 1.9073e-05 | 32.3508 | 32.3668 | 0.0036721 | 0.015995 | 32.3508 | 32.3868 | 0.0025859 | 0.035965
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 29.525 | 29.5248 | 0.16313 | -0.00019455 | 31.93 | 31.9172 | 0.015671 | -0.012842 | 31.93 | 31.9238 | 0.019722 | -0.006218
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 19.888 | 19.8884 | 0.13465 | 0.00040245 | 31.8571 | 31.8441 | 0.0032832 | -0.013012 | 31.8571 | 31.8557 | 0.0018914 | -0.0014057
2020-009-0133.mrg | 2020-009-0133.sam | 241 | 9.964 | 9.9639 | 0.12138 | -0.00014877 | 31.843 | 31.8282 | 0.0024508 | -0.014761 | 31.843 | 31.84 | 0.0015075 | -0.0029869
2020-009-0137.mrg | 2020-009-0137.sam | 131 | 1500.5 | 1500.39 | 0.30907 | -0.11255 | 34.5023 | 34.4887 | 0.00029526 | -0.013603 | 34.5023 | 34.5034 | 0.00028854 | 0.0011139
2020-009-0137.mrg | 2020-009-0137.sam | 238 | 1249.8 | 1249.8 | 0.32316 | 0.0021973 | 34.4483 | 34.4348 | 0.00023086 | -0.013493 | 34.4483 | 34.4496 | 0.00023978 | 0.0012512
2020-009-0137.mrg | 2020-009-0137.sam | 241 | 1000.06 | 1000.06 | 0.29643 | 0.00030518 | 34.3718 | 34.3598 | 0.00028416 | -0.012005 | 34.3718 | 34.3739 | 0.00037698 | 0.0020905
2020-009-0137.mrg | 2020-009-0137.sam | 241 | 749.935 | 749.935 | 0.21751 | -0.00024414 | 34.2631 | 34.2482 | 0.00045628 | -0.014934 | 34.2631 | 34.2624 | 0.00048882 | -0.00069809
2020-009-0175.mrg | 2020-009-0175.sam | 226 | 5.192 | 5.2128 | 0.25297 | 0.020823 | 31.647 | 31.6303 | 0.00025787 | -0.016724 | 31.647 | 31.644 | 0.00026465 | -0.003006
2020-009-0177.mrg | 2020-009-0177.sam | 123 | 2004.6 | 2004.64 | 0.30913 | 0.036865 | 34.5938 | 34.5811 | 0.00030629 | -0.012653 | 34.5938 | 34.5971 | 0.00026745 | 0.0033226
2020-009-0177.mrg | 2020-009-0177.sam | 118 | 1751.26 | 1751.2 | 0.31011 | -0.059204 | 34.5472 | 34.5337 | 0.00028904 | -0.013534 | 34.5472 | 34.5492 | 0.00024765 | 0.0020409
2020-009-0177.mrg | 2020-009-0177.sam | 55 | 1501.27 | 1501.27 | 0.29112 | 0.0079346 | 34.4907 | 34.4771 | 0.00039666 | -0.013649 | 34.4907 | 34.4925 | 0.00026386 | 0.0018387
2020-009-0177.mrg | 2020-009-0177.sam | 103 | 1251.96 | 1252.1 | 0.19463 | 0.13916 | 34.4375 | 34.4241 | 0.00027371 | -0.013355 | 34.4375 | 34.4395 | 0.00024237 | 0.0020142
2020-009-0177.mrg | 2020-009-0177.sam | 61 | 1001.16 | 1001.21 | 0.30695 | 0.057312 | 34.3667 | 34.354 | 0.00047798 | -0.012722 | 34.3667 | 34.3687 | 0.00054629 | 0.002018
2020-009-0188.mrg | 2020-009-0188.sam | 217 | 4.788 | 4.8468 | 0.26898 | 0.058829 | 31.5587 | 31.5399 | 0.0021193 | -0.018846 | 31.5587 | 31.5551 | 0.0010405 | -0.003582
2020-009-0190.mrg | 2020-009-0190.sam | 241 | 50.292 | 50.2916 | 0.11603 | -0.00041199 | 32.3454 | 32.3349 | 0.0010036 | -0.01054 | 32.3454 | 32.3503 | 0.0011064 | 0.0048752
2020-009-0190.mrg | 2020-009-0190.sam | 224 | 40.608 | 40.6548 | 0.25854 | 0.046829 | 32.0949 | 32.0805 | 0.0013655 | -0.014355 | 32.0949 | 32.0962 | 0.0024558 | 0.0013008
2020-009-0190.mrg | 2020-009-0190.sam | 241 | 30.537 | 30.5373 | 0.21935 | 0.00033188 | 31.6256 | 31.5733 | 0.01399 | -0.052298 | 31.6256 | 31.587 | 0.010877 | -0.038578
2020-009-0197.mrg | 2020-009-0197.sam | 225 | 299.324 | 299.376 | 0.14064 | 0.05191 | 31.8387 | 31.8252 | 0.00020299 | -0.013512 | 31.8387 | 31.839 | 0.00025303 | 0.00029182
2020-009-0201.mrg | 2020-009-0201.sam | 241 | 5.346 | 5.3457 | 0.037757 | -0.00028658 | 28.2151 | 28.3089 | 0.015173 | 0.093794 | 28.2151 | 28.3128 | 0.0072689 | 0.097715
2020-009-0204.mrg | 2020-009-0204.sam | 241 | 300.061 | 300.061 | 0.032189 | 0.00015259 | 30.878 | 30.861 | 0.00023845 | -0.01697 | 30.878 | 30.873 | 0.00021175 | -0.0050354
2020-009-0204.mrg | 2020-009-0204.sam | 241 | 5.209 | 5.2088 | 0.033069 | -0.00024891 | 27.513 | 27.8348 | 0.09733 | 0.32181 | 27.513 | 27.8327 | 0.068316 | 0.31974
2020-009-0205.mrg | 2020-009-0205.sam | 241 | 5.115 | 5.1147 | 0.03179 | -0.00034857 | 27.3817 | 26.7518 | 0.010438 | -0.62987 | 27.3817 | 26.7849 | 0.017792 | -0.59684
2020-009-0212.mrg | 2020-009-0212.sam | 241 | 5.023 | 5.0235 | 0.026997 | 0.00048161 | 27.5558 | 26.9474 | 0.071326 | -0.60843 | 27.5558 | 27.0308 | 0.089804 | -0.525
2020-009-0222.mrg | 2020-009-0222.sam | 241 | 300.254 | 300.254 | 0.062645 | 0.00030518 | 31.1783 | 31.1653 | 0.0002006 | -0.013031 | 31.1783 | 31.1771 | 0.00019735 | -0.0012245
2020-009-0222.mrg | 2020-009-0222.sam | 241 | 5.22 | 5.2197 | 0.030836 | -0.00026131 | 25.761 | 25.1824 | 0.033492 | -0.57864 | 25.761 | 25.3094 | 0.080208 | -0.45157
2020-009-0231.mrg | 2020-009-0231.sam | 241 | 299.944 | 299.944 | 0.0306 | -0.00024414 | 31.1905 | 31.1791 | 0.00020559 | -0.011435 | 31.1905 | 31.1907 | 0.00026426 | 0.00015068
2020-009-0231.mrg | 2020-009-0231.sam | 241 | 4.946 | 4.9459 | 0.030449 | -6.628e-05 | 26.9975 | 26.794 | 0.003763 | -0.2035 | 26.9975 | 26.8052 | 0.002619 | -0.19226
2020-009-0234.mrg | 2020-009-0234.sam | 241 | 5.165 | 5.1648 | 0.0287 | -0.00021172 | 29.5655 | 29.5887 | 0.00055474 | 0.023189 | 29.5655 | 29.602 | 0.0027457 | 0.036495
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 33
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2020-009-0022.mrg | 2020-009-0022.sam | 12 | 12 | 2020-009-0022.mrg & 2020-009-0022.sam
2 | 2020-009-0023.mrg | 2020-009-0023.sam | 14 | 14 | 2020-009-0023.mrg & 2020-009-0023.sam
3 | 2020-009-0027.mrg | 2020-009-0027.sam | 16 | 16 | 2020-009-0027.mrg & 2020-009-0027.sam
4 | 2020-009-0029.mrg | 2020-009-0029.sam | 18 | 22 | 2020-009-0029.mrg & 2020-009-0029.sam
5 | 2020-009-0033.mrg | 2020-009-0033.sam | 24 | 27 | 2020-009-0033.mrg & 2020-009-0033.sam
6 | 2020-009-0035.mrg | 2020-009-0035.sam | 29 | 31 | 2020-009-0035.mrg & 2020-009-0035.sam
7 | 2020-009-0039.mrg | 2020-009-0039.sam | 33 | 38 | 2020-009-0039.mrg & 2020-009-0039.sam
8 | 2020-009-0043.mrg | 2020-009-0043.sam | 40 | 41 | 2020-009-0043.mrg & 2020-009-0043.sam
9 | 2020-009-0057.mrg | 2020-009-0057.sam | 43 | 43 | 2020-009-0057.mrg & 2020-009-0057.sam
10 | 2020-009-0069.mrg | 2020-009-0069.sam | 45 | 45 | 2020-009-0069.mrg & 2020-009-0069.sam
11 | 2020-009-0077.mrg | 2020-009-0077.sam | 47 | 50 | 2020-009-0077.mrg & 2020-009-0077.sam
12 | 2020-009-0078.mrg | 2020-009-0078.sam | 52 | 56 | 2020-009-0078.mrg & 2020-009-0078.sam
13 | 2020-009-0085.mrg | 2020-009-0085.sam | 58 | 61 | 2020-009-0085.mrg & 2020-009-0085.sam
14 | 2020-009-0096.mrg | 2020-009-0096.sam | 63 | 63 | 2020-009-0096.mrg & 2020-009-0096.sam
15 | 2020-009-0103.mrg | 2020-009-0103.sam | 65 | 65 | 2020-009-0103.mrg & 2020-009-0103.sam
16 | 2020-009-0107.mrg | 2020-009-0107.sam | 67 | 70 | 2020-009-0107.mrg & 2020-009-0107.sam
17 | 2020-009-0115.mrg | 2020-009-0115.sam | 72 | 75 | 2020-009-0115.mrg & 2020-009-0115.sam
18 | 2020-009-0117.mrg | 2020-009-0117.sam | 77 | 80 | 2020-009-0117.mrg & 2020-009-0117.sam
19 | 2020-009-0122.mrg | 2020-009-0122.sam | 82 | 82 | 2020-009-0122.mrg & 2020-009-0122.sam
20 | 2020-009-0133.mrg | 2020-009-0133.sam | 84 | 102 | 2020-009-0133.mrg & 2020-009-0133.sam
21 | 2020-009-0137.mrg | 2020-009-0137.sam | 104 | 107 | 2020-009-0137.mrg & 2020-009-0137.sam
22 | 2020-009-0175.mrg | 2020-009-0175.sam | 109 | 109 | 2020-009-0175.mrg & 2020-009-0175.sam
23 | 2020-009-0177.mrg | 2020-009-0177.sam | 111 | 115 | 2020-009-0177.mrg & 2020-009-0177.sam
24 | 2020-009-0188.mrg | 2020-009-0188.sam | 117 | 117 | 2020-009-0188.mrg & 2020-009-0188.sam
25 | 2020-009-0190.mrg | 2020-009-0190.sam | 119 | 121 | 2020-009-0190.mrg & 2020-009-0190.sam
26 | 2020-009-0197.mrg | 2020-009-0197.sam | 123 | 123 | 2020-009-0197.mrg & 2020-009-0197.sam
27 | 2020-009-0201.mrg | 2020-009-0201.sam | 125 | 125 | 2020-009-0201.mrg & 2020-009-0201.sam
28 | 2020-009-0204.mrg | 2020-009-0204.sam | 127 | 128 | 2020-009-0204.mrg & 2020-009-0204.sam
29 | 2020-009-0205.mrg | 2020-009-0205.sam | 130 | 130 | 2020-009-0205.mrg & 2020-009-0205.sam
30 | 2020-009-0212.mrg | 2020-009-0212.sam | 132 | 132 | 2020-009-0212.mrg & 2020-009-0212.sam
31 | 2020-009-0222.mrg | 2020-009-0222.sam | 134 | 135 | 2020-009-0222.mrg & 2020-009-0222.sam
32 | 2020-009-0231.mrg | 2020-009-0231.sam | 137 | 138 | 2020-009-0231.mrg & 2020-009-0231.sam
33 | 2020-009-0234.mrg | 2020-009-0234.sam | 140 | 140 | 2020-009-0234.mrg & 2020-009-0234.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 33
2 | (Sal1 - Sal_Bot) | Pressure | 33
3 | (Sal0 - Sal_Bot) | File Pair # | 33
4 | (Sal1 - Sal_Bot) | File Pair # | 33

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.070846 |  | X | CALC Y
m= | 6.1943e-07 | 2.9205e-07 |  | Std Dev of Resid |  | 0.0011891 |  | 249.649 | -0.014392
b= | -0.014546 | 0.00033266 |  | Resid Sum of Squares |  | 8.3429e-05 |  | 2005.5 | -0.013304
# of Values Extracted = |  |  | 97
# of Values Used in Fit = |  |  | 61
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.5033 | 31.8902 | 0.0064545 | 0.0077324 | -0.014543 | -0.022275 |  | 400.684 | -0.014763
2 | NO | 5.4007 | 31.2491 | 0.013314 | -0.01536 | -0.014543 | 0.00081695 |  | 999.769 | -0.014614
3 | YES | 400.684 | 34.0217 | 0.00025242 | -0.014763 | -0.014298 | 0.0004648 |  | 749.78 | -0.014465
4 | YES | 999.769 | 34.3647 | 0.00031769 | -0.014614 | -0.013927 | 0.00068709 |  | 499.917 | -0.011948
4 | YES | 749.78 | 34.2324 | 0.00068821 | -0.014465 | -0.014082 | 0.00038344 |  | 1600.79 | -0.012897
4 | YES | 499.917 | 34.0443 | 0.00059345 | -0.011948 | -0.014237 | -0.002289 |  | 1499.99 | -0.015114
4 | NO | 399.818 | 34.0009 | 0.0015504 | -0.042351 | -0.014299 | 0.028052 |  | 1000.35 | -0.013493
4 | NO | 5.003 | 31.8691 | 0.00023901 | -0.015137 | -0.014543 | 0.00059351 |  | 2004.87 | -0.013332
5 | YES | 1600.79 | 34.5195 | 0.0003023 | -0.012897 | -0.013555 | -0.00065722 |  | 1499.69 | -0.01403
5 | YES | 1499.99 | 34.5022 | 0.00025571 | -0.015114 | -0.013617 | 0.0014966 |  | 1000.7 | -0.013901
5 | YES | 1000.35 | 34.3694 | 0.00026899 | -0.013493 | -0.013927 | -0.00043405 |  | 1250.42 | -0.01384
5 | NO | 5.0381 | 31.8884 | 0.00053255 | -0.017769 | -0.014543 | 0.0032257 |  | 1000.45 | -0.01458
6 | YES | 2004.87 | 34.5799 | 0.00026907 | -0.013332 | -0.013304 | 2.7981e-05 |  | 750.215 | -0.016006
6 | YES | 1499.69 | 34.5022 | 0.00026082 | -0.01403 | -0.013617 | 0.00041316 |  | 299.815 | -0.014969
6 | YES | 1000.7 | 34.363 | 0.00023335 | -0.013901 | -0.013926 | -2.563e-05 |  | 500.871 | -0.013901
7 | YES | 1250.42 | 34.4376 | 0.0002446 | -0.01384 | -0.013772 | 6.7954e-05 |  | 400.516 | -0.015083
7 | YES | 1000.45 | 34.3714 | 0.00050775 | -0.01458 | -0.013927 | 0.00065322 |  | 500.44 | -0.014561
7 | YES | 750.215 | 34.2573 | 0.00024963 | -0.016006 | -0.014082 | 0.0019249 |  | 1000.15 | -0.013794
7 | YES | 299.815 | 33.9196 | 0.00026635 | -0.014969 | -0.014361 | 0.00060832 |  | 749.87 | -0.013508
7 | NO | 200.472 | 33.8798 | 0.0032208 | -0.015739 | -0.014422 | 0.0013173 |  | 499.977 | -0.014938
7 | NO | 4.9019 | 31.9393 | 0.0002229 | -0.016613 | -0.014543 | 0.0020697 |  | 1999.98 | -0.013817
8 | YES | 500.871 | 34.0732 | 0.00026028 | -0.013901 | -0.014236 | -0.00033524 |  | 1500.14 | -0.014572
8 | YES | 400.516 | 34.0269 | 0.00024055 | -0.015083 | -0.014298 | 0.0007851 |  | 1249.94 | -0.012516
9 | NO | 4.84 | 30.2488 | 0.0048652 | -0.019524 | -0.014543 | 0.0049803 |  | 1000.25 | -0.015926
10 | YES | 500.44 | 34.0507 | 0.00041012 | -0.014561 | -0.014236 | 0.00032439 |  | 750.357 | -0.012958
11 | NO | 1249.9 | 34.4502 | 0.0014198 | -0.01535 | -0.013772 | 0.0015782 |  | 500.63 | -0.01577
11 | YES | 1000.15 | 34.3722 | 0.00021775 | -0.013794 | -0.013927 | -0.00013287 |  | 399.978 | -0.016491
11 | YES | 749.87 | 34.2412 | 0.00087882 | -0.013508 | -0.014082 | -0.000574 |  | 2005.5 | -0.011913
11 | YES | 499.977 | 34.0784 | 0.00028814 | -0.014938 | -0.014237 | 0.00070181 |  | 1500.67 | -0.014587
12 | YES | 1999.98 | 34.5829 | 0.00029701 | -0.013817 | -0.013307 | 0.00050935 |  | 1250.37 | -0.015213
12 | YES | 1500.14 | 34.4927 | 0.0002911 | -0.014572 | -0.013617 | 0.00095504 |  | 1000.06 | -0.013882
12 | YES | 1249.94 | 34.4495 | 0.00025847 | -0.012516 | -0.013772 | -0.001256 |  | 2001.16 | -0.01347
12 | YES | 1000.25 | 34.3696 | 0.00025999 | -0.015926 | -0.013927 | 0.0019997 |  | 1500.51 | -0.014763
12 | NO | 4.4433 | 32.083 | 0.00041769 | -0.015797 | -0.014544 | 0.0012532 |  | 1250.89 | -0.01424
13 | YES | 750.357 | 34.2284 | 0.00048026 | -0.012958 | -0.014081 | -0.001123 |  | 1001.2 | -0.013638
13 | YES | 500.63 | 34.1042 | 0.00023353 | -0.01577 | -0.014236 | 0.0015338 |  | 1249.88 | -0.015251
13 | YES | 399.978 | 34.0269 | 0.0003111 | -0.016491 | -0.014299 | 0.0021924 |  | 999.458 | -0.014343
13 | NO | 4.9374 | 31.1812 | 0.0012649 | -0.016546 | -0.014543 | 0.002003 |  | 749.064 | -0.014259
14 | NO | 4.887 | 31.6291 | 0.017306 | 0.11501 | -0.014543 | -0.12955 |  | 500.611 | -0.012966
15 | NO | 4.8299 | 31.609 | 0.00043197 | -0.016439 | -0.014543 | 0.0018961 |  | 1999.81 | -0.01144
16 | YES | 2005.5 | 34.575 | 0.0002952 | -0.011913 | -0.013304 | -0.0013907 |  | 1499.95 | -0.014046
16 | YES | 1500.67 | 34.4926 | 0.00028874 | -0.014587 | -0.013617 | 0.00097067 |  | 1249.5 | -0.012215
16 | YES | 1250.37 | 34.4381 | 0.00030415 | -0.015213 | -0.013772 | 0.0014412 |  | 999.526 | -0.013008
16 | YES | 1000.06 | 34.3571 | 0.00029141 | -0.013882 | -0.013927 | -4.5126e-05 |  | 749.773 | -0.012764
17 | YES | 2001.16 | 34.5805 | 0.00025166 | -0.01347 | -0.013307 | 0.00016298 |  | 499.713 | -0.016563
17 | YES | 1500.51 | 34.4931 | 0.00031121 | -0.014763 | -0.013617 | 0.0011461 |  | 400.008 | -0.013092
17 | YES | 1250.89 | 34.4405 | 0.00027121 | -0.01424 | -0.013771 | 0.00046884 |  | 300.351 | -0.012928
17 | YES | 1001.2 | 34.3674 | 0.00030601 | -0.013638 | -0.013926 | -0.00028862 |  | 249.649 | -0.01379
18 | YES | 1249.88 | 34.4185 | 0.00024692 | -0.015251 | -0.013772 | 0.0014791 |  | 1500.39 | -0.013603
18 | YES | 999.458 | 34.3466 | 0.00021032 | -0.014343 | -0.013927 | 0.0004161 |  | 1249.8 | -0.013493
18 | YES | 749.064 | 34.2108 | 0.00022792 | -0.014259 | -0.014082 | 0.000177 |  | 1000.06 | -0.012005
18 | YES | 500.611 | 34.0494 | 0.00054046 | -0.012966 | -0.014236 | -0.00127 |  | 749.935 | -0.014934
19 | NO | 87.4554 | 33.6795 | 0.0024446 | -0.025131 | -0.014492 | 0.010639 |  | 2004.64 | -0.012653
20 | YES | 1999.81 | 34.5761 | 0.00028433 | -0.01144 | -0.013308 | -0.0018673 |  | 1751.2 | -0.013534
20 | YES | 1499.95 | 34.4928 | 0.00030234 | -0.014046 | -0.013617 | 0.00042852 |  | 1501.27 | -0.013649
20 | YES | 1249.5 | 34.4407 | 0.00031627 | -0.012215 | -0.013772 | -0.0015576 |  | 1252.1 | -0.013355
20 | YES | 999.526 | 34.3588 | 0.0009639 | -0.013008 | -0.013927 | -0.00091906 |  | 1001.21 | -0.012722
20 | YES | 749.773 | 34.2521 | 0.0007455 | -0.012764 | -0.014082 | -0.0013179 |  | 299.376 | -0.013512
20 | YES | 499.713 | 34.0385 | 0.00041172 | -0.016563 | -0.014237 | 0.0023266 |  | 300.061 | -0.01697
20 | YES | 400.008 | 33.9698 | 0.00027041 | -0.013092 | -0.014299 | -0.0012065 |  | 300.254 | -0.013031
20 | YES | 300.351 | 33.9249 | 0.00045248 | -0.012928 | -0.01436 | -0.0014322 |  | 299.944 | -0.011435 | -0.013919 | 0.0012233
20 | YES | 249.649 | 33.898 | 0.00025133 | -0.01379 | -0.014392 | -0.00060155
20 | NO | 199.798 | 33.876 | 0.00028964 | -0.017773 | -0.014423 | 0.0033502
20 | NO | 175.224 | 33.8298 | 0.0011428 | -0.016277 | -0.014438 | 0.0018395
20 | NO | 149.667 | 33.7162 | 0.0023905 | -0.019627 | -0.014454 | 0.005173
20 | NO | 125.07 | 33.4654 | 0.0064108 | -0.012691 | -0.014469 | -0.0017773
20 | NO | 99.5247 | 33.2999 | 0.0021689 | 0.003437 | -0.014485 | -0.017922
20 | NO | 74.7872 | 32.8025 | 0.0093709 | -0.003067 | -0.0145 | -0.011433
20 | NO | 49.831 | 32.3668 | 0.0036721 | 0.015995 | -0.014515 | -0.03051
20 | NO | 29.5248 | 31.9172 | 0.015671 | -0.012842 | -0.014528 | -0.0016858
20 | NO | 19.8884 | 31.8441 | 0.0032832 | -0.013012 | -0.014534 | -0.0015221
20 | NO | 9.9639 | 31.8282 | 0.0024508 | -0.014761 | -0.01454 | 0.00022088
21 | YES | 1500.39 | 34.4887 | 0.00029526 | -0.013603 | -0.013617 | -1.3708e-05
21 | YES | 1249.8 | 34.4348 | 0.00023086 | -0.013493 | -0.013772 | -0.00027953
21 | YES | 1000.06 | 34.3598 | 0.00028416 | -0.012005 | -0.013927 | -0.0019219
21 | YES | 749.935 | 34.2482 | 0.00045628 | -0.014934 | -0.014082 | 0.00085274
22 | NO | 5.2128 | 31.6303 | 0.00025787 | -0.016724 | -0.014543 | 0.0021805
23 | YES | 2004.64 | 34.5811 | 0.00030629 | -0.012653 | -0.013305 | -0.00065116
23 | YES | 1751.2 | 34.5337 | 0.00028904 | -0.013534 | -0.013462 | 7.2951e-05
23 | YES | 1501.27 | 34.4771 | 0.00039666 | -0.013649 | -0.013616 | 3.2637e-05
23 | YES | 1252.1 | 34.4241 | 0.00027371 | -0.013355 | -0.013771 | -0.00041541
23 | YES | 1001.21 | 34.354 | 0.00047798 | -0.012722 | -0.013926 | -0.0012041
24 | NO | 4.8468 | 31.5399 | 0.0021193 | -0.018846 | -0.014543 | 0.0043032
25 | NO | 50.2916 | 32.3349 | 0.0010036 | -0.01054 | -0.014515 | -0.0039751
25 | NO | 40.6548 | 32.0805 | 0.0013655 | -0.014355 | -0.014521 | -0.00016641
25 | NO | 30.5373 | 31.5733 | 0.01399 | -0.052298 | -0.014527 | 0.03777
26 | YES | 299.376 | 31.8252 | 0.00020299 | -0.013512 | -0.014361 | -0.00084915
27 | NO | 5.3457 | 28.3089 | 0.015173 | 0.093794 | -0.014543 | -0.10834
28 | YES | 300.061 | 30.861 | 0.00023845 | -0.01697 | -0.01436 | 0.0026093
28 | NO | 5.2088 | 27.8348 | 0.09733 | 0.32181 | -0.014543 | -0.33635
29 | NO | 5.1147 | 26.7518 | 0.010438 | -0.62987 | -0.014543 | 0.61532
30 | NO | 5.0235 | 26.9474 | 0.071326 | -0.60843 | -0.014543 | 0.59388
31 | YES | 300.254 | 31.1653 | 0.0002006 | -0.013031 | -0.01436 | -0.0013293
31 | NO | 5.2197 | 25.1824 | 0.033492 | -0.57864 | -0.014543 | 0.5641
32 | YES | 299.944 | 31.1791 | 0.00020559 | -0.011435 | -0.01436 | -0.0029259
32 | NO | 4.9459 | 26.794 | 0.003763 | -0.2035 | -0.014543 | 0.18896
33 | NO | 5.1648 | 29.5887 | 0.00055474 | 0.023189 | -0.014543 | -0.037733
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
NO
NO
NO
YES
NO
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.3373 |  | X | CALC Y
m= | 1.693e-06 | 3.0385e-07 |  | Std Dev of Resid |  | 0.0012634 |  | 200.472 | 1.3628e-05
b= | -0.00032578 | 0.000344 |  | Resid Sum of Squares |  | 9.7373e-05 |  | 2005.5 | 0.0030696
# of Values Extracted = |  |  | 97
# of Values Used in Fit = |  |  | 63
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.5033 | 31.8993 | 0.006048 | 0.016819 | -0.00031646 | -0.017135 |  | 400.684 | 0.00098038
2 | NO | 5.4007 | 31.2652 | 0.020152 | 0.00069237 | -0.00031664 | -0.001009 |  | 999.769 | 0.00089264
3 | YES | 400.684 | 34.0375 | 0.00024373 | 0.00098038 | 0.0003526 | -0.00062778 |  | 749.78 | 0.00085449
4 | YES | 999.769 | 34.3802 | 0.00032798 | 0.00089264 | 0.0013669 | 0.00047423 |  | 499.917 | 0.0034943
4 | YES | 749.78 | 34.2478 | 0.00078512 | 0.00085449 | 0.00094363 | 8.9138e-05 |  | 1600.79 | 0.0036469
4 | YES | 499.917 | 34.0597 | 0.00074487 | 0.0034943 | 0.0005206 | -0.0029737 |  | 1499.99 | 0.0013466
4 | NO | 399.818 | 34.0174 | 0.001251 | -0.025841 | 0.00035113 | 0.026192 |  | 1000.35 | 0.0024338
4 | NO | 5.003 | 31.8831 | 0.00019925 | -0.00107 | -0.00031731 | 0.00075271 |  | 2004.87 | 0.0034866
5 | YES | 1600.79 | 34.536 | 0.00026636 | 0.0036469 | 0.0023844 | -0.0012624 |  | 1499.69 | 0.002243
5 | YES | 1499.99 | 34.5186 | 0.00027346 | 0.0013466 | 0.0022138 | 0.00086718 |  | 1000.7 | 0.0020256
5 | YES | 1000.35 | 34.3853 | 0.000337 | 0.0024338 | 0.0013679 | -0.0010659 |  | 1250.42 | 0.0024834
5 | NO | 5.0381 | 31.9029 | 0.00077753 | -0.0032654 | -0.00031725 | 0.0029481 |  | 1000.45 | 0.0016213
6 | YES | 2004.87 | 34.5967 | 0.00024344 | 0.0034866 | 0.0030685 | -0.00041808 |  | 750.215 | -0.00013351
6 | YES | 1499.69 | 34.5184 | 0.00025723 | 0.002243 | 0.0022133 | -2.9782e-05 |  | 299.815 | 0.00079346
6 | YES | 1000.7 | 34.3789 | 0.00023092 | 0.0020256 | 0.0013684 | -0.00065715 |  | 200.472 | 0.0004425
7 | YES | 1250.42 | 34.4539 | 0.00024976 | 0.0024834 | 0.0017912 | -0.00069214 |  | 500.871 | 0.0016403
7 | YES | 1000.45 | 34.3876 | 0.00037061 | 0.0016213 | 0.001368 | -0.00025323 |  | 400.516 | 0.00054169
7 | YES | 750.215 | 34.2732 | 0.0002485 | -0.00013351 | 0.00094437 | 0.0010779 |  | 500.44 | 0.0014114
7 | YES | 299.815 | 33.9354 | 0.00029319 | 0.00079346 | 0.00018182 | -0.00061164 |  | 1249.9 | 0.0010757
7 | YES | 200.472 | 33.8959 | 0.00031884 | 0.0004425 | 1.3628e-05 | -0.00042888 |  | 1000.15 | 0.0017891
7 | NO | 4.9019 | 31.954 | 0.00025289 | -0.0019417 | -0.00031748 | 0.0016242 |  | 749.87 | 0.0019455
8 | YES | 500.871 | 34.0887 | 0.00029972 | 0.0016403 | 0.00052222 | -0.0011181 |  | 499.977 | 0.00035858
8 | YES | 400.516 | 34.0425 | 0.00022015 | 0.00054169 | 0.00035231 | -0.00018938 |  | 1999.98 | 0.0028076
9 | NO | 4.84 | 30.2616 | 0.0043393 | -0.0066834 | -0.00031758 | 0.0063658 |  | 1500.14 | 0.0016174
10 | YES | 500.44 | 34.0667 | 0.00074509 | 0.0014114 | 0.00052149 | -0.00088995 |  | 1249.94 | 0.0036163
11 | YES | 1249.9 | 34.4667 | 0.00066694 | 0.0010757 | 0.0017904 | 0.00071461 |  | 1000.25 | -0.00016403
11 | YES | 1000.15 | 34.3878 | 0.00025064 | 0.0017891 | 0.0013675 | -0.00042157 |  | 750.357 | 0.0023422
11 | YES | 749.87 | 34.2566 | 0.00062783 | 0.0019455 | 0.00094378 | -0.0010017 |  | 500.63 | -0.00032806
11 | YES | 499.977 | 34.0937 | 0.00024801 | 0.00035858 | 0.0005207 | 0.00016212 |  | 399.978 | -0.0011253
12 | YES | 1999.98 | 34.5995 | 0.00026104 | 0.0028076 | 0.0030603 | 0.00025265 |  | 2005.5 | 0.0044365
12 | YES | 1500.14 | 34.5089 | 0.00029375 | 0.0016174 | 0.002214 | 0.00059659 |  | 1500.67 | 0.0013313
12 | YES | 1249.94 | 34.4656 | 0.00023307 | 0.0036163 | 0.0017904 | -0.0018259 |  | 1250.37 | 0.00027466
12 | YES | 1000.25 | 34.3853 | 0.00024216 | -0.00016403 | 0.0013677 | 0.0015317 |  | 1000.06 | 0.0014572
12 | NO | 4.4433 | 32.0971 | 0.00051837 | -0.0016556 | -0.00031826 | 0.0013373 |  | 2001.16 | 0.002903
13 | YES | 750.357 | 34.2437 | 0.00035157 | 0.0023422 | 0.00094461 | -0.0013976 |  | 1500.51 | 0.00072479
13 | YES | 500.63 | 34.1197 | 0.00023104 | -0.00032806 | 0.00052181 | 0.00084987 |  | 1250.89 | 0.0012779
13 | YES | 399.978 | 34.0423 | 0.00031916 | -0.0011253 | 0.0003514 | 0.0014767 |  | 1001.2 | 0.0017204
13 | NO | 4.9374 | 31.1932 | 0.0010302 | -0.004528 | -0.00031742 | 0.0042106 |  | 1249.88 | 0.00053024
14 | NO | 4.887 | 31.6379 | 0.010963 | 0.12385 | -0.00031751 | -0.12417 |  | 999.458 | 0.0012856
15 | NO | 4.8299 | 31.6232 | 0.0004461 | -0.0021973 | -0.0003176 | 0.0018797 |  | 749.064 | 0.0010414
16 | YES | 2005.5 | 34.5913 | 0.00022572 | 0.0044365 | 0.0030696 | -0.0013669 |  | 500.611 | 0.002533
16 | YES | 1500.67 | 34.5085 | 0.00025337 | 0.0013313 | 0.0022149 | 0.00088359 |  | 1999.81 | 0.004158
16 | YES | 1250.37 | 34.4536 | 0.00028762 | 0.00027466 | 0.0017911 | 0.0015165 |  | 1499.95 | 0.00092697
16 | YES | 1000.06 | 34.3725 | 0.00031473 | 0.0014572 | 0.0013674 | -8.9846e-05 |  | 1249.5 | 0.0026207
17 | YES | 2001.16 | 34.5969 | 0.00023106 | 0.002903 | 0.0030623 | 0.00015929 |  | 999.526 | 0.0024262
17 | YES | 1500.51 | 34.5086 | 0.0005391 | 0.00072479 | 0.0022146 | 0.0014899 |  | 749.773 | 0.002552
17 | YES | 1250.89 | 34.456 | 0.00023702 | 0.0012779 | 0.001792 | 0.00051411 |  | 499.713 | -0.0023766
17 | YES | 1001.2 | 34.3827 | 0.00032176 | 0.0017204 | 0.0013693 | -0.00035114 |  | 400.008 | 0.0013084
18 | YES | 1249.88 | 34.4343 | 0.00023681 | 0.00053024 | 0.0017903 | 0.0012601 |  | 300.351 | 0.0015297
18 | YES | 999.458 | 34.3622 | 0.00024156 | 0.0012856 | 0.0013663 | 8.0795e-05 |  | 249.649 | 0.00054932
18 | YES | 749.064 | 34.2261 | 0.0002192 | 0.0010414 | 0.00094242 | -9.8993e-05 |  | 1500.39 | 0.0011139
18 | YES | 500.611 | 34.0649 | 0.00040865 | 0.002533 | 0.00052178 | -0.0020112 |  | 1249.8 | 0.0012512
19 | NO | 87.4554 | 33.7063 | 0.00044582 | 0.0016861 | -0.00017771 | -0.0018638 |  | 1000.06 | 0.0020905
20 | YES | 1999.81 | 34.5917 | 0.00025465 | 0.004158 | 0.00306 | -0.001098 |  | 749.935 | -0.00069809
20 | YES | 1499.95 | 34.5077 | 0.00027991 | 0.00092697 | 0.0022137 | 0.0012867 |  | 2004.64 | 0.0033226
20 | YES | 1249.5 | 34.4555 | 0.00040275 | 0.0026207 | 0.0017897 | -0.00083102 |  | 1751.2 | 0.0020409
20 | YES | 999.526 | 34.3742 | 0.00052184 | 0.0024262 | 0.0013665 | -0.0010597 |  | 1501.27 | 0.0018387
20 | YES | 749.773 | 34.2675 | 0.00064808 | 0.002552 | 0.00094362 | -0.0016084 |  | 1252.1 | 0.0020142
20 | YES | 499.713 | 34.0527 | 0.00037488 | -0.0023766 | 0.00052026 | 0.0028968 |  | 1001.21 | 0.002018
20 | YES | 400.008 | 33.9842 | 0.00031366 | 0.0013084 | 0.00035145 | -0.00095699 |  | 299.376 | 0.00029182
20 | YES | 300.351 | 33.9393 | 0.00059728 | 0.0015297 | 0.00018273 | -0.001347 |  | 300.061 | -0.0050354
20 | YES | 249.649 | 33.9123 | 0.00026877 | 0.00054932 | 9.6887e-05 | -0.00045243 |  | 300.254 | -0.0012245
20 | NO | 199.798 | 33.8903 | 0.00038065 | -0.0035286 | 1.2487e-05 | 0.0035411 |  | 299.944 | 0.00015068 | 0.0013734 | 0.0015394
20 | NO | 175.224 | 33.8392 | 0.0025712 | -0.006916 | -2.9117e-05 | 0.0068869
20 | NO | 149.667 | 33.7287 | 0.0022898 | -0.007103 | -7.2387e-05 | 0.0070306
20 | NO | 125.07 | 33.4802 | 0.0039028 | 0.0020828 | -0.00011403 | -0.0021969
20 | NO | 99.5247 | 33.3147 | 0.0021777 | 0.018169 | -0.00015728 | -0.018327
20 | NO | 74.7872 | 32.818 | 0.0090986 | 0.012386 | -0.00019916 | -0.012585
20 | NO | 49.831 | 32.3868 | 0.0025859 | 0.035965 | -0.00024141 | -0.036206
20 | NO | 29.5248 | 31.9238 | 0.019722 | -0.006218 | -0.00027579 | 0.0059422
20 | NO | 19.8884 | 31.8557 | 0.0018914 | -0.0014057 | -0.00029211 | 0.0011136
20 | NO | 9.9639 | 31.84 | 0.0015075 | -0.0029869 | -0.00030891 | 0.002678
21 | YES | 1500.39 | 34.5034 | 0.00028854 | 0.0011139 | 0.0022144 | 0.0011006
21 | YES | 1249.8 | 34.4496 | 0.00023978 | 0.0012512 | 0.0017902 | 0.00053896
21 | YES | 1000.06 | 34.3739 | 0.00037698 | 0.0020905 | 0.0013674 | -0.00072309
21 | YES | 749.935 | 34.2624 | 0.00048882 | -0.00069809 | 0.00094389 | 0.001642
22 | NO | 5.2128 | 31.644 | 0.00026465 | -0.003006 | -0.00031695 | 0.002689
23 | YES | 2004.64 | 34.5971 | 0.00026745 | 0.0033226 | 0.0030682 | -0.00025444
23 | YES | 1751.2 | 34.5492 | 0.00024765 | 0.0020409 | 0.0026391 | 0.00059822
23 | YES | 1501.27 | 34.4925 | 0.00026386 | 0.0018387 | 0.0022159 | 0.00037725
23 | YES | 1252.1 | 34.4395 | 0.00024237 | 0.0020142 | 0.0017941 | -0.00022008
23 | YES | 1001.21 | 34.3687 | 0.00054629 | 0.002018 | 0.0013693 | -0.00064866
24 | NO | 4.8468 | 31.5551 | 0.0010405 | -0.003582 | -0.00031757 | 0.0032644
25 | NO | 50.2916 | 32.3503 | 0.0011064 | 0.0048752 | -0.00024063 | -0.0051158
25 | NO | 40.6548 | 32.0962 | 0.0024558 | 0.0013008 | -0.00025695 | -0.0015578
25 | NO | 30.5373 | 31.587 | 0.010877 | -0.038578 | -0.00027408 | 0.038304
26 | YES | 299.376 | 31.839 | 0.00025303 | 0.00029182 | 0.00018108 | -0.00011075
27 | NO | 5.3457 | 28.3128 | 0.0072689 | 0.097715 | -0.00031673 | -0.098032
28 | YES | 300.061 | 30.873 | 0.00021175 | -0.0050354 | 0.00018224 | 0.0052176
28 | NO | 5.2088 | 27.8327 | 0.068316 | 0.31974 | -0.00031696 | -0.32005
29 | NO | 5.1147 | 26.7849 | 0.017792 | -0.59684 | -0.00031712 | 0.59652
30 | NO | 5.0235 | 27.0308 | 0.089804 | -0.525 | -0.00031727 | 0.52469
31 | YES | 300.254 | 31.1771 | 0.00019735 | -0.0012245 | 0.00018256 | 0.0014071
31 | NO | 5.2197 | 25.3094 | 0.080208 | -0.45157 | -0.00031694 | 0.45125
32 | YES | 299.944 | 31.1907 | 0.00026426 | 0.00015068 | 0.00018204 | 3.1358e-05
32 | NO | 4.9459 | 26.8052 | 0.002619 | -0.19226 | -0.00031741 | 0.19194
33 | NO | 5.1648 | 29.602 | 0.0027457 | 0.036495 | -0.00031703 | -0.036812
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
NO
NO
NO
YES
NO
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.063514 |  | X | CALC Y
m= | -0.0035404 | 0.0013948 |  | Std Dev of Resid |  | 0.10942 |  | 1 | 0.025702
b= | 0.029242 | 0.025238 |  | Resid Sum of Squares |  | 1.1373 |  | 33 | -0.087589
# of Values Extracted = |  |  | 97
# of Values Used in Fit = |  |  | 97
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 5.5033 | 31.8902 | 0.0064545 | 0.0077324 | 0.025702 | 0.01797 |  | 1 | 0.0077324
2 | YES | 5.4007 | 31.2491 | 0.013314 | -0.01536 | 0.022162 | 0.037521 |  | 2 | -0.01536
3 | YES | 400.684 | 34.0217 | 0.00025242 | -0.014763 | 0.018621 | 0.033384 |  | 3 | -0.014763
4 | YES | 999.769 | 34.3647 | 0.00031769 | -0.014614 | 0.015081 | 0.029695 |  | 4 | -0.014614
4 | YES | 749.78 | 34.2324 | 0.00068821 | -0.014465 | 0.015081 | 0.029546 |  | 4 | -0.014465
4 | YES | 499.917 | 34.0443 | 0.00059345 | -0.011948 | 0.015081 | 0.027028 |  | 4 | -0.011948
4 | YES | 399.818 | 34.0009 | 0.0015504 | -0.042351 | 0.015081 | 0.057432 |  | 4 | -0.042351
4 | YES | 5.003 | 31.8691 | 0.00023901 | -0.015137 | 0.015081 | 0.030218 |  | 4 | -0.015137
5 | YES | 1600.79 | 34.5195 | 0.0003023 | -0.012897 | 0.011541 | 0.024438 |  | 5 | -0.012897
5 | YES | 1499.99 | 34.5022 | 0.00025571 | -0.015114 | 0.011541 | 0.026654 |  | 5 | -0.015114
5 | YES | 1000.35 | 34.3694 | 0.00026899 | -0.013493 | 0.011541 | 0.025033 |  | 5 | -0.013493
5 | YES | 5.0381 | 31.8884 | 0.00053255 | -0.017769 | 0.011541 | 0.029309 |  | 5 | -0.017769
6 | YES | 2004.87 | 34.5799 | 0.00026907 | -0.013332 | 0.0080001 | 0.021333 |  | 6 | -0.013332
6 | YES | 1499.69 | 34.5022 | 0.00026082 | -0.01403 | 0.0080001 | 0.022031 |  | 6 | -0.01403
6 | YES | 1000.7 | 34.363 | 0.00023335 | -0.013901 | 0.0080001 | 0.021901 |  | 6 | -0.013901
7 | YES | 1250.42 | 34.4376 | 0.0002446 | -0.01384 | 0.0044598 | 0.018299 |  | 7 | -0.01384
7 | YES | 1000.45 | 34.3714 | 0.00050775 | -0.01458 | 0.0044598 | 0.01904 |  | 7 | -0.01458
7 | YES | 750.215 | 34.2573 | 0.00024963 | -0.016006 | 0.0044598 | 0.020466 |  | 7 | -0.016006
7 | YES | 299.815 | 33.9196 | 0.00026635 | -0.014969 | 0.0044598 | 0.019429 |  | 7 | -0.014969
7 | YES | 200.472 | 33.8798 | 0.0032208 | -0.015739 | 0.0044598 | 0.020199 |  | 7 | -0.015739
7 | YES | 4.9019 | 31.9393 | 0.0002229 | -0.016613 | 0.0044598 | 0.021073 |  | 7 | -0.016613
8 | YES | 500.871 | 34.0732 | 0.00026028 | -0.013901 | 0.00091944 | 0.01482 |  | 8 | -0.013901
8 | YES | 400.516 | 34.0269 | 0.00024055 | -0.015083 | 0.00091944 | 0.016003 |  | 8 | -0.015083
9 | YES | 4.84 | 30.2488 | 0.0048652 | -0.019524 | -0.0026209 | 0.016903 |  | 9 | -0.019524
10 | YES | 500.44 | 34.0507 | 0.00041012 | -0.014561 | -0.0061613 | 0.0083994 |  | 10 | -0.014561
11 | YES | 1249.9 | 34.4502 | 0.0014198 | -0.01535 | -0.0097016 | 0.0056487 |  | 11 | -0.01535
11 | YES | 1000.15 | 34.3722 | 0.00021775 | -0.013794 | -0.0097016 | 0.0040923 |  | 11 | -0.013794
11 | YES | 749.87 | 34.2412 | 0.00087882 | -0.013508 | -0.0097016 | 0.0038062 |  | 11 | -0.013508
11 | YES | 499.977 | 34.0784 | 0.00028814 | -0.014938 | -0.0097016 | 0.0052368 |  | 11 | -0.014938
12 | YES | 1999.98 | 34.5829 | 0.00029701 | -0.013817 | -0.013242 | 0.00057482 |  | 12 | -0.013817
12 | YES | 1500.14 | 34.4927 | 0.0002911 | -0.014572 | -0.013242 | 0.0013301 |  | 12 | -0.014572
12 | YES | 1249.94 | 34.4495 | 0.00025847 | -0.012516 | -0.013242 | -0.00072598 |  | 12 | -0.012516
12 | YES | 1000.25 | 34.3696 | 0.00025999 | -0.015926 | -0.013242 | 0.0026844 |  | 12 | -0.015926
12 | YES | 4.4433 | 32.083 | 0.00041769 | -0.015797 | -0.013242 | 0.0025547 |  | 12 | -0.015797
13 | YES | 750.357 | 34.2284 | 0.00048026 | -0.012958 | -0.016782 | -0.0038238 |  | 13 | -0.012958
13 | YES | 500.63 | 34.1042 | 0.00023353 | -0.01577 | -0.016782 | -0.0010123 |  | 13 | -0.01577
13 | YES | 399.978 | 34.0269 | 0.0003111 | -0.016491 | -0.016782 | -0.00029144 |  | 13 | -0.016491
13 | YES | 4.9374 | 31.1812 | 0.0012649 | -0.016546 | -0.016782 | -0.00023614 |  | 13 | -0.016546
14 | YES | 4.887 | 31.6291 | 0.017306 | 0.11501 | -0.020323 | -0.13533 |  | 14 | 0.11501
15 | YES | 4.8299 | 31.609 | 0.00043197 | -0.016439 | -0.023863 | -0.0074236 |  | 15 | -0.016439
16 | YES | 2005.5 | 34.575 | 0.0002952 | -0.011913 | -0.027403 | -0.01549 |  | 16 | -0.011913
16 | YES | 1500.67 | 34.4926 | 0.00028874 | -0.014587 | -0.027403 | -0.012816 |  | 16 | -0.014587
16 | YES | 1250.37 | 34.4381 | 0.00030415 | -0.015213 | -0.027403 | -0.01219 |  | 16 | -0.015213
16 | YES | 1000.06 | 34.3571 | 0.00029141 | -0.013882 | -0.027403 | -0.013522 |  | 16 | -0.013882
17 | YES | 2001.16 | 34.5805 | 0.00025166 | -0.01347 | -0.030944 | -0.017474 |  | 17 | -0.01347
17 | YES | 1500.51 | 34.4931 | 0.00031121 | -0.014763 | -0.030944 | -0.016181 |  | 17 | -0.014763
17 | YES | 1250.89 | 34.4405 | 0.00027121 | -0.01424 | -0.030944 | -0.016703 |  | 17 | -0.01424
17 | YES | 1001.2 | 34.3674 | 0.00030601 | -0.013638 | -0.030944 | -0.017306 |  | 17 | -0.013638
18 | YES | 1249.88 | 34.4185 | 0.00024692 | -0.015251 | -0.034484 | -0.019233 |  | 18 | -0.015251
18 | YES | 999.458 | 34.3466 | 0.00021032 | -0.014343 | -0.034484 | -0.020141 |  | 18 | -0.014343
18 | YES | 749.064 | 34.2108 | 0.00022792 | -0.014259 | -0.034484 | -0.020225 |  | 18 | -0.014259
18 | YES | 500.611 | 34.0494 | 0.00054046 | -0.012966 | -0.034484 | -0.021518 |  | 18 | -0.012966
19 | YES | 87.4554 | 33.6795 | 0.0024446 | -0.025131 | -0.038024 | -0.012893 |  | 19 | -0.025131
20 | YES | 1999.81 | 34.5761 | 0.00028433 | -0.01144 | -0.041565 | -0.030125 |  | 20 | -0.01144
20 | YES | 1499.95 | 34.4928 | 0.00030234 | -0.014046 | -0.041565 | -0.027519 |  | 20 | -0.014046
20 | YES | 1249.5 | 34.4407 | 0.00031627 | -0.012215 | -0.041565 | -0.02935 |  | 20 | -0.012215
20 | YES | 999.526 | 34.3588 | 0.0009639 | -0.013008 | -0.041565 | -0.028557 |  | 20 | -0.013008
20 | YES | 749.773 | 34.2521 | 0.0007455 | -0.012764 | -0.041565 | -0.028801 |  | 20 | -0.012764
20 | YES | 499.713 | 34.0385 | 0.00041172 | -0.016563 | -0.041565 | -0.025001 |  | 20 | -0.016563
20 | YES | 400.008 | 33.9698 | 0.00027041 | -0.013092 | -0.041565 | -0.028473 |  | 20 | -0.013092
20 | YES | 300.351 | 33.9249 | 0.00045248 | -0.012928 | -0.041565 | -0.028637 |  | 20 | -0.012928
20 | YES | 249.649 | 33.898 | 0.00025133 | -0.01379 | -0.041565 | -0.027775 |  | 20 | -0.01379
20 | YES | 199.798 | 33.876 | 0.00028964 | -0.017773 | -0.041565 | -0.023792 |  | 20 | -0.017773
20 | YES | 175.224 | 33.8298 | 0.0011428 | -0.016277 | -0.041565 | -0.025288 |  | 20 | -0.016277
20 | YES | 149.667 | 33.7162 | 0.0023905 | -0.019627 | -0.041565 | -0.021938 |  | 20 | -0.019627
20 | YES | 125.07 | 33.4654 | 0.0064108 | -0.012691 | -0.041565 | -0.028873 |  | 20 | -0.012691
20 | YES | 99.5247 | 33.2999 | 0.0021689 | 0.003437 | -0.041565 | -0.045002 |  | 20 | 0.003437
20 | YES | 74.7872 | 32.8025 | 0.0093709 | -0.003067 | -0.041565 | -0.038498 |  | 20 | -0.003067
20 | YES | 49.831 | 32.3668 | 0.0036721 | 0.015995 | -0.041565 | -0.05756 |  | 20 | 0.015995
20 | YES | 29.5248 | 31.9172 | 0.015671 | -0.012842 | -0.041565 | -0.028723 |  | 20 | -0.012842
20 | YES | 19.8884 | 31.8441 | 0.0032832 | -0.013012 | -0.041565 | -0.028553 |  | 20 | -0.013012
20 | YES | 9.9639 | 31.8282 | 0.0024508 | -0.014761 | -0.041565 | -0.026804 |  | 20 | -0.014761
21 | YES | 1500.39 | 34.4887 | 0.00029526 | -0.013603 | -0.045105 | -0.031502 |  | 21 | -0.013603
21 | YES | 1249.8 | 34.4348 | 0.00023086 | -0.013493 | -0.045105 | -0.031613 |  | 21 | -0.013493
21 | YES | 1000.06 | 34.3598 | 0.00028416 | -0.012005 | -0.045105 | -0.0331 |  | 21 | -0.012005
21 | YES | 749.935 | 34.2482 | 0.00045628 | -0.014934 | -0.045105 | -0.030171 |  | 21 | -0.014934
22 | YES | 5.2128 | 31.6303 | 0.00025787 | -0.016724 | -0.048646 | -0.031922 |  | 22 | -0.016724
23 | YES | 2004.64 | 34.5811 | 0.00030629 | -0.012653 | -0.052186 | -0.039532 |  | 23 | -0.012653
23 | YES | 1751.2 | 34.5337 | 0.00028904 | -0.013534 | -0.052186 | -0.038651 |  | 23 | -0.013534
23 | YES | 1501.27 | 34.4771 | 0.00039666 | -0.013649 | -0.052186 | -0.038537 |  | 23 | -0.013649
23 | YES | 1252.1 | 34.4241 | 0.00027371 | -0.013355 | -0.052186 | -0.038831 |  | 23 | -0.013355
23 | YES | 1001.21 | 34.354 | 0.00047798 | -0.012722 | -0.052186 | -0.039464 |  | 23 | -0.012722
24 | YES | 4.8468 | 31.5399 | 0.0021193 | -0.018846 | -0.055726 | -0.03688 |  | 24 | -0.018846
25 | YES | 50.2916 | 32.3349 | 0.0010036 | -0.01054 | -0.059267 | -0.048727 |  | 25 | -0.01054
25 | YES | 40.6548 | 32.0805 | 0.0013655 | -0.014355 | -0.059267 | -0.044912 |  | 25 | -0.014355
25 | YES | 30.5373 | 31.5733 | 0.01399 | -0.052298 | -0.059267 | -0.006969 |  | 25 | -0.052298
26 | YES | 299.376 | 31.8252 | 0.00020299 | -0.013512 | -0.062807 | -0.049295 |  | 26 | -0.013512
27 | YES | 5.3457 | 28.3089 | 0.015173 | 0.093794 | -0.066347 | -0.16014 |  | 27 | 0.093794
28 | YES | 300.061 | 30.861 | 0.00023845 | -0.01697 | -0.069888 | -0.052918 |  | 28 | -0.01697
28 | YES | 5.2088 | 27.8348 | 0.09733 | 0.32181 | -0.069888 | -0.3917 |  | 28 | 0.32181
29 | YES | 5.1147 | 26.7518 | 0.010438 | -0.62987 | -0.073428 | 0.55644 |  | 29 | -0.62987
30 | YES | 5.0235 | 26.9474 | 0.071326 | -0.60843 | -0.076968 | 0.53146 |  | 30 | -0.60843
31 | YES | 300.254 | 31.1653 | 0.0002006 | -0.013031 | -0.080509 | -0.067478 |  | 31 | -0.013031
31 | YES | 5.2197 | 25.1824 | 0.033492 | -0.57864 | -0.080509 | 0.49813 |  | 31 | -0.57864
32 | YES | 299.944 | 31.1791 | 0.00020559 | -0.011435 | -0.084049 | -0.072614 |  | 32 | -0.011435
32 | YES | 4.9459 | 26.794 | 0.003763 | -0.2035 | -0.084049 | 0.11945 |  | 32 | -0.2035
33 | YES | 5.1648 | 29.5887 | 0.00055474 | 0.023189 | -0.087589 | -0.11078 |  | 33 | 0.023189
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.063224 |  | X | CALC Y
m= | -0.0031938 | 0.0012613 |  | Std Dev of Resid |  | 0.098946 |  | 1 | 0.037069
b= | 0.040263 | 0.022823 |  | Resid Sum of Squares |  | 0.93007 |  | 33 | -0.065131
# of Values Extracted = |  |  | 97
# of Values Used in Fit = |  |  | 97
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 5.5033 | 31.8993 | 0.006048 | 0.016819 | 0.037069 | 0.02025 |  | 1 | 0.016819
2 | YES | 5.4007 | 31.2652 | 0.020152 | 0.00069237 | 0.033875 | 0.033183 |  | 2 | 0.00069237
3 | YES | 400.684 | 34.0375 | 0.00024373 | 0.00098038 | 0.030681 | 0.029701 |  | 3 | 0.00098038
4 | YES | 999.769 | 34.3802 | 0.00032798 | 0.00089264 | 0.027488 | 0.026595 |  | 4 | 0.00089264
4 | YES | 749.78 | 34.2478 | 0.00078512 | 0.00085449 | 0.027488 | 0.026633 |  | 4 | 0.00085449
4 | YES | 499.917 | 34.0597 | 0.00074487 | 0.0034943 | 0.027488 | 0.023993 |  | 4 | 0.0034943
4 | YES | 399.818 | 34.0174 | 0.001251 | -0.025841 | 0.027488 | 0.053329 |  | 4 | -0.025841
4 | YES | 5.003 | 31.8831 | 0.00019925 | -0.00107 | 0.027488 | 0.028558 |  | 4 | -0.00107
5 | YES | 1600.79 | 34.536 | 0.00026636 | 0.0036469 | 0.024294 | 0.020647 |  | 5 | 0.0036469
5 | YES | 1499.99 | 34.5186 | 0.00027346 | 0.0013466 | 0.024294 | 0.022947 |  | 5 | 0.0013466
5 | YES | 1000.35 | 34.3853 | 0.000337 | 0.0024338 | 0.024294 | 0.02186 |  | 5 | 0.0024338
5 | YES | 5.0381 | 31.9029 | 0.00077753 | -0.0032654 | 0.024294 | 0.027559 |  | 5 | -0.0032654
6 | YES | 2004.87 | 34.5967 | 0.00024344 | 0.0034866 | 0.0211 | 0.017614 |  | 6 | 0.0034866
6 | YES | 1499.69 | 34.5184 | 0.00025723 | 0.002243 | 0.0211 | 0.018857 |  | 6 | 0.002243
6 | YES | 1000.7 | 34.3789 | 0.00023092 | 0.0020256 | 0.0211 | 0.019075 |  | 6 | 0.0020256
7 | YES | 1250.42 | 34.4539 | 0.00024976 | 0.0024834 | 0.017906 | 0.015423 |  | 7 | 0.0024834
7 | YES | 1000.45 | 34.3876 | 0.00037061 | 0.0016213 | 0.017906 | 0.016285 |  | 7 | 0.0016213
7 | YES | 750.215 | 34.2732 | 0.0002485 | -0.00013351 | 0.017906 | 0.01804 |  | 7 | -0.00013351
7 | YES | 299.815 | 33.9354 | 0.00029319 | 0.00079346 | 0.017906 | 0.017113 |  | 7 | 0.00079346
7 | YES | 200.472 | 33.8959 | 0.00031884 | 0.0004425 | 0.017906 | 0.017464 |  | 7 | 0.0004425
7 | YES | 4.9019 | 31.954 | 0.00025289 | -0.0019417 | 0.017906 | 0.019848 |  | 7 | -0.0019417
8 | YES | 500.871 | 34.0887 | 0.00029972 | 0.0016403 | 0.014713 | 0.013072 |  | 8 | 0.0016403
8 | YES | 400.516 | 34.0425 | 0.00022015 | 0.00054169 | 0.014713 | 0.014171 |  | 8 | 0.00054169
9 | YES | 4.84 | 30.2616 | 0.0043393 | -0.0066834 | 0.011519 | 0.018202 |  | 9 | -0.0066834
10 | YES | 500.44 | 34.0667 | 0.00074509 | 0.0014114 | 0.0083252 | 0.0069137 |  | 10 | 0.0014114
11 | YES | 1249.9 | 34.4667 | 0.00066694 | 0.0010757 | 0.0051314 | 0.0040557 |  | 11 | 0.0010757
11 | YES | 1000.15 | 34.3878 | 0.00025064 | 0.0017891 | 0.0051314 | 0.0033423 |  | 11 | 0.0017891
11 | YES | 749.87 | 34.2566 | 0.00062783 | 0.0019455 | 0.0051314 | 0.0031859 |  | 11 | 0.0019455
11 | YES | 499.977 | 34.0937 | 0.00024801 | 0.00035858 | 0.0051314 | 0.0047728 |  | 11 | 0.00035858
12 | YES | 1999.98 | 34.5995 | 0.00026104 | 0.0028076 | 0.0019376 | -0.00086998 |  | 12 | 0.0028076
12 | YES | 1500.14 | 34.5089 | 0.00029375 | 0.0016174 | 0.0019376 | 0.00032021 |  | 12 | 0.0016174
12 | YES | 1249.94 | 34.4656 | 0.00023307 | 0.0036163 | 0.0019376 | -0.0016787 |  | 12 | 0.0036163
12 | YES | 1000.25 | 34.3853 | 0.00024216 | -0.00016403 | 0.0019376 | 0.0021017 |  | 12 | -0.00016403
12 | YES | 4.4433 | 32.0971 | 0.00051837 | -0.0016556 | 0.0019376 | 0.0035932 |  | 12 | -0.0016556
13 | YES | 750.357 | 34.2437 | 0.00035157 | 0.0023422 | -0.0012561 | -0.0035983 |  | 13 | 0.0023422
13 | YES | 500.63 | 34.1197 | 0.00023104 | -0.00032806 | -0.0012561 | -0.00092805 |  | 13 | -0.00032806
13 | YES | 399.978 | 34.0423 | 0.00031916 | -0.0011253 | -0.0012561 | -0.00013078 |  | 13 | -0.0011253
13 | YES | 4.9374 | 31.1932 | 0.0010302 | -0.004528 | -0.0012561 | 0.0032719 |  | 13 | -0.004528
14 | YES | 4.887 | 31.6379 | 0.010963 | 0.12385 | -0.0044499 | -0.1283 |  | 14 | 0.12385
15 | YES | 4.8299 | 31.6232 | 0.0004461 | -0.0021973 | -0.0076436 | -0.0054464 |  | 15 | -0.0021973
16 | YES | 2005.5 | 34.5913 | 0.00022572 | 0.0044365 | -0.010837 | -0.015274 |  | 16 | 0.0044365
16 | YES | 1500.67 | 34.5085 | 0.00025337 | 0.0013313 | -0.010837 | -0.012169 |  | 16 | 0.0013313
16 | YES | 1250.37 | 34.4536 | 0.00028762 | 0.00027466 | -0.010837 | -0.011112 |  | 16 | 0.00027466
16 | YES | 1000.06 | 34.3725 | 0.00031473 | 0.0014572 | -0.010837 | -0.012295 |  | 16 | 0.0014572
17 | YES | 2001.16 | 34.5969 | 0.00023106 | 0.002903 | -0.014031 | -0.016934 |  | 17 | 0.002903
17 | YES | 1500.51 | 34.5086 | 0.0005391 | 0.00072479 | -0.014031 | -0.014756 |  | 17 | 0.00072479
17 | YES | 1250.89 | 34.456 | 0.00023702 | 0.0012779 | -0.014031 | -0.015309 |  | 17 | 0.0012779
17 | YES | 1001.2 | 34.3827 | 0.00032176 | 0.0017204 | -0.014031 | -0.015752 |  | 17 | 0.0017204
18 | YES | 1249.88 | 34.4343 | 0.00023681 | 0.00053024 | -0.017225 | -0.017755 |  | 18 | 0.00053024
18 | YES | 999.458 | 34.3622 | 0.00024156 | 0.0012856 | -0.017225 | -0.01851 |  | 18 | 0.0012856
18 | YES | 749.064 | 34.2261 | 0.0002192 | 0.0010414 | -0.017225 | -0.018266 |  | 18 | 0.0010414
18 | YES | 500.611 | 34.0649 | 0.00040865 | 0.002533 | -0.017225 | -0.019758 |  | 18 | 0.002533
19 | YES | 87.4554 | 33.7063 | 0.00044582 | 0.0016861 | -0.020419 | -0.022105 |  | 19 | 0.0016861
20 | YES | 1999.81 | 34.5917 | 0.00025465 | 0.004158 | -0.023612 | -0.02777 |  | 20 | 0.004158
20 | YES | 1499.95 | 34.5077 | 0.00027991 | 0.00092697 | -0.023612 | -0.024539 |  | 20 | 0.00092697
20 | YES | 1249.5 | 34.4555 | 0.00040275 | 0.0026207 | -0.023612 | -0.026233 |  | 20 | 0.0026207
20 | YES | 999.526 | 34.3742 | 0.00052184 | 0.0024262 | -0.023612 | -0.026039 |  | 20 | 0.0024262
20 | YES | 749.773 | 34.2675 | 0.00064808 | 0.002552 | -0.023612 | -0.026164 |  | 20 | 0.002552
20 | YES | 499.713 | 34.0527 | 0.00037488 | -0.0023766 | -0.023612 | -0.021236 |  | 20 | -0.0023766
20 | YES | 400.008 | 33.9842 | 0.00031366 | 0.0013084 | -0.023612 | -0.024921 |  | 20 | 0.0013084
20 | YES | 300.351 | 33.9393 | 0.00059728 | 0.0015297 | -0.023612 | -0.025142 |  | 20 | 0.0015297
20 | YES | 249.649 | 33.9123 | 0.00026877 | 0.00054932 | -0.023612 | -0.024162 |  | 20 | 0.00054932
20 | YES | 199.798 | 33.8903 | 0.00038065 | -0.0035286 | -0.023612 | -0.020084 |  | 20 | -0.0035286
20 | YES | 175.224 | 33.8392 | 0.0025712 | -0.006916 | -0.023612 | -0.016696 |  | 20 | -0.006916
20 | YES | 149.667 | 33.7287 | 0.0022898 | -0.007103 | -0.023612 | -0.016509 |  | 20 | -0.007103
20 | YES | 125.07 | 33.4802 | 0.0039028 | 0.0020828 | -0.023612 | -0.025695 |  | 20 | 0.0020828
20 | YES | 99.5247 | 33.3147 | 0.0021777 | 0.018169 | -0.023612 | -0.041782 |  | 20 | 0.018169
20 | YES | 74.7872 | 32.818 | 0.0090986 | 0.012386 | -0.023612 | -0.035999 |  | 20 | 0.012386
20 | YES | 49.831 | 32.3868 | 0.0025859 | 0.035965 | -0.023612 | -0.059577 |  | 20 | 0.035965
20 | YES | 29.5248 | 31.9238 | 0.019722 | -0.006218 | -0.023612 | -0.017394 |  | 20 | -0.006218
20 | YES | 19.8884 | 31.8557 | 0.0018914 | -0.0014057 | -0.023612 | -0.022207 |  | 20 | -0.0014057
20 | YES | 9.9639 | 31.84 | 0.0015075 | -0.0029869 | -0.023612 | -0.020626 |  | 20 | -0.0029869
21 | YES | 1500.39 | 34.5034 | 0.00028854 | 0.0011139 | -0.026806 | -0.02792 |  | 21 | 0.0011139
21 | YES | 1249.8 | 34.4496 | 0.00023978 | 0.0012512 | -0.026806 | -0.028057 |  | 21 | 0.0012512
21 | YES | 1000.06 | 34.3739 | 0.00037698 | 0.0020905 | -0.026806 | -0.028897 |  | 21 | 0.0020905
21 | YES | 749.935 | 34.2624 | 0.00048882 | -0.00069809 | -0.026806 | -0.026108 |  | 21 | -0.00069809
22 | YES | 5.2128 | 31.644 | 0.00026465 | -0.003006 | -0.03 | -0.026994 |  | 22 | -0.003006
23 | YES | 2004.64 | 34.5971 | 0.00026745 | 0.0033226 | -0.033194 | -0.036516 |  | 23 | 0.0033226
23 | YES | 1751.2 | 34.5492 | 0.00024765 | 0.0020409 | -0.033194 | -0.035235 |  | 23 | 0.0020409
23 | YES | 1501.27 | 34.4925 | 0.00026386 | 0.0018387 | -0.033194 | -0.035032 |  | 23 | 0.0018387
23 | YES | 1252.1 | 34.4395 | 0.00024237 | 0.0020142 | -0.033194 | -0.035208 |  | 23 | 0.0020142
23 | YES | 1001.21 | 34.3687 | 0.00054629 | 0.002018 | -0.033194 | -0.035212 |  | 23 | 0.002018
24 | YES | 4.8468 | 31.5551 | 0.0010405 | -0.003582 | -0.036387 | -0.032805 |  | 24 | -0.003582
25 | YES | 50.2916 | 32.3503 | 0.0011064 | 0.0048752 | -0.039581 | -0.044456 |  | 25 | 0.0048752
25 | YES | 40.6548 | 32.0962 | 0.0024558 | 0.0013008 | -0.039581 | -0.040882 |  | 25 | 0.0013008
25 | YES | 30.5373 | 31.587 | 0.010877 | -0.038578 | -0.039581 | -0.0010032 |  | 25 | -0.038578
26 | YES | 299.376 | 31.839 | 0.00025303 | 0.00029182 | -0.042775 | -0.043067 |  | 26 | 0.00029182
27 | YES | 5.3457 | 28.3128 | 0.0072689 | 0.097715 | -0.045969 | -0.14368 |  | 27 | 0.097715
28 | YES | 300.061 | 30.873 | 0.00021175 | -0.0050354 | -0.049162 | -0.044127 |  | 28 | -0.0050354
28 | YES | 5.2088 | 27.8327 | 0.068316 | 0.31974 | -0.049162 | -0.3689 |  | 28 | 0.31974
29 | YES | 5.1147 | 26.7849 | 0.017792 | -0.59684 | -0.052356 | 0.54448 |  | 29 | -0.59684
30 | YES | 5.0235 | 27.0308 | 0.089804 | -0.525 | -0.05555 | 0.46945 |  | 30 | -0.525
31 | YES | 300.254 | 31.1771 | 0.00019735 | -0.0012245 | -0.058744 | -0.057519 |  | 31 | -0.0012245
31 | YES | 5.2197 | 25.3094 | 0.080208 | -0.45157 | -0.058744 | 0.39283 |  | 31 | -0.45157
32 | YES | 299.944 | 31.1907 | 0.00026426 | 0.00015068 | -0.061938 | -0.062088 |  | 32 | 0.00015068
32 | YES | 4.9459 | 26.8052 | 0.002619 | -0.19226 | -0.061938 | 0.13032 |  | 32 | -0.19226
33 | YES | 5.1648 | 29.602 | 0.0027457 | 0.036495 | -0.065131 | -0.10163 |  | 33 | 0.036495
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE