## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2020-017

Agency: Ecosystem Science Division (ESD)

Location: Queen Charlotte Sound, Hecate Strait, Dixon Entrance

Project: October Juvenile Salmon

Party Chief: Anderson E

Platform: Sir John Franklin

Date: 6 October 2020 – 15 October 2020

Processed by: Germaine Gatien

Date of Processing: 22 December 2020 – 16 April 2021

Number of original HEX files: 22 (2 test files)   	Number of CTD files:  (2 SBE911, 18 SBE25)

Number of bottle files:	3 (1 test)   		Number of bottle casts processed: 2 CHE 18 BOT

# INSTRUMENT SUMMARY

1, SBE911+ CTD #0506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), a SBE 43 DO sensor (#0047), a Wet Labs ECO-AFL Fluorometer (#2215) and a Teledyne Benthos 916P altimeter (#1252).

Seasave version 7.26.7.121 was used for acquisition.

The data logging computer was WP #103.

The deck unit was Seabird model 11+ #1043.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

2. SBE25+ #334 was mounted with an SBE 43 DO sensor (#1483) and Wet Labs ECO-AFL Fluorometer (#2215). Sampling was done using a bucket at the surface near the beginning of CTD casts.

# SUMMARY OF QUALITY AND CONCERNS

An SBE911+ was in use at the beginning of the cruise but the cable broke while moving in the LARS. A backup SBE25+ was used for the rest of the cruise but the smaller Niskin bottles did not fit on the conductivity cable used. After event #12 all sampling was taken from a surface bucket at the beginning of CTD casts. The CTD was lowered to between 10 and 12m for a pre-cast soak then raised to near the surface.

Bottle files with extensions CHE contain analysis results for samples collected from the rosette together with CTD data acquired at the same time as the bottle closed.

Bottle files with extensions BOT contain analysis results from samples collected from surface buckets together with the shallowest CTD data acquired after the 10m soak.

The Daily Science Log was in good order with details about problems encountered. The rosette logs were available in single spreadsheets. Given the many problems arising from a new ship, a broken cable, a need to switch CTDs, an inability to use Niskins on a wire and COVID restrictions, the record-keeping was remarkably good.

File names were non-standard; event numbers should have 4 digits. The cruise number was wrong in a few files. Times and dates were wrong in the SeaBird headers for the SBE25. They were corrected in processing.

The configuration file provided did not have the most recent values. This led to considerable confusion as to why the pressure data looked so bad. Working remotely made this

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available. For most casts chlorophyll is only available right at the surface.

Recalibration of salinity and dissolved oxygen were applied to CTD and CHE files from events #5 and #12. No recalibration was applied to the SBE25 CTD files, but fluorescence data was corrected because there was an error in the original conversion. The fluorescence was corrected in SBE25 BOT files earlier in processing.

A study of oxygen saturation near the surface was used to determine a correction factor for SBE dissolved oxygen for the 2 casts using the SBE911. No adjustment was made to that channel for the SBE25 casts because it used a different sensor for which there was insufficient evidence to justify a correction.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained

PART I – SBE911

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2020-017-ctd911.xmlcon. Depth was included.

(Note: There was an error in the configuration file that led to fluorescence values being too high by 0.169ug/L at all depths. This was corrected later.)

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. None were found.

There are no positions in the headers so obviously the NMEA download failed.

A text editor was used to enter the information based on log entries.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2020-017-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2020-017chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2020-017chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file 2020-017SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 46 days after collection. The files were simplified and saved as 2020-017sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2020-017_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2020-017-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were examined and no data were missing.

##### Compare

Salinity

There were 3 salinity samples at 2100db, 51db and 3db.

The primary salinity was found to be high by 0.0055, 0.0088 and 0.0033psu, respectively.

The secondary salinity was found to be high by 0.0030, 0.0058 and 0.0008psu respectively.

The deepest values tend to most reliable due to low gradients, but bottles fired near the ocean bottom may have somewhat lower salinity due to incomplete flushing of Niskin bottles as shed wakes arrive from above. That would lead to the CTD reading higher than it really is. At 50db there is a fairly high gradient and incomplete flushing of Niskin bottles could make the CTD salinity look lower than it really is, though for both sensors the CTD values are especially high relative to bottles. The near-surface waters are well mixed so that value may be the most reliable.

There was a 46 day delay in analysis which is likely to raise the bottle values slightly, and that would lead to the CTD looking lower than it really is. With these competing bits of evidence and only 3 bottles to judge by, the only conclusion is that the secondary salinity is likely the better choice for accuracy and that the two channels differ by about 0.0025db.

Fluorescence

An error in the fluorescence offset in the configuration file led to values too high by 0.169ug/L. The values listed below are after correction of that error.

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. There were only 5 samples and CHL values were all <1ug/L. Fluorescence values were close to CHL for the 2 deepest samples but much higher than CHL in the top 25db.

| Event | Pressure | CHL | FL | FL/CHL |
| --- | --- | --- | --- | --- |
| 5 | 75.1 | 0.32 | 0.36 | 0.29 |
| 5 | 50.9 | 0.94 | 0.84 | 0.61 |
| 5 | 25.4 | 0.71 | 2.21 | 2.74 |
| 5 | 11.1 | 0.38 | 1.84 | 4.15 |
| 5 | 5.2 | 0.75 | 1.79 | 2.03 |
| 12 | 9.8 | 0.91 | 2.15 | 2.07 |

Fluorometers tend to read higher than CHL at low CHL values, so this is not out of line with expected performance. For full details of the comparison see file 2020-017-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using 2020-017-ctd.

(Note: There was an error in the configuration file that led to fluorescence values being too high by 0.169ug/L at all depths. This was corrected later.)

The Tau function was selected and also the hysteresis function as there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

- There was a lot of noise in the temperature and conductivity, mostly in the upcast.

- The Dissolved Oxygen. and transmissivity looked reasonable.

- Fluorescence had a dark value of ~0.28ug/L which is unusually high.

- The altimetry looks reasonable near the bottom for the one cast that got close enough.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. That setting was used during 2018-01.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between the channel pairs was examined:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2020-017-0005 | 900 | -0.0006 | -0.00025 | -0.003 | High, Noisy |
| 2020-017-0012 | 1030 | -0.0006 | -0.00025 | -0.003 | High, Moderate |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

The positions were missing from the headers so Merge Headers was run to add that information to the headers using input file 2020-017-merge-hdr.csv.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

There was an initial soak between 10 and 12m and those data were included in the data files. CLIP was run to remove the first 10785 records from cast #5 and 8270 records from cast #12.

##### Checking Headers

The SBE25 casts were processed up to the point of header checks so they could all be run together.

The header check was run and a few problems were found and corrected.

A cross-reference list was checked against the log book. The only problem found was for event #32. The log entries were wrong for the NET and CTD events at DE01 but 51° was changed to 54° in the log entry for the SET event. This looks correct, so the file header was changed to 54°.

Cruise track plots were added to the end of this report including the SBE25 casts.

The altimeter worked well for cast #12 and confirms the water depth entry as being reasonable. For cast #5 the CTD did not get close enough to the bottom for useful altimetry. There was no altimeter mounted on the SBE25.

For the SBE911 there was one deck pressure measurement of +1.1db. Examination of the near-surface data was limited by few casts and pumps not being on, but from transmissivity at the beginning of casts it appears that the CTD was in water at 1db. Salinity at the end of one upcast when pumps were still running suggests that the CTD was still in the water at 0.4db. There is insufficient evidence to justify recalibration of pressure.

##### Shift

Fluorescence

SHIFT was not run on the fluorescence channel since it was not pumped.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on the 2 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space.

SHIFT was run twice on all SBE911 casts using -0.3 records for the primary conductivity and -0.8 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning pertained to upcast data.

##### Other Comparisons

Previous experience with these sensors –

- SBE911+ sensors – Temperature, conductivity and dissolved oxygen sensors have not been used since the last factory calibration. The pressure sensor was used during many cruises in 2019 and no problems were reported.

- SBE25+ sensors – The temperature, conductivity and dissolved oxygen sensors have been used on 4 previous cruise. No recalibration was applied to salinity since there were few bottles and the comparisons were not considered reliable. Dissolved oxygen was not recalibrated as there was insufficient sampling but it appeared to be low by about 2%. The pressure sensor was recalibrated and the range changed since it was last used.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All T and S data fell within the climatology.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary T-S pair were chosen for editing and eventual archiving because they were slightly less noisy in T-S plots and were closer to bottle salinity values.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and bottom of casts that appear to be corrupted by shed wakes or ship effects. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. Both files required fairly heavy editing due to noisy descent rates.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts and while there are small unstable features they may be real so no further editing was applied.

##### Recalibration

There was no dissolved oxygen sampling. The 2 911+ casts were in deep offshore waters where surface oxygen saturation values to be in the range 100-103%. For these 2 casts it was in the range 96-99%. SBE DO sensors typically drift to lower values. Multiplying values by 1.02 is a reasonable estimate.

It was discovered late that the fluorescence dark values were unreasonably high. The fluorometer offset had been set wrong which lead to values high by 0.169ug/L at all depths.

File 2020-017-SBE911-recal.ccf was prepared to subtract 0.169ug/L from fluorescence, subtract  0.0035psu from channel Salinity T0:C0, subtract 0.001psu from channel Salinity:T1:C1 and to apply the following corrections to channels Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.02

This correction was first applied to the SAM and MRGCLN2 files.

The results were examined and the corrections were applied properly.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Fluorescence Processing

The fluorometer was an ECO and did not need filtering. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

- A second SBE DO channel (with umol/kg units) was added.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found. This will be repeated after the SBE25 data are completed so all can be reported together.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

- A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

- Standards check was run. 1 format error was found and corrected.

- A header check were run. No problems were found.

- Plots of each file were examined and no problems were found.

- A cross-reference listing was produced for the CHE files.

PART II – SBE25

There were rosette spreadsheet files prepared for each cast but these were prepared in advance and there may have been some changes later. File “2020-017-Event_Log.xlsx contains corrected station data and sampling. There is a note in the log that times were wrong on the computer (local time and wrong day).

##### Seasave

This step was completed at sea.

##### Preliminary Steps

The cruise summary sheet was completed.

A configuration file was provided and was saved as 2020-017-ctd25.xmlcon.

##### Conversion of Raw Data

All files were converted using configuration file 2020-017-ctd25.xmlcon.

There was a turbidity meter listed in the log and entered under User Polynomial but all parameters are 0 except for some surface spikes, and converted values are all 0. So data were reconverted without that channel.

The pressure is clearly wrong. Investigation led to the discovery that the range of the sensor had been changed and the pressure sensor had been calibrated in 2020. The configuration file was updated and conversion was rerun.

Plots were examined and all expected channels were present and values look reasonable except for spikes at the surface. The deployment method was to lower the CTD to about 10-12m and return to the surface before running the full cast.

Based on conductivity values the pressure looks accurate within ±0.2db; ±1db is the accuracy quoted by SeaBird for the pressure sensor so this looks good.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

There are frequent small steps in the pressure and some very small reversals. Tests were run during 2008-005 to see if just filtering pressure was sufficient to improve the salinity and it was not. A variety of filters were tested and the best choice was found to be a Windows cosine filter size 7.

WFILTER was run on pressure, depth, temperature, conductivity and oxygen voltage with cosine file width 7. This setting smooths steps but leaves very small reversals in pressure.

##### ALIGNCTD

During 2016-40, 2017-41, 2017-52, 2017-53 and 2018-05 when this DO sensor was also used the DO channel was advanced by 2.5s. That setting looks reasonable for these data as well.

ALIGNCTD was run to apply a 2.5s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Times recorded from the computer were in PST plus about 12 minutes and a day ahead, so ADD TIME CHANNEL was run to subtract 16.2 hours to the Start Times so they match the log entries reasonably well.

Station names, locations and water depth were not included in the CNV file headers so were not automatically added to the IOS Header files upon conversion. A spreadsheet was prepared with those details, 2020-017-consolidation.xlsx. Merge Headers was used to add that information to the headers.

CLEAN was run to add event numbers in case any were missed and to interpolate any pad values in pressure. The only negative fluorescence values are at the surface and will be removed when the soak data are removed.

##### Checking Headers

The cross-reference check was run. No further problems were noted.

Files were examined to make an estimate of the # of records that should be removed to eliminate data from the 10m soak.

CLIP was used to remove initial records from the soak period.  Plots were examined until the number of records removed were appropriate

Track plot and Header Check were run together with the SBE911 data and no problems were found.

The track plots were added to the end of this report.

The surface check shows an average of -0.07db on files before CLIP was applied. That included 2 casts from the SBE911 which did not come on until the CTD was in the water and had surface pressure of 1.46 and 1.26. Most of the SBE25 casts start in air and have values from -0.03 to -0.20db. A few were slightly higher at 0.32 and 0.15db but those values only occur in the first few records, so look like spikes.

The SBE25 pressure appears to be accurate.

##### SHIFT

Fluorescence

The fluorometer on the SBE25 was not pumped so alignment was not needed.

Conductivity

Applying a shift of +0.5records to the conductivity improved the salinity as judged by stability in T-S space after running DELETE. A larger shift improved some features while making others worse.

SHIFT was applied to all casts using +0.5 records.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; the alignment looks reasonable. SHIFT was not run on DO.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure. Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

Examination of T-S plots shows a need to edit surface and near-bottom data for most casts and some other editing for a few.

CTDEDIT was used to remove near-surface records and some data corrupted by shed wakes from most casts; salinity was cleaned lightly in some casts. Notes of editing details were made in the headers.

All casts except #53 required some editing.

The edited files were copied to *.EDT.

T-S plots were examined and no problems were found.

##### Other calibration checks

Sensor History – The sensors were used on 4 previous cruises since they were last serviced except for the pressure sensor which was newly serviced. There was no calibration sampling for 2 of those cruises and for the other 2 sampling was from a bottle on a wire so the comparison was not considered reliable enough to justify recalibration. The most recent comparison suggested that salinity might be high by ~0.005psu.

Comparison of repeat casts – None were available.

Historic Ranges – All temperature and salinity data were within the climatology.

Post-cruise calibrations – None were available.

Inter-comparison of the 2 CTDs – There was no inter-comparison between the SBE911 and SBE25.

##### Bottle File Preparation

##### As there are no surface data available from the level of the bucket samples, the shallowest data available were found so that temperature, salinity, dissolved oxygen and fluorescence data could be placed in the BOT files containing analysis results.

##### The DEL files were clipped so they contained only the first record. Data from those files were output to file 2020-017-ctd-output.csv. That file plus information from the analysis spreadsheets, event log and Daily Science Log Book were used to prepare spreadsheet 2020-017-bottles_plus_CTD-6linehdr.csv.

##### The fluorescence values were corrected by subtracting 0.169ug/L so fluorescence recalibration will not be needed later.

There was no salinity sampling.

##### The spreadsheet was then converted to IOS files.

CLEAN was run to add 0s to empty flag channels.

REMOVE was used to remove time and date channels.

EDIT HEADERS was used to correct some formats and add a comment to the headers.

Standards Check was run and no problems were found.

##### CALIBRATE

There were bucket salinity samples from the surface. The only comparison possible with samples is between the CHL and the corrected fluorescence from the BOT files. When 1 outlier is excluded (one of 2 samples that were flagged 3 by the analyst) the fit is:

Given the slightly different depths this is a good correspondence.

One other check on calibration was done by deriving oxygen saturation. Plots show that the dissolved oxygen values right at the surface are lower than at 2 or 3db. Values at about 2db ranged from 90% to 99% except for station QCSD01 where it was 75%. The low value was in Johnstone Strait where tidal mixing tends to keep values low. The highest value was 99% at station H02. All casts in Hecate Strait had DO saturation >95% at 3db. The sensor appears not to have been used since the last factory visit 2 years before this cruise, but some drift can occur on the shelf.

There was a sub-surface maximum consistent through the region. The values do not suggest that the sub-surface maximum is due to continuing equilibration nor are the lower surface values likely due to shed wake corruption due to the rise from 10db.

There is no history of the use of this sensor since it was recalibrated 2 years previous to this cruise. The temperature and salinity profiles show that the surface waters were well-mixed so dissolved oxygen may be mixed down keeping surface saturation low. There is insufficient information to justify recalibration.

File 2020-017-sbe25-recal.ccf was prepared to subtract 0.169ug/L from the fluorescence channel.

##### Filter

This step was skipped since the fluorescence data did not look spiky.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units.

REORODER was used to get the 2 DO values together.

Saturation rates were calculated and plots show near-surface saturation range from 73% to 99% .

##### HEADER EDIT and final checks of CTD files

Header Edit was used to fix headers, fix formats and channel names and to add some notes concerning processing to the headers:

The standards check was run and no problems were found.

##### Final Checks for all CTD data

A cross-reference listing was produced for the whole cruise. No problems were found.

A header check was run including files from both CTDs. 3 very small negative fluorescence values were found; they were replaced with 0 values in the header and data section. No further problems were found.

A header check was run on the BOT and CHE files and a few header problems were found and fixed in the BOT files.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

Particulars

General – times and dates entered in the SeaBird headers in the SBE25 files are wrong, but the IOS  header entries are correct.

1. Test cast - surface data only. Not processed.

2. Test cast – to 126db. Not processed.

12. Deck pressure 1.1db

16. No event 16.

17. SBE911 cable broke while moving in the LARS. No more rosette casts. An SBE25 cast did not acquire any data because the batteries were drained. No data acquired.

45. File named 43 – fixed.

CRUISE SUMMARY - CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  |  | CTD Calibration Sheet Competed? |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  |  | Yes |  |  |  |  |  |
| 2 | SEABIRD | 25 |  | 334 |  | Yes |  |  | No |  |  |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  | Post Cruise |  |  |  |  |
| Name |  |  | S/N |  | Date |  | Location |  |  | Date |  | Location |  |  |
| Temperature |  |  | 2013 |  | 05Dec2019 |  | Factory |  |  |  |  |  |  |  |
| Conductivity |  |  | 4434 |  | 12Dec2019 |  | Factory |  |  |  |  |  |  |  |
| Secondary Temp. |  |  | 5825 |  | 03Dec2019 |  | Factory |  |  |  |  |  |  |  |
| Secondary Cond. |  |  | 4448 |  | 12Dec2019 |  | Factory |  |  |  |  |  |  |  |
| Transmissometer |  |  | 1005 |  | Dec2018 |  | Factory |  |  |  |  |  |  |  |
| SBE 43 DO sensor |  |  | 0047 |  | 20Nov2018 |  | Factory |  |  |  |  |  |  |  |
| SeaPoint Fluor. |  |  | 2215 |  | 27Nov2018 |  | Factory |  |  |  |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 13Mar2018 |  | Factory |  |  |  |  |  |  |  |
| Altimeter |  |  | 1252 |  | 05Dec2019 |  | Factory |  |  |  |  |  |  |  |
| Calibration Information - 0334 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  |  |  | Post Cruise |  |  |  |
| Name |  | S/N |  |  | Date |  |  | Location |  |  | Date |  | Location |  |
| Temperature |  | 4484 |  |  | 08Nov2018 |  |  | Factory |  |  |  |  |  |  |
| Conductivity |  | 2128 |  |  | 08Nov2018 |  |  | Factory |  |  |  |  |  |  |
| ECO Fluorometer |  | 2215 |  |  | 03Nov2018 |  |  | Factory |  |  |  |  |  |  |
| Pressure |  | 482 |  |  | 05Nov2018 |  |  | Factory |  |  |  |  |  |  |