## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 25 Feb 2025 | Removed original Silicate & Flag:Silicate channels and renamed corrected versions of those channels. TSG channel names/units updated. G.G. |
| 18 February 2021 | Salinity:Bottle data were corrected. CTD Salinity recalibrated. See §24 for steps taken to correct CHE, CTD & Loop files G.G. & S.H. |
| 27 March 2023 | Added HPLC data. J.R. |

## PROCESSING NOTES

Cruise: 2020-028

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdF Water Properties Survey

Chief Scientist: Robert M.

Platform: John P. Tully

Date: 16 June 2020 –  22 June 2020

Processed by: Germaine Gatien

Date of Processing:  7 October 2020 – 5 November 2020

Number of original HEX files: 161 (includes 2 not needed)

Number of HEX files: 73  (1 split cast)		Number of CTD files processed: 72

Number of rosette files:	 27			Number of bottle casts processed: 27

Number of original TSG files: 2			Number of processed TSG files:

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (1185DR & #1883DG), a SBE 43 DO sensor (#3791) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensor (#70613) and an altimeter (#62355).  An SBE pH sensor ($692) was mounted for events #1 to #24.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.121was used for acquisition

The data logging computer WP #102 for casts 1-14 & 87-106 with Seasave version 7.26.7.121 used for acquisition.

The old Tully laptop was used for casts 16-86. with Seasave V 7.23.2 used for acquisition.

The deck unit was a Seabird model 11+ #425.

A Guildline model 8400B Autosal serial # 68572 was used to analyze salinity samples.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in excellent order with comments about problems encountered. Sampling notes provided by the Chief Scientist were very helpful.

The standard procedure was used for all CTD casts was: The rosette was brought to the surface.  Pumps were turned ON.  The rosette was brought down to 10m and kept there for 30 seconds.  Once back at the surface, the data started to be archived, with the rosette at the surface for 30 seconds longer.  Then the cast would start.

Early in the cruise it was noted that the differences between the 2 salinity channels were large and erratic. The primary salinity looked poor; other evidence suggested that the problem was associated with the primary system rather than specifically with the primary conductivity. When the pump was replaced the quality of the data improved greatly. It was  also noted during this cruise that the secondary salinity was higher than the primary while the opposite was true when these sensors were used during 2020-001. Also during 2020-001 the secondary salinity was very noisy. It is possible that the pumps were reversed and that a failing pump accounted for the problems on the earlier cruise as well. It is recommended that a record of which pumps are in use be added to the log book equipment page.

The fluorometer was also affected by the pump failure, so channel Fluorescence:URU:Seapoint was removed from events #1 to 6.

It was noted in the log that fluorescence was very noisy between Events #63 and 82 after which the cable was changed. Comparisons with extracted chlorophyll samples show no difference between this section and the fit for the whole cruise. Comparison with nearby casts are not very useful as most of the noisy ones are in the Haro Strait /Gulf Islands section where variability is very high. Event #88 also looks quite noisy and that was after the cable change so possibly the noise reflects real variability.

Problems were discovered in the pressure readings at the surface, so the pressure offset was adjusted.

There were 2 WetLabs CStar transmissometers in use during this cruise:

Channel Transmissometer refers to sensor #1185DR (650nm - red)

Channel Transmissometer2 refers to sensor #1883DG (530nm - green)

For comparison with other Institute of Ocean Sciences cruises, note that the transmissometer wavelength is 650nm unless otherwise stated.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The pH:SBE:Nominal data should be used with caution; no field calibration data are available.     Calibration is required for each cast to get absolute values, although general trends within a cast are likely real. pH data were removed from 2 casts because the cap had obviously been left on the sensor.

The Thermosalinograph system functioned well. The salinity traces contained very few spikes. Comparisons with CTD and loop samples were too noisy to justify recalibration of salinity. In February 2020 the TSG salinity was found to be within 0.002psu. TSG Temperature appears to be high by ~0.01C, similar to observations in February, but the shallow CTD data are noisy and slight heating may occur due to ship effects. No recalibration was applied. The fluorescence data compared well with the CTD fluorescence except when the latter is <1ug/L. There was a good level of detail in the traces and no sign of upward drift.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as sampling notes from the Chief Scientist.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. The temperature, conductivity and dissolved oxygen sensors had been used only once since the last factory recalibrations. The 2 transmissometers had been recalibrated at IOS fairly recently.

- Conversion of a test file from 2020-028 showed that the 2 salinity channel were closer than during the previous cruise.

Early in the cruise it was noted that the differences between the 2 salinity channels were large and erratic. The primary sensor looked like the problem but other evidence suggested that the problem was associated with the primary system rather than specifically with the primary conductivity. When the pump was replaced the quality of the data improved greatly.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2020-028-ctd.xmlcon.

They were converted to IOS Header format with extension *.IOS.

The IOS files were put through CLEAN to create BOT files.

Due to delays in some analyses, steps were run out of the usual order.

The depths in the headers were adjusted based on the study described in section 11.

Temperature and salinity were plotted for all BOT files to check for outliers. Cast #72 was opened in

CTDEDIT. Primary salinity looks poor around 20m but simple editing won’t help so it was left unchanged.

A preliminary header check was run and a problem was found in the pH channel for one cast. Checks will be made of all casts later.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The file was sorted on sample number.

There was one cast that was split between 2 files - #13 and #14. File 13.BOT was renamed as .BOTa and 14.BOT was renamed 13.BOTb. Ultraedit was used to adjust the Bottle Number channel in file “b”, then JOIN was run to join the 2 files with output 13.bot.

The ADDSAMP file was adjusted and used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2020-001-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2020-028_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2020-028oxy.csv. That file was converted into individual *.OXY files.

At this point note was made of all comments in the DO file that had comments starting with “ALL:”. The rosette sheets were checked to note what other samples were taken from the same bottle so the flags will be applied to all samples from the other csv files that are affected by the DO analysts observation.

EXTRACTED CHLOROPHYLL NOT AVAILABLE AT TIME OF ORIGINAL PREP

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2020-028_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2020-028chl.csv. There were no bottles with comments noted as ALL by the DO analyst. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2020-028_SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 29 to 33 days after collection. The files were simplified and saved as 2020-028sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2020-028_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as 2020-028nuts.csv and converted to individual NUT files.

The flags were changed to 3 for the 1 sample noted by the oxygen analyst as having a problem that would affect all samples.

The next steps will need to be repeated when the CHL data become available.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. No problems were found.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were only 13 salinity bottles from 12 casts, 2 of which were in the top 5m.

One of the surface bottles had an extremely high standard deviation in the CTD Salinity for both sensor pairs and is clearly way out of line with all other bottles. The local vertical salinity gradient was very high so any flushing error would have a large effect. The other surface bottle is in an area with a lower salinity gradient.

When the 2 surface bottles are excluded the primary salinity is low by an average of  0.0460 and the secondary is low by 0.0402psu with standard deviations in both ~0.017psu.

The bottles from events #102 and 105 have much smaller differences than the others and they are both from the western end of Juan de Fuca Strait. That area had, as usual, rougher waters as demonstrated by the descent rate of the CTD, so that the contents of Niskin bottles would be mixed better and tend to be closer to ambient waters. For most casts the descent rate was very steady indicating calm conditions.

The 2 bottles from Juan de Fuca indicate that the primary is low by 0.0127 and 0.0077psu respectively.

The differences between the primary and secondary salinity correspond well to the results shown in section 9.

The results from the cruise that followed, 2020-005, include much deeper sampling and probably better flushing of bottles. The primary salinity was found to be low by 0.0042psu and the secondary high by 0.0018psu. The results from this cruise if all bottles are included except the 2 surface ones show differences lower than those of 2020-005 by ~0.042psu. The Juan de Fuca bottles are closer to the 2020-005 results but still lower than those by ~0.009psu.

There are no outliers that cannot be explained by poor flushing in the presence of significant vertical gradients.

For full details for the COMPARE run see file 2020-028-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was a lot of scatter in the plots of (CTD DO – Bottle CTD) versus CTD DO with the most severe outliers being at the high DO end where most indicate the CTD was reading higher than the bottles.

No reasonably objective method could be found to identify outliers from the full data set in a way that produced a good fit. Based on the results of the salinity comparison a fit was made using only casts #102 and #105 plus one other cast, #72, that had a fairly noisy descent rate and low vertical gradients so that flushing errors should be minimal. When the average of replicates was replaced by a single sample for 1 bottle and 2 outliers with high standard deviations in the CTD DO were excluded the fit was:

CTD DO Corrected = CTD DO * 1.0332 - 0.0207 R2 = 0.71

This compares reasonably well with the fit used to recalibrate 2020-005:

CTD DO Corrected = CTD DO * 1.0308 - 0.0085 R2 = 0.97

If the fit for 2020-028 is forced to have the same offset as 2020-005 the fit is:

CTD DO Corrected = CTD DO * 1.0288 - 0.0085 R2 = 0.70

Cruise 2020-005 had many more low DO samples so the offset is likely to be more reliable than from this cruise. These results confirm that using the results for 2020-005 is appropriate for this data set.

Outliers were examined and none seemed far enough out of line to add quality flags.

Cases flagged 3 or 4 were examined. The only one that was noticeably out of line was sample #269; it was flagged 4 and the bottle value appears to be a little too low.  When the replicates were examined it was found that one was a good fit and one poor. Using 5.258mL/L produces a much better fit while 5.094mL produces a much worse fit. The first value was selected and the flag changed to 2.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The first 2 casts were not included in the comparison because the pump on which the sensor mounted malfunctioned. The traces look very poor with very low values on upcasts with almost no variability. Neither downcast nor upcast values compared well with extracted chlorophyll samples.

When all casts except the first 2 are included the fit of CTD Fluorescence against extracted CHL samples, the fluorescence values are about 61% of CHL. As is typical of these sensors fluorescence values tend to be higher when CHL<1ug/L and lower than CHL at higher CHL values. For this cruise there were few very low CHL values.

Because there were noisy fluorescence traces during events #63-82 a separate fit was done using only those casts and the CTD fluorescence was ~60% of CHL. So the noise does not appear to have affected CTD average values, at least while stopped. Filtering will be applied to CTD fluorescence in the downcast sections in the normal course of processing.

##### Conversion of Full Files from Raw Data

All files were converted using 2020-028-ctd.xmlcon.

The Tau function was selected but not the hysteresis function since there was no deep sampling. Depth was included in the conversion.

A few casts were examined and problems were noted casts #1, 2 and 6 in the primary temperature and conductivity and hence in primary salinity. The fluorescence data looks odd with a large vertical offset between downcast and upcast profiles. The fluorometer was mounted on the primary pump. This problem was discovered at sea, and after event #6 the primary pump was replaced. Thereafter the channel pairs look close. The T and C pairs were reasonably close during downcasts with upcasts very noisy, especially the primary. Fluorescence, PAR, SPAR and Dissolved Oxygen profiles looked normal. Fluorescence has values ~0.11ug/L at 300m. The “Green” transmissometer was generally higher than the “Red”; the profiles had similar shapes. The altimetry looked noisy at the bottom for some casts so header entries should be checked carefully.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both sensors. That setting has worked well for many SBE DO sensors in recent years.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on all casts to find the differences between the pairs of temperature, conductivity and salinity channels. There are no deep casts so comparing to previous and subsequent use is of limited value. The differences were very noisy so these are very rough estimates.  The shaded values come from 2020-001 and 2020-005 and are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2020-001-0025 | 500 | -0.0003 | +0.00060 | +0.0040 | “ |
|  | 1000 | 0 | +0.00060 | +0.0075 | High, V Noisy |
|  | 1900 | -0.0004 | +0.00063 | +0.0080 | “ |
| 2020-028-0016 | 300 | -0.0002 | +0.0004 | +0.0044 | High, V Steady |
| 2020-028-0035 | 290 | -0.0003 | +0.0005 | +0.0053 | High, V Steady |
| 2020-028-0050 | 295 | -0.0004 | +0.0004 | +0.0045 | High, F Steady |
| 2020-005-0075 | 500 | -0.0009 | +0.0004 | +0.0055 | High, Noisy |
| “ | 1000 | -0.0006 | +0.0004 | +0.0055 | “ |
| “ | 1900 | -0.0005 | +0.00046 | +0.0060 | “ |
| 2020-005-0140 | 500 | -0.0003 | +0.0004 | +0.0050 | High, X Noisy |
| “ | 1000 | -0.0002 | +0.00043 | +0.0055 | “ |
| “ | 1900 | -0.0002 | +0.00048 | +0.0060 | “ |
| 2020-005-0218 | 500 | -0.0004 | +0.00042 | +0.0055 | High, Moderate |
|  | 1000 | -0.0005 | +0.00046 | +0.0060 |  |

The salinity difference is lower than in February. From the bottle comparison it is clear that the big change is in the secondary salinity. This may be due to a primary pump problem in February. The differences during 2020-005 are similar with salinity slightly higher but this may be partly due to the data coming from different depths.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check and header check were run.

The only problem in the cross-reference check is that times in the log are generally 5minutes later than in the header files, which is likely due to the computer being turned on somewhat earlier than CTD deployment. In most cases this is not significant, but in a few cases the depth recorded proved to be questionable.

The header check turned up a few potential problems:

- An off-scale fluorescence was seen in 1 file. That will likely disappear in processing.

- Negative pressure was found at the end of one cast but conductivity was very low and the averaged values are positive, with some negative values in the “jittter” around the average.

- pH values were very low for 1 cast. Profile plots were examined for all casts and events #55 and 84 had values ~4. The cap must have been left on, so the pH channel should be removed later for those 2 casts.

Surface check was run and found an average  of 2.0db which is reasonable. The pressure looks good.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header -1 (since the altimetry averages over 2m the value are likely high by ~1m.). Where that number was > 4 or <-4 plots of the altimetry were checked. In 3 cases the altimetry was sufficiently messy that it can not be trusted, so the altimetry header was removed. For most the altimetry looked ok. Then the log entry for the depth was examined and in 10 cases the log entry was found to produce a better Check Value. In one case the altimetry looked fine and the log agreed with the header; for that case a new water depth was calculated based on maximum depth sampled plus altimetry header reading.

See file 2020-028-altimeter-ctd.xlsx.

The same changes were made to the bottle files.

Cruise tracks were plotted and added to the end of this report.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on a few casts to assess what settings are best to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space). The best settings were -0.4records for the primary and -1.2records for the secondary. The same settings were found appropriate for the cruise that followed.

SHIFT was run twice on all SBE911 casts using -0.4 records for the primary and -1.2 records for the secondary conductivity. Salinity was recalculated for both channels.

pH:SBE

During 2020-005 an advance of +30 records brought the SBE:pH sensor data into better alignment with temperature and tests on a few of the current casts shows the same value works reasonably well.

SHIFT was run on channel pH:SBE to advance the records by 30 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned event #14 which was run to collect 1 bottle not included in file #13, so no downcast file is required..

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature, dissolved oxygen and conductivity sensors were used during 2020-001 when primary salinity was found to be low by 0.0035pus and secondary low by 0.0112. Dissolved oxygen was recalibrated using slope/offset=1.0242/-0.0022. Cruise 2020-005 was run immediately after this cruise and the primary salinity was low by 0.0042psu, secondary salinity high by 0.0018psu and a DO correction of 1.0308/-0.0085. So the secondary salinity had changed markedly between February and July.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There were only a few excursions from the climatology. In Saanich Inlet there were high temperatures and low salinity. There were small areas with high temperatures in the northern part of the Strait of Georgia and one case of high salinity  and 1 of low salinity in the southern Strait of Georgia and in Juan de Fuca Strait. All excursions came from near-shore areas where the climatology is not suitable.

The excursions do not suggest any problems with calibration.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The decision on which channels to use is not as obvious as it was during 2020-001. For that cruise the secondary salinity had very noisy data and compared poorly with the CTD. For this cruise the primary channels were bad for the first few casts but ok after that. The comparison with bottles is unreliable. The results from cruise 2020-005 suggests that the primary salinity is low by ~0.0042 and the secondary high by ~0.0018 but small flushing errors are likely which would imply that the primary is closer to bottles than that and secondary further away.

A preliminary examination of a few files in CTDEDIT was done to determine which sensor pair is spikier. The secondary pair looked slightly better and for one cast the primary looked bad near the bottom. Given the problems with the pump at the beginning, the secondary channels were selected for editing and eventual archiving.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files except #18 required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. The only unstable features found were from near the surface in areas where some unstable features are expected and could be real or where it is not clear which data are “good” and which “bad”. No further editing was done.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

There is no indication of a problem with pressure calibration.

The results of cruise 2020-005 will be used to recalibrate salinity and dissolved oxygen.

File 2020-028-SBE911-recal.ccf was prepared to add 0.003psu to channel Salinity:T0:C0 and to subtract 0.003psu from channel Salinity:T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0308 - 0.0085

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity to ensure that corrections were appropriate and they were. When the same outliers were removed as in the original comparison, the both salinity channels were found to be low by an average of 0.004psu. See file 2020-028-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded using the same points as in the original fit, the CTD DO was high by an average of 0.0006mL/L. This shows that the recalibration was done correctly though some error due to incomplete flushing likely remains. See file 2020-028-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.048mL/L but the standard deviation was 0.45mL/L. This falls within expectations since the CTD DO is being compared with water from deeper in the water column. There is no way to estimate the errors in the CTD DO data from this comparison.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

The problem with noisy fluorescence in events #63-82 was explored further. Unfortunately all the affected casts were in the Gulf Islands area where variability can be very high with some casts very well mixed. But casts #82 and #88 which are both towards the north in eastern Juan de Fuca Strait look similar with #88 actually noisier than #82. There is no evidence in the profiles or comparison with extracted CHL samples that the average values are unreliable.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. The only unstable feature that was not very small was investigated and came from an area of active mixing.

Profile plots were examined to see if there any problems. No problems were noted other than the pH problem already noted in files 55 and 84.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

The pH:SBE channel was removed from casts #55 and 84.

The Fluorescence:URU:Seapoint channel was removed from casts #1, 2 and 6.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the comments about processing. Separate header comments were added where pH or Fluorescence channels were removed.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 2 to 3m ranged between ~65% to 135%. The lowest values were in Haro Strait and eastern Juan de Fuca Strait.

Values in the Strait of Georgia were higher towards the north and lower to the south. The highest reading was from Saanich Inlet. These observations look typical of this region in June with low values where there is strong vertical mixing and higher where the near-surface gradients are high.

##### Final Bottle Files

MRGSORT was run to get files in pressure order.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

The pH:SBE channel was removed from cast #84.

The Fluorescence:URU:Seapoint channel was removed from casts #1and 6.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEAD EDIT was run to add comments to the headers. Separate header was used for the casts with no fluorescence channel and a note was added to the one cast with no pH:SBE:Nominal channel.

Data were exported from the CHE files to file 2020-028-bottles-final.xlsx. The entries were compared with the rosette log sheets and no data were missing.

A Header Check was prepared and showed that the bottle DO in mass units were missing. A few steps were rerun and the report was recreated and no further errors were found..

Data were exported from the CHE files to file 2020-028-bottles-final.xlsx. The entries were compared with the rosette log sheets and no sample data were missing.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise. The data were delivered in 2 files.

The IOS SBE TSG45 files have extensions RAW but are in csv format, so the files were opened in EXCEL and combined in a single CSV file. (In opening I use DELIMITED, deselect TAB, select COMMA and OTHER (*).

It is necessary to choose TEXT for the time on the 2nd page of the text import wizard.)

The spreadsheets were adjusted as follows:

- 2 lines of headers were added – channel names and units.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to a separate column. Then a concentration value was calculated in column F using offset -0.044 and scale 15.8 as determined in the most recent recalibration of the fluorometer. Copy/Special Paste was used to save those values and then the voltage channel was removed.

- There was only a 11s gap between the 2 TSG files, so the data were combined in one CSV file.

- Time and Date formats are a problem – when converting from RAW choose TEXT but once opened in EXCEL set Time Format  to HH:MM:SS and save the file again.

- There were no NaN entires.

- The file break column was completed so that new files would be created at the beginning of each day by assigned file names like 20200210-000000. Only the first file has a time other than 000000.

- The flow was off at the beginning of the cruise, so temperature, conductivity, salinity and fluorescence data were replaced with pad vales for all records with flow rate = 0 and for the first few records after the flow began.

The file was then converted to IOS Header format with header info added. This produced 8 IOS files, but an initial plot showed that the last 3 of these files really belong to cruise 2020-005; they will be processed with 2020-005.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. (A record number was also added to enable averaging (for use in comparison to CTD files).  Time zero was set to 31 December 2019 0:00:00.

DERIVED QUANTITIES was run to derive salinity using the lab temperature.

DERIVED QUANTITIES was run a second time to derive sigma-T.

REORDER was run to get the separate DATE and TIME channels before the Julian time to get plots with hours on the axes. Other channels were also reordered.

Time-series plots were produced and the traces look very good, with no spikes in salinity though there is a lot of variability in all channels, but that looks real.

a.) Plots

A track plot was produced and added to the end of this report.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2020-028-ctd-tsg-loop-rosette-comp.xlsx. There were 70 CTD casts that overlapped with TSG records with flow turned on.

The TSG files were averaged over 24 records (2 minutes) on record number to reduce the noise and file size. Standard deviations were included. Then records were extracted for the times of CTDs and added to file 2020-028-ctd-tsg-loop-rosette-comp.xlsx..

Comparisons were made of positions to check for good matches; one error was found and corrected. The differences in positions are expected to be small despite the averaging because the ship was stopped at these times. The average differences were 0.0001º for latitude and 0.0000º  for longitude.

c.) Comparisons

- Comparison of T, S and fluorescence from TSG and CTD data

The initial comparison between TSG and CTD data has enormous variability between events #19 and #69.

The standard deviations in the intake temperature follow a similar pattern. This is most likely due to large vertical temperature gradients in one region and well-mixed surface waters in another. The latter are the most useful for establishing how well the TSG is working as it lessens the importance of matching sampling depth between CTD and TSG.

When all data are included the standard deviations are high in all channels except for fluorescence.

|  | Tint-Tctd | Tlab-Tctd | Stsg-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| min | -2.3622 | -0.4690 | -7.6415 | 0.651 |
| max | 1.9224 | 2.5862 | 1.9844 | 3.104 |
| median | 0.0229 | 0.5577 | -0.0558 | 1.097 |
| average | 0.1033 | 0.7181 | -0.7894 | 1.219 |
| stdev | 0.5273 | 0.5920 | 1.7600 | 0.455 |

When the data were sorted on standard deviation in the intake temperature and only differences included for which std dev < 0.008, then there were 19 casts included and the differences are:

|  | Tint-Tctd | Tlab-Tctd | Stsg-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| min | -0.1018 | 0.2780 | -4.5630 | 0.6506 |
| max | 1.9224 | 2.0362 | 0.2922 | 3.1039 |
| median | 0.0164 | 0.4900 | -0.0115 | 1.0256 |
| average | 0.1466 | 0.5865 | -0.2248 | 1.2573 |
| stdev | 0.4710 | 0.4123 | 1.0565 | 0.6360 |

When the data were sorted on standard deviation in the salinity and only differences included for which std dev < 0.012psu there are 19 casts included and the differences are:

|  | Tint-Tctd | Tlab-Tctd | Stsg-Sctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| min | -0.2700 | 0.2928 | -0.0839 | 0.8248 |
| max | 0.3000 | 1.0400 | 0.1490 | 1.4405 |
| median | 0.0103 | 0.4909 | -0.0137 | 0.9628 |
| average | 0.0080 | 0.5238 | -0.0065 | 1.0455 |
| stdev | 0.1105 | 0.1798 | 0.0527 | 0.1949 |

The results using the low salinity standard deviation standard produced better results for temperature and fluorescence as well as for salinity. Using the 19 best casts for salinity the TSG  intake temperature is higher than the CTD by a median of 0.010Cº and the TSG salinity is low by 0.014psu. The fact that temperature is a little high and salinity a little low may be due to the TSG data in the comparison coming from a little higher than the level from which most of the CTD data,~4m, though the intake is actually at 4.5m. There may also be some slight heating of water from the ship before it enters the loop and small bubbles might account for salinity being slightly low. Whatever the case the differences are not large enough to justify recalibration of the TSG especially given the standard deviations in the comparison.

Heating in the loop is about 0.5Cº from the table above, but most of the data included in that  comparison comes from the latter part of the cruise when the temperatures at 4m were at their lowest. This is expected to be a temperature-dependent variable but there is too much noise in TSG temperature in the Strait of Georgia for a useful comparison. This is likely due to imperfect matching of times between intake and lab measurement in an area with high variability. The values from Juan de Fuca Strait look consistent and similar to results from 2019-005 when intake temperatures were about 10ºC.

The TSG fluorescence compares very well with the CTD fluorescence, especially for values >1ug/L. Below 1ug/L either the CTD is reading a little low or the TSG is reading a little high.

- Comparisons of 3 Loop samples, 1 rosette sample with TSG data

There were only 4 loop samples and 1 had no time recorded. Salinity differences (TSG – Bottle) were +0.0359, -0.0558 and -0.0255. The first of those samples came from a time when the TSG salinity had a very high standard deviation. The other 2 samples show the TSG salinity to be low.

The loop CHL samples had values from 2.8 to 3.9ug/L and, as usual, the fluorometer read lower than CHL, averaging about 50% of CHL. This is based on only 3 samples but the result looks reasonable.

There was only one loop sample that occurred during a rosette cast. The salinity and extracted CHL from the loops and rosette compare very well; the loop salinity was higher than the rosette by 0.005psu and the two chlorophyll samples were virtually identical at 2.82 and 2.78ug/L. For that same cast the TSG salinity read low by 0.02 and the TSG fluorescence was 40% of the Rosette CHL. Bubbles may explain the salinity difference as they would affect the salinometer more than the bottle samples. The fluorescence difference is similar to other observations with these fluorometers. The relationship between fluorescence and extracted chlorophyll is complex varying from one region to another and is also dependent on light levels.

d.) Calibration History

- The TSG and fluorometer were recalibrated in late 2019. The only other use since then was during 2020-001. At that time  the TSG temperature was found to be higher than the CTD by about 0.007C° for casts with well-mixed surface waters. Salinity differences varied according to area; where surface waters were well-mixed the salinity looked slightly high but it was lower than loops samples by 0.002psu. TSG fluorescence traces versus time were similar to CTD traces but there was an offset ~0.4ug/L but most values were low for this cruise.

e.) Conclusions re IOS TSG

1. The TSG clock worked well and position information was available and reliable.

2. There was an opening section with no flow and then a few records with a very high flow rate after flow was turned on. Thereafter the flow rate was mostly very steady with values ~1.1L/min with short sections of variability but records all in the 1 to 1.1 L/min range

3. For the casts with well-mixed surface waters, the TSG temperature is higher than the CTD temperature by about 0.01C° with a standard deviation of 0.008 C°. This is similar to the results when the system was last used.

4. Heating in the loop was ~0.5C° in the areas with low surface temperature variability. This is similar to results for 2019-005 in sections where there were similar intake temperatures.

5. The TSG salinity is remarkably free of spikes. It is lower than the CTD salinity by about 0.014psu but the standard deviation is 0.18psu. The comparison with 3 loop samples shows the TSG salinity to be ±0.05psu with 2 out of 3 samples showing the TSG salinity to be low. There is insufficient data to justify recalibration.

6. The TSG fluorescence is in excellent agreement with the CTD fluorescence except when the CTD FL <1ug/L when it seems to vary little so reads higher than the CTD.

7. No recalibration will be applied.

- The salinity comparison is too noisy to be relied upon and on the previous use salinity was found to be within 0.002psu.

- Temperature appears to be high by ~0.01C°; it is generally the case that the intake temperature is a little higher than CTD temperatures. This may be due to some slight warming at the entrance to the loop. Shallow CTD data in this region are not reliable enough to warrant recalibration of temperature based on this evidence.

- TSG Fluorescence compares well with CTD fluorescence except when the latter is <1ug/L. No recalibration was applied.

g.) Editing

No editing is needed.

h.) Preparing Final Files

REMOVE was used to remove channel Pressure, Temperature:Difference , and record #.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

A cross-reference list was prepared:

Filename                        Latitude      Longitude      Date                    Time

----------------  -----         -------------    ---------------    --------------------    -------

20200618-012447        48 39.54 N  125  8.88 W  UTC 2020/06/18 01:24

20200619-000000        49 20.34 N  125  7.62 W  UTC 2020/06/19 00:00

20200620-000000        48 43.32 N  123 50.94 W  UTC 2020/06/20 00:00

20200621-000000        48 10.98 N  124  6.54 W  UTC 2020/06/21 00:00

20200622-000000        48 19.20 N  125  9.00 W  UTC 2020/06/22 00:00

##### CORRECTIONS

February 2021

Problems were discovered in the software used to derive salinity from conductivity measured on the Autosal. The effects were most notable in the lower salinity range. The corrected data remove the pressure dependence noted in the original processing.

Based on corrected bottle data the primary salinity was found to be high by an average of 0.0026psu and the secondary salinity high by 0.0086psu. In the original processing the corrections applied were -0.003 to the primary and +0.003 to the secondary. To correct this recalibration was applied by subtracting 0.006psu from the secondary salinity. Since the primary salinity was not archived, no recalibration was required for that channel.

Salinity:Bottle data was replaced in CHE files.

Salinity:T1:C1 was corrected in CHE and CTD files

COMPARE was run to see if the corrections were applied correctly and they were, although the differences between CTD and Bottle salinity are large. This was investigated in the original processing and suspicion fell on a combination of poor flushing of Niskin bottles in the presence of large vertical gradients and problems with how the bottles were sealed. Both would lead to the CTD looking lower than it really was. The original comparison was considered unreliable so recalibration was based on the comparison from cruise 2020-005. For post-correction comparison see file 2020-028-sal-comp3.xls.

Particulars  - Notes from Daily Science Log and Sampling Notes

PAR off: None

pH on: All

Casts with bottle fired out of order: None

Casts with no Niskin closed: Many

There are errors in the SeaBird headers re ship name and cruise #. These were not fixed as the final files will have IOS Headers with the correct information.

1. Station name (SI) and depth (175) missing from headers. Fixed.

2. Pause @150 to test salinity alignment.

1-6. Many problems with primary pump. Replaced after cast 6.

1-28. Con file used at sea produced poor pressures so it was updated after #28 with offset = -0.8db.

1-14, 86-106. run with computer 102

16-86. run with old Tully laptop

13/14. Cast 13 interrupted by computer reboot – last bottle in file 14. Need to join bottle files.

34. Niskin 1 closed at surface just to check pressure. No CHE file needed.

62. Niskins 2 and 3 not closed – not needed. Station shallower than planned.

63. Computer switched from local to UTC time.

63-82. Fuzzy Fluorometer signal. Cable changed after cast #82.

74. Wrong depth in header should be 224. Fixed.

104. Wrong depth in header should be 193. Fixed.

105. Wrong depth in header should be 221. Fixed.

2020-028

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 443 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 4700 |  | 11Dec2019 |  | Factory |  |  |  |  |
| Conductivity |  |  | 3531 |  | 12 ec2019 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 4888 |  | 11Dec2019 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 4513 |  | 28Jan2020 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 8Feb2020 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 8Feb2020 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3791 |  | 7Dec2019 |  | Factory |  |  |  |  |
| SBE43 pH |  |  | 692 |  | 14 Jan2020 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SPAR |  |  | 20518 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3640 |  |  |  | Factory |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 9Jan2020 |  | Factory |  |  |  |  |
| Altimeter |  |  | 62355 |  |  |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 21Jan20 | Factory |  |  |
| Conductivity | 0620 | 21Jan20 | Factory |  |  |
| WETStar Fluorometer | 953 | 30Dec2019 | Factory |  |  |