## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Feb 2025 | Removed original Silicate & Flag:Silicate channels and renamed corrected versions of those channels. TSG channel names/units updated. G.G. |
| 27 March 2023 | Added HPLC data. J.R. |

## PROCESSING NOTES

Cruise: 2020-069

Agency: OSD

Location: Strait of Georgia, Johnstone Strait, WCVI, Queen Charlotte Sound, Hecate Strait

Project: Moorings

Chief Scientist: Spear D.

Platform: John P. Tully

Date: 15 July 2020 –  11 August 2020

Processed by: Germaine Gatien

Date of Processing: 13 November 2020 – 27 January 2021

Number of original HEX files: 301

Number of original CTD files: 301 		Number of processed CTD files: 300

Number of rosette files:	 133			Number of processed CHE files: 127

Number of raw TSG files: 5			Number of processed TOB files: 26

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (1185DR & #1883DG), a SBE 43 DO sensor (#3791) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensor (#70613), a SPAR sensor (#20518) and an altimeter (#62355).

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlab/Wetstar fluorometer (WS3S-953P) and flow meter.

Seasave version 7.26.7.121 was used for acquisition.

The data logging computer WP #102.

The deck unit was a Seabird model 11+ #425.

A Guildline model 8400B Autosal serial # 68572 was used to analyze salinity samples.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets were in good order with many comments about problems encountered. There was no list of personnel and no information about the thermosalinograph.

There were a number of cases of the same sample #s being used for 2 different casts. To ensure proper assignment of data to bottle files the analysts added a leading digit to the sample numbers for 4 events. For event #280 (a leading  9), event #290 (a leading 8) and for event #298 (a leading 7) and for event #311 (a leading 9). So, for example, sample 1123 from event #180 became sample 91123.

During this cruise there was a need for sampling undisturbed surface waters for most casts, so the 10m soak was skipped except for a few casts in Grenville Channel and the Strait of Georgia. Such soaks are done to improve the dissolved oxygen sensor data in the top 10 or 15m. For the casts without the soak the temperature and salinity data were likely sufficiently equilibrated, though that is hard to judge where temporal variability is high. Dissolved oxygen values may be a little high in the top 20m due to bubble release, but there is no obvious evidence of that having occurred.

There were samples taken from a small boat with no accompanying CTD data. Those results are not included in this processing job but there is an intent to archive them later. (Cast #319 – Kitlope)

There were 2 WetLabs CStar transmissometers in use during this cruise:

Channel Transmissometer refers to sensor #1185DR (650nm - red)

Channel Transmissometer:Green refers to sensor #1883DG (530nm - green)

For comparison with other Institute of Ocean Sciences cruises, note that the transmissometer wavelength is 650nm unless otherwise stated.

An error was discovered in the conversion of the Transmissometer:Green channel part-way through processing but was corrected using CALIBRATE.

There were a few deep salinity samples but they all came from about 5m off the bottom; that suits the needs of geochemists but is of limited value for checking the CTD salinity calibration. This is particularly problematic as there were problems with the Autosal that affected a number of 2020 cruises, so calibration cannot be based on the previous cruise. There were good Autosal results for the cruise that followed, 2020-008. Because the 2020-008 primary CTD salinity was poor, the secondary CTD channels were selected for this cruise so that recalibration could be based on sampling during the later cruise. It is recommended that samples for calibration of salinity be taken in deep water where vertical gradients are low and at least 10m off the bottom.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples. It is recommended that users check extracted chlorophyll values where available.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.40 mL/L from 0-100db except where DO gradients are very large

±0.20 mL/L from 100db-250db

±0.10 mL/L from 250db-500db

±0.04 mL/L below 500db

The usual averaging interval of 1db was used for most casts, but for events #198 to 219 and #230 to 333 the bins were set to 0.5db to ensure that structure near the surface was resolved.

The Thermosalinograph system functioned well with lots of detail in the traces and no significant spiking in salinity. There was a major problem with fluorescence values drifting upwards until July 30th when the fluorometer was serviced; thereafter it performed well. Comparisons with CTD data suggest that some contamination was lowering the TSG salinity in the latter half of the cruise, but the TSG values were fairly close to the few loop samples collected during the cruise. It is possible that both TSG and loops were affected by suspended sediments in the loop or in the TSG itself. During the cruise that followed the differences between TSG and CTD were close to expected values. More loops samples are recommended covering the full cruise; the few available were all taken in the second half.

The TSG data was divided into daily files with the exception of July 30th for which there are 2 files due to the interruption for fluorometer cleaning.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

There was usually no 10m soak in order to sample undisturbed surface waters. For a few casts there was a soak. Where possible acquisition started with the CTD right at the surface. Initial data usually include data recorded before pumps came on. There was generally a wait of at least a minute before the full cast was run. There was a wait of 30s before bottles were fired.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The histories of the pressure sensor, conductivity and dissolved oxygen sensors were checked. The temperature, conductivity and dissolved oxygen sensors had been used on 3 other cruises since the last factory recalibrations. See section 14 for details. The 2 transmissometers had been recalibrated at IOS fairly recently.

- During 2020-028, using the same equipment it was discovered that the pressure sensor was reading too high. Tests were done and the offset was adjusted in the configuration file for 2020-028 and 2020-005. The corrected configuration will be used for 2020-069.

- Early in cruise 2020-028 a problem was found in the secondary pump so it was replaced. It appears to have worked well during 2020-005.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2020-069-ctd.xmlcon for all other files.

Due to delays in chlorophyll and nutrient analyses, steps were run out of the usual order.

The depths in the headers were adjusted based on the study described in section 11.

The ROS files were converted to IOS format.

The IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Problems were noted in a few casts so they were opened in CTDEDIT:

- Event 93 – Both salinity channels were cleaned lightly at ~50db.

- Event 97 – Channel SalinityT1:C1 was cleaned ~20db.

- Event 171 – The CTD never stopped during the 40db level - Niskin 3 was closed “on the fly”. No editing was applied but nutrient samples #680 should be flagged 4. No other sampling at that level.

- Event 205 – The CTD started moving up immediately after firing so the 10s window caught some “on the fly” data. The data were removed using CTDEDIT that came after the bottle

closed.

- Event 394 – Channel Salinity:T0:C0 was cleaned lightly at the bottom of the cast.

The output files were copied to *.BOT.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number.

The output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

There were a number of cases of sample numbers being repeated. A decision was made to change all sample numbers in the affected casts by adding a leading 7, 8 or 9. Following the information given in the salinity analysis notes the following changes were made:

Event 280 at Kit1: Sample #s changed to 91118-91131

Event 290 at DOUG4: Sample #s changed to 81118-81123

Event #298 at FOC1: Sample #s changed to 71131-71135

The file was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2020-001-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2020-069_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2020-069oxy.csv. That file was converted into individual *.OXY files.

At this point note was made of all comments in the DO file that had comments starting with “ALL:”. The rosette sheets were checked to note what other samples were taken from the same bottle so the flags will be applied to all samples from the other csv files that are affected by the DO analysts observation.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2020-069_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2020-069chl.csv. Flags were changed to 3 for samples 329, 516, 986 and 1254 due to problems noted as applying to all by the DO analyst. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2020-069_SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 13 to 25 days after collection. The files were simplified and saved as 2020-069sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2020-069_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as 2020-069nuts.csv and converted to individual NUT files.

Flags were changed to 3 for samples 120, 329, 516, 986, 1056 and 1254 due to problems noted as applying to all variable by the DO analyst.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few problems were found:

- First cases where the DO analyst indicated a flag should be applied to all samples where checked to ensure that had been done. The only missing cases were for the nutrient samples #1055 for Event 231. The flag 3 had been added to the wrong samples. That was fixed.

- Event 34 - Nutrient samples #59 shown on rosette sheet but missing and the analysis sheets show no sign that the sample was received. These samples were entered with flag 9.

- Event 207 –The nutrients were mislabelled as being from event #208.

- Event 280 – Chlorophyll Sample 1130 was renamed 91130. Sample 1123 not indicated on rosette sheet and unlikely a sample would be taken at that depth.

- Event 290 – Chlorophyll Sample 1123 renamed 81123 and reassigned from 280 to 290.

- Event 284 – The rosette entry was unclear and the chlorophyll sample had been entered as event 234 and the Salinity sample as 294. Both were reassigned to event 284.

- Event 309 – All samples were named correctly but the CTD file was named 310. The CTD file was corrected.

- Event 396 at station 6 – Chlorophyll entered with event #6. Corrected.

After these corrections the merge process was repeated for the affected casts.

##### Compare

Salinity

Compare was run with pressure as reference channel.

All salinity bottles below 250db were fired about 5m off the bottom of casts, so any flushing errors will be of the opposite sign to the usual and possibly larger due to being so close to the bottom which might lead to reflection of shed wakes.

The primary CTD salinity was very close to bottles and the secondary higher by about 0.006psu. The difference between the channels is close to what has been seen in previous cruises and is line with the differences seen between the 2 CTD channels. Both channels are higher relative to bottles than in previous cruises. This is unlikely to be due to calibration drift since both changed by the same amount. It is more likely due to firing close to the bottom. The results of cruise 2020-008 that followed 2020-069 suggest that that the secondary salinity was high by ~0.002psu. (The primary salinity from 2020-008 looks unreliable as it was drifting.) Flushing errors from 2020-008 are likely smaller due there being very deep samples that were well away from the bottom.

It is known that there was a problem with the analysis of bottles for several recent cruises, 2020-028, 2020-005, 2020-017, 2020-031 and 2020-083 but it did not affect this cruise.

The Autosal error did not occur during 2020-008 when there was good deep sampling. At that time the primary salinity was found to be drifting and secondary salinity was high by about 0.0018psu. During 2020-083 the primary salinity looked much better and it is thought that thorough cleaning and reconnecting cables might have led to improved data. The secondary salinity was found to be high by about 0.0018psu. A preliminary look at cruise 2020-083 suggests the primary was low by ~0.004 and the secondary high by ~0.002psu. The difference of 0.006 between channels is consistent.

Subtracting 0.0018psu from the secondary salinity looks appropriate for this cruise.

Almost all outliers came from the portions of the cruise in inlets and close to shore. Below 350m there was only 1 outlier (event 42) and it affected the primary sensor only so was clearly a problem with the CTD data. The differences between sensors are frequently very large at the surface but the only case with a large difference below 50m was for event #42 and the problem appears to have affected the whole cast.

Outliers that came from 300db or lower were investigated. In all but one case the outliers were from bottles fired at the bottom and the sign of the error suggests the Niskin bottles contain water from above:

Cast #42 – As reported above the primary salinity looks bad.

Cast #212 – Bottle fired at bottom; bottle value similar to data collected as shed wake passed through from above.

Cast #293 – In a narrow inlet. Minor outlier. Bottle fired at bottom; bottle value similar to data collected as shed wake passed through from above.

Cast #314 – In a narrow inlet. Bottle fired at bottom. Likely poor flushing.

Cast #329 – Hecate Strait – standard deviation in the CTD data high.

For full details for the COMPARE run see file 2020-069-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

This cruise included 5 deep offshore casts and many casts in inlets and/or shallow and close to shore.

Plots were made of differences between bottles and CTD DO versus CTD DO.

Choosing all casts and removing outliers based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0376 - 0.0215 R2 = 0.92	(1)

Selecting just the 6 deep casts and removing 3 outliers the fit was:

CTD DO Corrected = CTD DO * 1.0369 - 0.0197 R2 = 0.94	(2)

These 2 fits were very close. When the data in the first fit were plotted against file pair # it showed that most of the outliers were in the inlet sections of the cruise.

So a further fit was made of just the inlet section of the cruise. There was a lot of scatter, so identifying outliers was highly subjective. Fits included:

CTD DO Corrected = CTD DO * 1.0241 + 0.013 R2 = 0.43	(3)

And when enough outliers were rejected to get a good R2 value the fit was:

CTD DO Corrected = CTD DO * 1.0233 + 0.0181 R2 = 0.85	(4)

but this removed half the records. There was only 1 bottle with DO <2.5 mL/L and none <1, so the offset is not reliable. If the offset is set to the offshore value, the fit looks poor and that is reflected in an R2 of 0.58.

These results confirm the expectation that flushing of bottles will be poor in inlets where descent rates of CTD are very steady so the Niskin contents don’t get stirred up. Since dissolved oxygen values generally decrease with depth, incomplete flushing leads to bottles reading closer to the CTD DO than we would expect since the CTD generally reads lower than ambient values. We simply cannot get a good fit for the inshore as flushing errors will be non-linear changing with local vertical gradient in DO.

Fit #2 looks like the best fit for accurate CTD values. The slope and offset are higher than when the sensor was last used but some drift is expected and an initial look at cruise 2020-008 (which followed this cruise) had a higher slope.

The fit used for 2020-001 was

CTD DO Corrected = CTD DO * 1.0242 - 0.0022

And for 2020-005 it was:

CTD DO Corrected = CTD DO * 1.0308 - 0.0085

And for 2020-008 it was:

CTD DO Corrected = CTD DO * 1.0449 - 0.0298

Outliers were investigated. Most came from near the surface. Those from inlets generally had bottle values that look low and this is likely due to incomplete flushing of Niskin bottles. There was one 1 inlet cast with a bottle value higher than the CTD but there was very high variability in the CTD data with some values matching the bottle and others lower. There were 3 outliers from the Strait of Georgia that also that had bottle values that are higher than the CTD; in each case these were near the surface and the differences were easily explained by minor errors due to incomplete flushing. There were 2 outliers of the west coast of Vancouver Island but close to shore; these were both near the surface in high gradients so the differences do not look significant. None of the outliers look likely to be due to analysis or sampling problems so no flag changes were suggested.

Plots of DO versus salinity suggest a high value at cast 57 at 5m but examination of the CTD data shows very high variability in DO during the bottle stop, so it is quite likely the bottle contents really did have a high DO concentration.

For full details for the COMPARE run see file 2020-069-dox-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The fit of CTD Fluorescence against Extracted CHL samples has the usual pattern with Fluorescence reading higher at low CHL and falling relative to CHL as the latter values increase. The CTD fluorescence is 0.66 times the CHL values when a fit is forced through the origin. The performance is typical of this type of sensor.

The dark values for this fluorometer are gradaully increasing with minimum fluorescence being 0.07ug/L in February, 0.08ug/L in June and 0.09ug/L for this cruise. They are not high enough to have a significant effect on the data but should be monitored.

##### Conversion of Full Files from Raw Data

All files were converted using 2020-069-ctd.xmlcon. There was a problem in the headers of  file #35, but when reconverted it was fine.

The Tau function and the hysteresis function were selected since there was deep sampling. Depth was included in the conversion.

A few casts were examined and all expected channels are present. The T and C pairs were reasonably close during downcasts with upcasts very noisy. Fluorescence, PAR, SPAR and Dissolved Oxygen profiles looked normal. Fluorescence has deep values <0.1ug/L. The “Green” transmissometer was generally higher than the “Red”; the profiles had similar shapes. The altimetry looked noisy near the bottom but got close enough to allow reasonable estimates if the header estimates are found to be poor.

The PAR sensor was removed for casts

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined after this step; using +2.5s certainly improves the alignment. That setting has worked well for many SBE DO sensors in recent years. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to find the differences between the pairs of temperature, conductivity and salinity channels. There were no deep casts for the 2nd half of the cruise so 1 shallow cast is included. The differences were noisy except for the last cast so these are very rough estimates.  Some deep casts from 2020-001 and 2020-005 are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2020-001-0025 | 500 | -0.0003 | +0.00060 | +0.0040 | “High, V Noisy |
|  | 1000 | 0 | +0.00060 | +0.0075 | “ |
|  | 1900 | -0.0004 | +0.00063 | +0.0080 | “ |
| 2020-005-0075 | 500 | -0.0009 | +0.0004 | +0.0055 | High, Noisy |
| “ | 1000 | -0.0006 | +0.0004 | +0.0055 | “ |
| “ | 1900 | -0.0005 | +0.00046 | +0.0060 | “ |
| 2020-005-0140 | 500 | -0.0003 | +0.0004 | +0.0050 | High, X Noisy |
| “ | 1000 | -0.0002 | +0.00043 | +0.0055 | “ |
| “ | 1900 | -0.0002 | +0.00048 | +0.0060 | “ |
| 2020-005-0218 | 500 | -0.0004 | +0.00042 | +0.0055 | High, Moderate |
|  | 1000 | -0.0005 | +0.00046 | +0.0060 |  |
| 2020-069-0043 | 400 | -0.0005N | +0.00046 | +0.0050 | High, Mod |
|  | 1000 | -0.0006N | +0.00040N | +0.0054 | “ |
|  | 1900 | -00002 | +0.00046 | +0.0058 | High, Noisy |
| 2020-069-0068 | 400 | -0.0005N | +0.00040 N | +0.0053 | High, Mod |
|  | 1000 | -0.0004 | +0.00044 | +0.0057 | High, Noisy |
|  | 1900 | -0.0005N | +0.00050N | +0.0068 | “ |
| 2020-069-0122 | 400 | -0.0003 | +0.00050 | +0.0060 | High, Mod |
|  | 1000 | -0.0002 | +0.00051 | +0.0064 | “ |
|  | 1900 | -0.0002 | +0.00053 | +0.0068 | “ |
|  | 2900 | -0.0003 | +0.00055 | +0.0072 | “ |
| 2020-069-0401 | 400 | -0.0002 | +0.00060 | +0.0062 | High, V Steady |

In the early part of the cruise the differences are very similar to those seen during the previous cruise. Later there is a slight increase in the conductivity and salinity differences but by no more than 0.001psu. Given the rough nature of this comparison the change is too small to imply calibration drift without supporting evidence from the bottle comparison.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert Sea-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The cross-reference check was run and turned up no errors.

The header check were run and checks were made against the logs.

Changes were made to station names to maintain consistent formats and to match the logs in the following CLN files and affected SAM files: 1, 49, 112, 138, 203, 215, 216, 219, 301, 302, 305, 374, 379, 386 and 398.

A few checks of pressure calibration were made:

- Surface check was run and found an average  of 0.25db, which is much lower than usual but many casts were in inland waters where surface measurements were sought. None of the casts had sufficiently well-mixed surface waters to be useful in comparisons to bottles.

- During event #59 the CTD appears to have passed through the surface about-0.1db towards the end of the soak but pumps were off.

- Cast #210 – long time at surface with pumps off – conductivity fell at  -0.03db then rose when pumps came on. Transmissivity fell to ~0 at pressure +0db.

- Two deck measurements were recorded, +0.2db and +0.3db but after a few scans values dropped to -0.1db to -0.2db as the pressure sensor likely equilibrated.

- A few casts were found where the CTD went through the surface at the end. The pumps were off so there is no way to judge where the surface was from pumped channels but the Green Transmissivity shifted very suddenly from ~0% at 0db to high values at -0.2db or -0.3db.

So it looks like the pressure is within ±0.02db. No recalibration is required.

The altimeter and water depth readings from the headers of the CLN files were exported to spreadsheets. A check value was calculated as:

= Max Depth Sampled – Header Water Depth + Hdr Altimetry - 1

The -1 is because the altimetry algorithm averages over 2m so the value may be high by ~1m.

Where that number was > 4 or <-4 plots of the altimetry and header water depth entries were checked.

Changes were made to headers as follows:

- If the header depth differed from the log entry and the latter led to a lower Check Value, then the header entry was changed.

- If the altimetry header value is clearly wrong, it was removed or replaced with an estimate based on examination of a plot.

- If difference is small (<8m) and no obvious error, likely just variations through cast, no changes were made.

- If none of the above applied a depth estimate was made based on altimetry plus maximum depth sampled and that was used for the header entry.

There were many cases where the depth was obviously changing through a cast or possibly in narrow channels the sounder was not giving good values.

The changes were made to the SAM files which were then bin-averaged again and to the CLN files.

Cruise tracks were plotted and added to the end of this report.

A study was made of 5 casts that had a sub-surface soak to see what could be learned from them:

- Events #225 and 226 were cases where the pumps had not been turned on so there was a return to the surface to start over.

- Event #289 just had a brief dip to 3m.

- Event #234 had pumps turned on at the surface and the soak period is recorded so we can study how the two downcasts sections compare and look at entrainment on the upcast.

  - The salinity settled quickly during the surface soak with variability of ±0.005psu. During the 10m soak it settled after about 30s with variability of ±0.02psu. When the CTD returned to 10m during the full downcast salinity was slightly lower than the minimum measured.

  - During the surface soak DO quickly rose to ~7.7mL/L, fell to about 5.3 during the drop to 10m. A shed wake passed through made obvious by the salinity signal. DO rose slightly due to the wake, fell during the 10m soak from about 5.3 to 4.9. There is no way to separate equilibration of DO to lower values and DO rising due to bubble release. When the CTD returned to 10m during the full cast the DO was ~4.8. We might expect it to be slightly high due to slow response. So being low may imply that bubbles did get released earlier and DO might have been too high if the soak had not occurred.

- Event #235 also had the pumps turned on at the surface and the soak period is recorded so we can study how the two downcasts sections compare and look at entrainment on the upcast. After a large shed wake passed through DO decreased and salinity increased and then reversed. There is too much variability to see evidence of bubble release or equilibration of sensors.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on a few casts to assess what settings are best to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space). The best settings were -0.4 records for the primary though the differences were slight and -1.4 records for the secondary.

SHIFT was run twice on all SBE911 casts using -0.4 records for the primary and -1.4 records for the secondary conductivity. Salinity was recalculated for both channels.

##### DELETE

For all casts acquisition was started immediately with pumps off. For most there was no 10m soak. There was often a drop to between 1 and 2m with pumps off then the CTD was raised a little and pumps were turned on. DELETE is likely to pick data with pumps off or remove some good data if these initial records are not removed first.

Plots were made of all casts to determine when pumps came on. In general, clipping casts based on when the pumps came on will work appropriately, but sometimes the CTD rose a little after pumps came on. For a few casts many good records would be removed because there was a 10m soak. So for some casts a record was kept of how many records should be removed in order to get as shallow data as possible.

CLIP was first run on all casts using Pumps On as the criteria.

Then CLIP was run individually on those casts noted as needing individual settings based on scan #: 151, 185, 191, 225, 234, 235, 268 and 289.

DELETE was then run.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <  0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was only one warning, concerning cast #288. Pumps came on and off several times so most of the data had been lost. CLIP was rerun based on # of records to remove and DELETE was rerun on that cast.

All DEL files were copied to *.EDT.

##### Other Comparisons

Experience with these sensors since last factory service –

The pressure, temperature, dissolved oxygen and conductivity sensors were used a lot during 2020

- 2020-001 – Primary salinity was low by 0.0035 and secondary performed poorly. Dissolved oxygen was recalibrated using slope/offset =1.0242/-0.0022

- 2020-028 – Shallow sampling not trusted for recalibration. Autosal problem discovered after initial processing but result was not trusted anyway. Used 2020-005 result. Dissolved oxygen was recalibrated using slope/offset=1.0308/-0.0085. Pressure was recalibrated during 2020-028 by increasing the offset to -0.8db. The same offset was used for cruises that followed.

- 2020-005 – Autosal problem discovered after processing but error very small for deep bottles used in comparison. Primary low by 0.004 and secondary high by 0.0018. Made small allowance for flushing error in recalibration. Dissolved oxygen was recalibrated using slope/offset=1.0308/-0.0085.

The following later cruises used this equipment and some information is available about the sensors.

- 2020-008 – Primary salinity poor. Secondary high by 0.0018psu. Preliminary Dissolved oxygen  recalibration found slope/offset=1.0449/-0.0298.

- 2020-009 – No results yet.

- 2020-083 – Autosal problem – based on tentative correction of Autosal results the primary was low by 0.004 and the secondary high by 0.0018psu. So similar to 2020-008.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data from the offshore area and Georgia Strait fell within the climatology as well as casts in Queen Charlotte Sound, Hecate Strait and Dixon Entrance that were well away from shore. There were many cases of salinity looking low and temperature high, especially near the surface and bottom. but all came from inlets or in shallow water close to shore. No climatology was available for many of the inlets. None of these excursions suggest calibration drift or instrumental problems.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

For the past 2 cruises using these sensors the primary channels were selected for editing, and eventual archiving, but for this cruise the primary channels produced poor results for 3 casts. In general, the secondary channels look a little more stable in T-S space, so they were selected for editing.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. Most files required some editing. Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing, T-S plots were examined for all casts. Some unstable features remain but come from areas where such features are expected and likely real or where it is impossible to say which data are poor. A little further editing was applied to cast #25 to remove obviously bad surface records.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

There is no indication of a problem with pressure calibration.

File 2020-069-SBE911-recal.ccf was prepared to recalibrate by subtracting 0.0018psu from channel Salinity :T1:C1 and to apply the following correction to channel Oxygen:Dissolved:

CTD DO Corrected = CTD DO * 1.0369 - 0.0197

Transmisssivity:Green was also recalibrated to correct an error in the parameters used in conversion from raw:

Transmissivity2(Corrected) = 1.44*Transmissivity2 -0.02

This correction was first applied to the SAM and MRGCLN2 files.

COMPARE was rerun for salinity. When the same outliers were removed as in the original comparison, the secondary salinity channel was found to be high by an average of 0.0043psu. That difference is believed to be due to flushing errors in the bottles. See file 2020-069-sal-comp2.xls for details.

COMPARE was rerun for dissolved oxygen and shows that the correction was applied properly. When data are excluded using the same points as in the original fit for the 3 deep casts, the CTD DO was low by an average of <0.00001mL/L. For all casts the average was -0.0007mL/L.

This shows that the recalibration was applied corrected and was appropriate.

See file 2020-069-DO-comp2.xls for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

The CTD DO was higher than the titrated samples by an average of ~0.048mL/L when data from the top 12m were excluded. The CTD DO was high by an average of <0.01mL/L when data from 12db to 75db plus a few outliers were excluded. The DO gradient tended to be largest between 12 and 75db so the CTD DO may be high due to slow response and the bottle values may be low due to incomplete flushing.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to most of the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

For events #198 to 219 and #230 to 333 the bins were set to 0.5db to ensure that structure near the surface was resolved.

Profile plots were examined. Dissolved oxygen had an odd shape at station A1 but looking at a cast from the same site during 2020-008 has a similar look. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

For casts #39-152 and 328-412 channel PAR was also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the comments about processing.

The Standards Check routine was run and no problems were found.

The Header Check was run. PAR:Reference was missing so the steps in this § were repeated; no further problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 2 to 3m ranged between ~70% to 140%, but the deeper offshore casts had values from 104% to 106%.  The values are a little higher than usually seen offshore possibly due to the lack of a 10m soak though bubble release is more likely to occur lower in the water column.

##### Final Bottle Files

SORT was run to arrange casts in pressure order.

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

EDIT HEADERS was run to fix formats and channel names and to add comments about analyses and CTD processing.

Checks were made of comments from the log to see if adjustments were needed to CHE files:

- One line was removed from cast #184 as there was no sampling.

- Cast #126 was missing.

- Cast #396 had 2 bottles fired at one level only one sampled, so one removed from file.

Data were exported from the CHE files to file 2020-069-bottles-final.xlsx. The entries were compared with the rosette log sheets. A few casts had been lost along the way, so they were put through the missing steps and compared again. Flags were added to the nutrients at 40m of cast #171 because the bottle was fired on the fly. No further problems were found.

Standards check and a header check were run. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise. The data were delivered in 5 files, 2 very short.

The IOS SBE TSG45 files were opened in EXCEL.

The files have extensions RAW but are in csv format, so the files were opened in EXCEL and combined in a single CSV file.

Note: In opening such files use DELIMITED, deselect TAB, select COMMA and OTHER (*).

Choose TEXT for time on the 2nd page of the text import wizard and set the DATE format.)

Two unneeded columns were removed after the files were merged.

The file was saved as 2020-069-tsg.csv.

The spreadsheets were adjusted as follows:

- 2 lines of headers were added – channel names and units.

- A column with pressure was added with all values set to 4.5 (to enable derivation of salinity).

- A temperature difference column was added (Lab-Intake).

- The fluorescence channel is in volts. It was moved to column M. Then a concentration value was calculated in column F using offset -0.044 and scale 15.8 as determined in the most recent recalibration of the fluorometer. Copy/Special Paste was used to save those values and then the voltage channel was removed.

- Flow was not on for the first few hours so temperature, conductivity and fluorescence were padded until flow had been on for 1 minute.

- There are short gaps at 15:25 on July 24 (~4 minutes), 15:31 on July 24 (~15seconds) and at 19:09 on July 30 (~21 minutes).

- A file break column was filled with the cruise #-date/time info from the original file name. When the date changed a new file break entry was created so there will be a single file for each day. The only exception was to split July 20 into 2 files given the 21 minute interruption. The time is set to 000000 where a file starts at the beginning of a day and from the first time in the file otherwise. The file break entries were manually adjusted (using fill) as needed.

- Time and Date formats are a problem – when converting from RAW choose TEXT but once opened in EXCEL set Time Format  to HH:MM:SS and save the file again.

- There were no NaN entries.

- The flow was turned off about 30minutes before the end of the file and for a short period before that. The last file was edited to remove data with zero flow, but positions and times were left in place. The steps were then repeated.

The file was then converted to IOS Header format with header info added. There are 22 IOS files.

CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

ADD TIME CHANNEL was used to add Julian dates – i.e. Decimal Year. (A record number was also added to enable averaging (for use in comparison to CTD files). Time zero was set to 31 December 2019 0:00:00.

DERIVED QUANTITIES was run twice, first to derive salinity using the lab temperature and again to derive sigma-T.

REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end.

Time-series plots were produced and the only problem noted was a suspicious steady increase in fluorescence from July 20 at about 0600hours to July 29 and then on July values include pad values and negative values until the TSG was stopped at 19:09. At 19:30 a new file was started after a 21minute break and the fluorescence looked fine thereafter. Steady rises in TSG fluorescence were seen frequently in 2019 but was not seen during 2020-001, 2020-028 and 2020-005. Lindsey Mazzei reported that the fluorometer was removed on July 30th, disassembled, cleaned and reinstalled.

The fluorescence channel should be removed from all files up to that servicing.

The first file that has good fluorescence is file 20200730-193021

a.) Plots

A track plot was produced and added to the end of this report.

b.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet which was saved as 2020-069-ctd-tsg-loop-rosette-comp.xlsx. All CTD casts overlapped with TSG records with flow turned on but CTD data were not available at the appropriate depth for a few casts.

The TSG files were averaged over 24 records (2 minutes) on record number to reduce the noise and file size. Standard deviations were included. Then records were extracted for the times of CTDs and added to file 2020-069-ctd-tsg-loop-rosette-comp.xlsx..

Comparisons were made of positions to check for good matches. The differences in positions are expected to be small despite the averaging because the ship was stopped at these times. The average differences were 0.0000º for both latitude and longitude. There were no differences >0.0007 º.

The flow rate was high (~2.4) up to event #103. After that it ranged from 1.2 to 1.9 but values were  mostly ~1.5. A plot of heating in the loop versus flow rate shows the least heating associated with high flow rates but that could also be related to intake temperatures and local gradients. The greatest variability in heating occurs in the Vancouver Island inlets. The flow rates were mostly ~1.5 there, but also local gradients are often high so that exact time matches are more important. A plot of (Tlab-Tintake) versus standard deviation in the TSG intake temperature confirms that the outliers in the differences are mostly associated with high standard deviations.

It has been noted in other cruises that high flow rates were associated with bubbles in the salinity but that does not appear to be an issue for this thermosalinograph.

c.) Comparisons

The differences between TSG and CTD temperature and salinity had very high standard deviations. A plot of salinity differences versus event numbers indicate that this is primarily a geographic issue. For the offshore and Hecate Strait the differences show little variations. The inlet sampling has extremely noisy differences which is likely due to high gradients near the surface so small mismatches in depth are significant. Slight time mismatches may also be significant in inlets where local variability is high. In the deeper casts of Queen Charlotte Sound there is moderate variability.

- Comparison of T, S and fluorescence from TSG and CTD data

Comparisons between TSG and CTD data were first made using all casts. The standard deviations are high in temperature and salinity differences and extremely high for fluorescence.

|  | TSG Intake TEMP –CTD TEMP | TSG SAL – CTD SAL | TSG FL/ CTD FL |
| --- | --- | --- | --- |
| min | -2.3480 | -12.8802 | -70.56 |
| max | 3.1433 | 4.6682 | 104.62 |
| average | 0.3111 | -1.2317 | 9.61 |
| median | 0.1861 | -0.3051 | 0.74 |
| stdev | 0.6090 | 2.2030 | 21.57 |

The plot of salinity differences against event number shows that the Offshore section plus Hecate Strait has fairly consistent differences. The casts in Queen Charlotte Sound section had more scatter in differences than those in the offshore but a lot less than seen in the inlet casts. So the Offshore/Hecate casts and the deeper casts in Queen Charlotte Sound were studied.

| Offshore Events 5-151 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | TSG Intake TEMP –CTD TEMP |  | TSG SAL – CTD SAL | TSG FL/ CTD FL |
|  | min | -0.4919 |  | -1.2273 | 0.56 |
|  | max | 0.7955 |  | 0.7726 | 104.62 |
|  | average | 0.1780 |  | -0.0655 | 22.27 |
|  | median | 0.1339 |  | -0.0611 | 18.10 |
|  | stdev | 0.2033 |  | 0.1543 | 20.25 |
| QCS Events 337, 339, 341-349, 354-361 |  |  |  |  |  |
|  |  | TSG Intake TEMP –CTD TEMP |  | TSG SAL – CTD SAL | TSG FL/ CTD FL |
|  | min | -0.1138 |  | -0.4913 | 0.37 |
|  | max | 0.4765 |  | -0.2403 | 0.92 |
|  | average | 0.0863 |  | -0.3219 | 0.57 |
|  | median | 0.0166 |  | -0.3086 | 0.59 |
|  | stdev | 0.1618 |  | 0.0613 | 0.14 |

The temperature and salinity differences are usually lower in the offshore area because the surface waters are generally better mixed minimizing the error caused by small vertical and temporal offsets in the TSG and CTD data chosen for the comparison. However, the standard deviations are large due to many of the casts coming from near-shore. Using all the offshore data the TSG temperature looks high by a median of 0.134Cº (standard deviation 0.203) but when only the 5 deepest casts are included the TSG temperature was high by a median of 0.105Cº with a standard deviation of 0.073Cº. The QCS temperature differences are smaller, though the standard deviation is very high; there are fewer casts and more noise in that area, so this result is likely not significant.

Offshore the TSG salinity is lower than the CTD salinity by a median of 0.061psu and standard deviation of 0.15pus. Using only the 5 deepest casts it was low by a median of 0.059psu with a standard deviation of 0.023psu. In QCS it is lower by 0.31psu with a standard deviation of 0.06psu.. If we choose the 38 offshore casts with the lowest standard deviation in salinity the median is 0.062psu and the standard deviation is 0.024psu. So the offshore salinity differences look similar to other recent observations. There were no offshore loop samples. The inshore observations show the TSG salinity to be consistently low with a median of -0.309psu and standard deviation of 0.061psu and the loop samples support that. The final sample in the Strait of Georgia has a slightly smaller difference but the bottle value is flagged due to a loose screw-top which could lead to higher bottle salinity due to evaporation. Given the many casts in narrow channels it is possible that suspended sediments may have gotten into the TSG and loop sample bottles.

A quick check was made of the offshore cruise that followed, 2020-008, to see if the differences between TSG salinity and loop salinity samples continued to be large. Four loop samples were compared with unprocessed TSG  data and the differences found were ~0.1, 0.1, 0.06 and 0.006psu. So the large differences are not seen.

Variations in the flow rate do not seem to be relevant. The smallest differences are associated with a high flow rate but they are also from the offshore. The lowest flow rates are also associated with small differences though there are few such casts so this is not significant. Flow rates in general were higher than we usually see for this TSG system.

Heating in the loop goes up in the inlet section but that is likely because the flow rate is lower and temperatures are generally a little lower there. Bubbles can lead to low TSG salinity but there is no evidence of this in the TSG time-series plots.

One possibility is that sediements were affecting the TSG and loop samples. Looking at transmissivity there does appear to be a correspondence between very low transmissivity and large salinity differences but since both transmissivity and large vertical salinity gradients tend to be high in inlets this may not imply that sediements are affecting salinity values.

TSG fluorescence data are only reliable for the latter half of the cruise. The TSG fluorescence was about 60% of the CTD fluorescence; this is lower than during the previous cruise. But looking at the plot below it is clear that it is much lower in the inlet section but close to the CTD in QCS and the Strait of Georgia. Even in the inlet section it is much lower when CTD fluorescence is high. The gradients in the inlets are highly variable but the peak is often close to 5m. The TSG may draw water from a little higher in the water column where values may be lower. While this may be part of the explanation there is also a tendency of this type of fluorometer to read low when CHL is high and perhaps that is a bigger problem for the TSG fluorometer than for the one on the CTD. We have no rosette chlorophyll from the section in Queen Charlotte Strait. In the inlets the TSG fluorescence follows the usual pattern of reading high when CHL is low and vice versa, but overall reads lower than we would normally expect. The CHL samples are likely from deeper water than the TSG reading so this may well be due to offsets in sampling levels. For the 6 loop samples available the TSG fluorescence ranged from 2.1 times extracted CHL when the latter was low (0.58ug/L) and 0.33 times CHL when CHL was high (8.2ug/L) and 2.2 times CHL when the loop was 1ug/L. Again there are too few points of comparison to be useful. The close correspondence between the 2 fluorometers between events #342 and #412 does suggest that the TSG fluorometer was performing well. While TSG values were generally rising during that period the trace follows the pattern in the CTD fluorescence quite closely.

For more details see file 2020-069-tsg-ctd-loops-comp.txt

d.) Calibration History

- The TSG and fluorometer were recalibrated shortly before cruise 2020-001.

- During 2020-001 which mostly sampled offshore and with mostly well-mixed surface waters, the TSG salinity was found to be lower than that from the CTD and loops by ~0.002psu. TSG temperature was higher than the CTD by ~0.01Cº. Salinity and fluorescence traces looked good.

- During cruise 2020-028, which mostly sampled inshore, the TSG salinity was found to be lower than that from the CTD by ~0.015psu but there was a high standard deviation. TSG temperature was higher than the CTD by ~0.01Cº. Salinity and fluorescence traces looked excellent.

- During cruise 2020-005, which mostly sampled offshore, the TSG salinity was found to be lower than that from the CTD by ~0.014psu and lower than loop samples by 0.05psu. The loops were higher than rosette samples despite the fact that the latter likely came from lower in the water column. TSG temperature was higher than the CTD by ~0.03Cº though the standard deviation was high. Salinity and fluorescence traces looked excellent. TSG Fluorescence tracked the CTD trace quite well though values at the low end of the range generally looked higher than those measured by the CTD sensor. At the high end they read lower.

e.) Conclusions re TSG

1. The TSG clock worked well and position information was available and reliable.

2. The flow rate was variable at ~2.5L/min early in the cruise and closer to 1.5 later.

3. The variability in the comparisons is high due to variations in the near-surface vertical and horizontal gradients, as well and in flow rates in the loop. Inlet sampling may be affected by sediments as well.

4. The TSG temperature was higher than that from the CTD by 0.134Cº (std. dev. 0.20Cº) in the offshore area and by 0.017Cº (std. dev. 0.16Cº) for the deeper casts in QCS. The offshore values are higher than noted during 2020-005 but the standard deviations are very high and many of the casts in that group are quite close to shore and are not well mixed at the surface.

5. The TSG salinity was lower than CTD salinity by a median of 0.06psu in the offshore area (std dev  0.15psu). In the inlets the TSG salinity was lower by a median of 1.5psu (std. dev. 2.7psu). The inlet data cannot be used to recalibrate as they are clearly affected by temporal and vertical variability. In a few deeper casts in QCS it was low by 0.31psu (std. dev. 0.06psu).

6. Are the Queen Charlotte Sound data indicative of a change in calibration? The vertical gradients in QCS are large and a displacement of 2m between TSG and CTD would explain many of the differences. There is also an issue of extremely low transmissivity at the surface in some of the inlets which may indicate suspended sediments that could affect salinity values in both the TSG and loop samples. It could affect CTD salinity as well but that would likely be a lesser problem since the sensor would get flushed at depth and after a cast so there would not be the buildup that could occur in the loop or TSG itself.

7. The comparisons with CTD data do not justify recalibration of temperature or salinity. The few loop samples do suggest that the TSG salinity was reading significantly low in the inlets. Unfortunately the one bottle flagged due to a loose cap is the one that might have shown a return to expected differences as the ship moved of the Strait of Georgia. It does show a smaller difference than in the inlet section and if there was evaporation of the sample then the “real” difference would be even lower. Further complicating the issue is that the loop samples may also be affected by suspended sediments.

8. It appears that the TSG salinity was reading low in the inlets, likely due to suspended sediments or other contaminants. This may also explain the high variability as the amount of such contaminants would vary from place to place.

9. The CTD rosette samples could also be affected although given the problems with flushing in inlets they probably contain samples that mostly come from below the turbid layer. Nonetheless, examination of a  few casts that were outliers in COMPARE do come from areas of very low transmissivity at 5m.

10. No recalibration will be applied to the TSG data.

g.) Editing

The only editing required was done using a text editor to pad temperature, salinity, fluorescence and sigma-T in the opening section of file 20200730-193021.reo and a short section of negative values in fluorescence at 19:31 hours.

h.) Preparing Final Files

REMOVE was used to remove channel Pressure, Temperature:Difference , and record # in all files and to remove channel Fluorescence:URU:Wetlabs in all files up to the end of file 2020-00730-000000.reo.

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments.

A cross-reference list was prepared.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series an all looks fine.

Particulars  - Notes from Daily Science Log

PAR sensor off: 39-152, 328-412

1. All bottles tripped – no sampling – no CHE file needed.

33. Variability at bottom of cast – combination of shed wake & high gradient near bottom.

39-42. Problems with primary CTD salinity – started during upcast 39; ended during upcast 42.

44. Bottles 1 & 2 fired but no sampling. No CHE file needed.

51. Duplicate sample #161 for bottle 16 – changed to next available #164.

59. Problem with NMEA – needed to reboot. File looks ok. Long time at surface with pumps off; TR & C looks out of water but suddenly in water when pressure -0.1db.

68. Altimetry messy at bottom. Estimate ~4m.

70. Station name should be MT11B. Fixed in CNV & ROS file.

74. Deck pressure 0.2db.

89. Fluorescence trace very jagged – probably correct based on net tow.

95. Station name should be JPS2. Fixed in CNV file. Depth in log clearly wrong.

105. Crane leaking hydraulic fluid on upcast. Stopped to trip all bottles. No CHE file needed.

113. Pumps off – return to surface – pumps on – full cast. Removed original drop using text editor.

124. Test fire bottle 17 – No CHE file needed.

126. 2 bottles sampled. Bottle 3 fired but not used. Remove line from CHE.

141. Large jelly left tendril all over rosette.

151. Touched bottom. No evidence of that in transmissivity.

152. Brief stop at 55m during downcast.

154. PAR reinstalled before this cast.

162. Vent on bottle 3 not closed; several others bit leaky.

170. Bottle 2 (75m) fired with only 5s wait. Nutrients sampled, no others.

171. Bottle fired on the fly at 41m

179. Fired bottle 8 (30m) after only 5s wait. Salinity and carbon sampling..

184. Fired Niskin 11 at wrong depth – no sampling. Remove from CHE.

193. Niskin 2 not tripped.

196. 2 bottles fired – only 1 sampled. Remove bottle2 from CHE.

201-202. Spikes in fluorescence at 80m. Both Transmissometers have local minimum ~50m.

221. Station name should be HECS7. Fixed in CNV & ROS file.

227. Up cast stopped ~16m due to mass of kelp tangled on the wire – LARS jibbed in, kelp set itself free.

232. Missed 100m bottle, went back down.

233-236. Did 10m soak; subsurface oxygen increase 50-100m. Needs 10m soak removed.

236. Deck pressure 0.3db

245. Pause at 238m to correct large aft lead.

261. Station name should be DOUG32. Corrected in CNV file.

268. Pumps off during 10m soak. Turned on at surface but soak not repeated. Needs 10m soak removed.

270. Bottles not sampled because surface bottles not tripped. No CHE file needed.

275. Station name should be DOUG8D.

310. Event # should be 309. File name fixed in CNV file.

311. Sample # repeat so changed to 91141.

315-317. Surface salinity very low, <1psu.

318. 3 bottles missed – went back down to trip them.

319. 3 surface bottles from small boat – no CTD data. Not processed – may be archived elsewhere.

328. PAR removed before this cast.

372. Pump and archiving on deck after cast

374. Lost connection part-way through cast – returned to surface, restarted Seasave, over-wrote hex file.

396. Bottle 5 & 6 both tripped at 50m. Use Bottle 5, remove Niskin #6.

2020-069

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 443 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 4700 |  | 11Dec2019 |  | Factory |  |  |  |  |
| Conductivity |  |  | 3531 |  | 12 ec2019 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 4888 |  | 11Dec2019 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 4513 |  | 28Jan2020 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 8Feb2020 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 8Feb2020 |  | IOS |  |  |  |  |
| SBE 43 DO sensor |  |  | 3791 |  | 7Dec2019 |  | Factory |  |  |  |  |
| PAR |  |  | 70613 |  | 21Mar2016 |  | Factory |  |  |  |  |
| SPAR |  |  | 20518 |  | 21Mar2016 |  | Factory |  |  |  |  |
| pH:SBE |  |  | 692 |  | 14Jan2020 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3640 |  |  |  | Factory |  |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 9Jan2020 |  | Factory |  |  |  |  |
| Altimeter |  |  | 20518 |  | 21Mar2016 |  | Factory |  |  |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 0620 | 21Jan20 | Factory |  |  |
| Conductivity | 0620 | 21Jan20 | Factory |  |  |
| WETStar Fluorometer | 953 | 30Dec2019 | Factory |  |  |