## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 12 January 2020 – 3 February 2021

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 19 February 2021 –  22 February 2021

Number of original Hex files: 9 	Number of casts processed: 9

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (serial #1449) was used with an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

These data form part of a long-standing time-series rather than a cruise. They cover sampling by the Canadian Forces Maritime Experimental and Test Ranges during 2019 in Nanoose Bay, BC.

Calibration information for the CTD could not be confirmed; the sensors have not been recalibrated since 2014.

To assist in analysis, event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However since this is a time series, a 7-character system is used for this project based on the date. A 4-character format was derived from those IDs based on the month and day of each cast. There were no cases of more than 1 cast on a single day.

The soak period as seen in the files varied greatly from 21 seconds to 80 seconds. It is unknown what pump delay was set or if data logging was delayed. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality. However, conductivity appears to have equilibrated before downcasts began.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005. Other errors are expected due to calibration drift.

The times given in the SeaBird headers do not include a time zone, but it is believed that local time was used. All times were converted to UTC by adding 7 or 8 hours depending on whether local time was PDT or PST.

One file (N21020301.ctd) from 2021 was received with the 2020 files, so was processed and moved to osd_data_archive:\Cruise_data\2020\2020-099\History.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

9 hex and con files were delivered. One of the files was from 2021.

The file names were non-standard and contain the letter (U)N followed by the date in YYMMDD format and 2 digits to indicate the sequential casts on a single day. The non-standard names are in the usual format for this project except for the initial (U) so that part was removed.

- Conversion of Raw Data

- A test conversion was done to see if the 2 voltage channels contained useful data. Voltage 0 corresponds to the altimeter channel and Voltage 1 contained only noise.

- The raw data were converted by matching the configuration file to the input file name.

- A spreadsheet “NanooseCtdStationLocationss2020.xls” was provided with positions for all casts.

- The spreadsheet provided was saved as NanooseCtdStations-merge-hdr.csv. Some editing was required:

- The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this.

- A column LOC:Station was added with every row having “Nanoose” entered.

- A column of event numbers was added to the CSV file using MMDD part of the file name as event number.

- Data were acquired at the surface for between 21s to 80s. It is unknown what pump delay was set or if data logging was delayed. A 2-minute soak is generally advised for the best salinity data. The conductivity seems to have reached equilibrium before the downcasts began.

- Temperature and conductivity generally look ok.

- The altimetry looks fine for 7 casts. The CTD did not get within 15m for the other 2 casts so the altimetry will not get entered into the IOS Headers.

- WILDEDIT

No spikes were noted in the data so WILDEDIT was not run.

- FILTER

As usual for SBE 19 CTDs the conductivity was low-pass filtered with a variety of time constants and the best results were with 0.7 seconds to force it to have the same response as the temperature.

The pressure was not filtered since it is quite smooth already other than a couple of reversals that look like they might be. In previous years tests showed that filtering made things worse in sections where pressure was constant.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Temperature was advanced by 0.73s relative to pressure since that has worked well for data from this project in previous years. Two casts were checked and the results were excellent. Alignment can be fine-tuned later using SHIFT if further adjustment is found useful. Tests were also run using slightly different advancements but 0.73s looked best.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.
The program MERGE CSV FILES TO HEADERS was then used to add positions, station name, event numbers and project name to the headers.

8. Checking Headers

Track Plot was used to ensure positions look normal. There are no problematic outliers though, as usual, there is some scatter in locations. The plot may be found at the end of this report.

Header Check was run. All variable ranges look reasonable.

- The altimetry header information from each file was exported to a spreadsheet. 7 casts got within 15m of the bottom.  The sum of the maximum depth sampled plus altimeter header entry was used to add water depth to file NanooseCtdStations-merge-hdr.csv.

- The header entries for water depth were blank for 2 casts that did not get within 15m of the bottom so they were removed.

Program MERGE CSV FILES TO HEADERS was then used to add positions, station name, event numbers and project name to the headers.

A cross-reference list was produced. The times appear to be in local time which is the usual case for this project. The times of casts are not consistent enough to notice if clocks might have changed from standard to daylight time, but in the past the Nanoose project was thought to use local time with both PST and PDT in use.

ADD TIME CHANNEL was used to add 8 hours to casts up to March 7th and from November 1rd to the end of the year and for the 1 cast from February 2021.

ADD TIME CHANNEL was used to add 7 hour to all other casts.

Surface check was run and values ranged from +0.3 to +1.4db. The average surface was 1.06db. , So the pressure looks reasonable given the low resolution of strain gauge sensors.

Profile plots were examined. Many casts will need some editing.

The need to edit will be re-examined after DELETE.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features.  For a few casts a shift of -0.5records made slight improvements while making most casts worse; variable descent rates and gradients limit the ability to improve alignment. SHIFT was not run.

10. CLIP

To remove records that are likely to have poor conductivity, CLIP was used to remove the first 50 scans for most files, but only 40 for the first 2 casts that started down before the 50-record point.

11. DELETE

DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings.

12. CTDEDIT

The DEL files were copied to EDT.

CTDEDIT was used to edit all casts .

Salinity was smoothed in cases where unstable features appeared to be due to misalignments of T and C in the presence of temperature gradients. A few records were removed from the bottom of some casts.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen.

T-S plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – Profile plots of temperature and salinity data were examined with the local seasonal climatology ranges superimposed. All casts feel within the climatology.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

-----------------------------------------------------------------------------------------

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The CTD sensors have not been calibrated since 2014.

For processing details see 2020-099_Processing_Report.doc.

-----------------------------------------------------------------------------------------

The standards check routine was run and no problems were found.

T-S plots were examined and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing and header check were produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	 12 January 2020	    End:	3 February 2021

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

CTDs

| CTD# | Make | Model | Serial# | Altimeter | Calibration data – Pressure, Temperature, Conductivity |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 19 | 1449 | 1094 | September / October 2014 |

The calibration information used in conversion can be found in the headers of the CTD.