#### 2021-008

CCGS John P Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### August 24-September 07, 2021

Prepared by Michael Arychuk for Danielle Caleb

- Sample Collection

Samples were collected from stations  P2, P4, P12, P16, P20 and P26 for DMS.  Samples were collected from P2, P4, P12, P16, P20 and P26 for DMSPD (dissolved) & DMSPT (total).

  - DMS

Fourteen water samples from various depths (400m, 200m, 150m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at station P4, P12, P16, P20 and P26.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface). Duplicates were taken at 20m.  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in a fridge, in the dark and removed one at a time before analysis.

  - DMSP

Seven samples for both DMSPD and DMSPT were collected ( 400m, 200m, 150m, 30m, 20m, 5m, surface).  The only exception to this was P2 where there were no 400m, 200m or 150m samples, hence, only 4 samples were collected. Duplicates were taken at 20m.

- Analysis

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Helium for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasil 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a 15mL or 5mL polypropylene tube.  The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

Exactly 3.5mL of seawater was collected directly from the niskin into a 15mL or 5mL polypropylene tube. The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by either using MQ water, some of the bottled SQ water brought from the lab or using some deep sea water (600 m or deeper) and treating them as actual samples.    For example, in the case of DMSPD,   3.5mL was collected from a separate funnel and for DMSPT    3.5mL  was added directly to the polypropylene tube.  Like the samples, they were then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The blank was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Data & Results

  - DMS

Data from this cruise showed some very high levels of DMS at P2 and P20.  In fact, overall the DMS values were higher than what has been seen traditionally.

There were also issues with the MQ water on this cruise (see final remarks in the “Conclusion.”)  The MQ water was clean enough  to make up the blanks and standards for station P2 and P4 but after P4 it was contaminated with sulfur compounds and could not be used.  Water was collected from P8, P12 and P20 at 2000m, and used within 12 hours of collection, as the blank and to make up the standards at subsequent stations. It was free from DMS contamination as the blanks were not detectible.

Initially the P8  2000m water was stored under the counter, in the dark, in the DMS lab container.  Trace levels of DMS was detected in the water collected from P8 after 24 hours and 0.5nM of DMS detected in the deep water collected from P8 within 48 hours. This could possibly be the result of biological activity in the water collected but is important to note because it indicates that the deep water (if being used as blank and standard water)needs to be fresh for it to be free of contamination. The water collected from P12 and P20 was subsequently stored in the  walk-in cooler and used within 12 hours of collection.

  - DMSP

The DMSP samples were collected for all stations.  Samples were run back in the lab in late September 2021.

- Conclusions

6.1 DMS

There is a strong suspicion that the ship water is affecting the quality of the MQ water on board the ship and it would be beneficial to get some written clarification from the Engineers regarding the treatment process involved and if it has changed in the past couple of years.  Furthermore it would be good to know how long the ship goes before they have to make water to replace what was in the holding tanks (that were likely filled with municipal water)?  Typically, in the past, the quality of the MQ water would improve as the cruise went on but over the past year or so the trend seems to have reversed and more sulphur contamination is being detected in the MQ water.  The problem is not with the MQ system because the system is tested in the lab before the cruise and it is 100% clean when packed up. Once on the ship it  starts to get worse as each day progresses to the point where the contamination overwhelms the blanks and alternate sources of clean water need to be found (i.e. deep sea water).  Deep water access is a problem for P2 and P4 because of the depth (typically the water needs to be deeper than 600m and the timing of the stations.  If the problem with the MQ water on the ship cannot be resolved then additional testing will have to be done to determine how long deep water can be stored and under what conditions before the biology starts to impact its purity with respect DMS blanks and standards.

  - DMSP

Sample collection went well without any issues.