## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2022/04/18 15:14:25 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2021-036-0004.mrg | 2021-036-0004.sam | 241 | 1721.19 | 1721.19 | 0.31433 | -0.00012207 | 34.569 | 34.5523 | 0.00027374 | -0.016705 | 34.569 | 34.5568 | 0.00023047 | -0.012192
2021-036-0008.mrg | 2021-036-0008.sam | 220 | 2578.93 | 2579.04 | 0.43167 | 0.10986 | 34.6384 | 34.6283 | 0.00026853 | -0.010059 | 34.6384 | 34.6334 | 0.00025222 | -0.005024
2021-036-0015.mrg | 2021-036-0015.sam | 241 | 552.774 | 552.774 | 0.27248 | 0.00018311 | 34.1526 | 34.1449 | 0.0060964 | -0.0076675 | 34.1526 | 34.151 | 0.0050416 | -0.0015717
2021-036-0026.mrg | 2021-036-0026.sam | 241 | 1351.16 | 1351.16 | 0.28131 | -0.00012207 | 34.4691 | 34.4564 | 0.00068006 | -0.012726 | 34.4691 | 34.4608 | 0.00079143 | -0.0082512
2021-036-0032.mrg | 2021-036-0032.sam | 241 | 1333.6 | 1333.6 | 0.44855 | 0.00024414 | 34.4643 | 34.4482 | 0.00027835 | -0.016109 | 34.4643 | 34.4522 | 0.00024542 | -0.012081
2021-036-0035.mrg | 2021-036-0035.sam | 241 | 1110.28 | 1110.28 | 0.28508 | 0.00024414 | 34.3527 | 34.3413 | 0.00028621 | -0.011448 | 34.3527 | 34.345 | 0.00023379 | -0.0077324
2021-036-0037.mrg | 2021-036-0037.sam | 209 | 2161.2 | 2161.36 | 0.5697 | 0.15698 | 34.6049 | 34.5949 | 0.00035066 | -0.010029 | 34.6049 | 34.5995 | 0.0002549 | -0.0053787
2021-036-0052.mrg | 2021-036-0052.sam | 241 | 3185.41 | 3185.41 | 0.39766 | 0 | 34.6516 | 34.6383 | 0.00025351 | -0.013329 | 34.6516 | 34.6427 | 0.00025519 | -0.0089111
2021-036-0056.mrg | 2021-036-0056.sam | 182 | 3214.43 | 3214.57 | 0.66976 | 0.14062 | 34.6601 | 34.6469 | 0.00031385 | -0.013176 | 34.6601 | 34.6516 | 0.00025195 | -0.0084915
2021-036-0056.mrg | 2021-036-0056.sam | 170 | 3197.64 | 3197.69 | 0.63464 | 0.043457 | 34.6545 | 34.6446 | 0.00036327 | -0.009861 | 34.6545 | 34.6492 | 0.00024773 | -0.0053139
2021-036-0056.mrg | 2021-036-0056.sam | 183 | 3000.34 | 3000.29 | 0.58571 | -0.049561 | 34.6506 | 34.6377 | 0.00030456 | -0.012909 | 34.6506 | 34.6423 | 0.00023787 | -0.0083389
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 2750.59 | 2750.59 | 0.18476 | 0.00048828 | 34.641 | 34.6281 | 0.00029127 | -0.012875 | 34.641 | 34.6328 | 0.00023803 | -0.0081978
2021-036-0056.mrg | 2021-036-0056.sam | 104 | 2498.65 | 2498.66 | 0.61379 | 0.01123 | 34.6255 | 34.613 | 0.00037396 | -0.012524 | 34.6255 | 34.6174 | 0.0002332 | -0.0081024
2021-036-0056.mrg | 2021-036-0056.sam | 240 | 2250.13 | 2250.13 | 0.59218 | -0.0046387 | 34.6091 | 34.5943 | 0.00030734 | -0.014759 | 34.6091 | 34.5988 | 0.00026409 | -0.010307
2021-036-0056.mrg | 2021-036-0056.sam | 196 | 1998.2 | 1997.92 | 0.50031 | -0.27344 | 34.5817 | 34.5687 | 0.00032206 | -0.01302 | 34.5817 | 34.573 | 0.00023205 | -0.0087128
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 1751.71 | 1751.71 | 0.39147 | -0.00012207 | 34.5393 | 34.5259 | 0.00030082 | -0.013443 | 34.5393 | 34.5302 | 0.00025208 | -0.0091095
2021-036-0056.mrg | 2021-036-0056.sam | 238 | 1499.89 | 1499.91 | 0.62188 | 0.013306 | 34.491 | 34.4791 | 0.00038575 | -0.011948 | 34.491 | 34.4832 | 0.00031847 | -0.0077972
2021-036-0056.mrg | 2021-036-0056.sam | 188 | 1349.88 | 1350.08 | 0.61762 | 0.19702 | 34.4541 | 34.44 | 0.00041443 | -0.014091 | 34.4541 | 34.4439 | 0.00042921 | -0.0102
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 1250.08 | 1250.08 | 0.28458 | -0.00024414 | 34.4352 | 34.4158 | 0.00025853 | -0.019421 | 34.4352 | 34.4199 | 0.00021976 | -0.015293
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 1101.41 | 1101.41 | 0.5784 | 0.00036621 | 34.3891 | 34.3759 | 0.00035828 | -0.013245 | 34.3891 | 34.3802 | 0.00033137 | -0.0089188
2021-036-0056.mrg | 2021-036-0056.sam | 229 | 1048.56 | 1048.51 | 0.32912 | -0.057007 | 34.3714 | 34.3559 | 0.00032745 | -0.015476 | 34.3714 | 34.3603 | 0.0003974 | -0.011112
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 950.634 | 950.634 | 0.24631 | 0.00048828 | 34.3372 | 34.321 | 0.00027289 | -0.016151 | 34.3372 | 34.325 | 0.00024858 | -0.012219
2021-036-0056.mrg | 2021-036-0056.sam | 217 | 900.292 | 900.295 | 0.55024 | 0.0031128 | 34.3107 | 34.2969 | 0.00034131 | -0.013809 | 34.3107 | 34.3007 | 0.0003129 | -0.0099945
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 799.562 | 799.562 | 0.29657 | 0.00018311 | 34.2784 | 34.2633 | 0.00030765 | -0.015095 | 34.2784 | 34.2671 | 0.00032455 | -0.01133
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 700.391 | 700.391 | 0.58316 | -0.00030518 | 34.2213 | 34.2083 | 0.00023683 | -0.012974 | 34.2213 | 34.2122 | 0.00023032 | -0.0091095
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 601.094 | 601.094 | 0.32236 | -0.00042725 | 34.1521 | 34.1403 | 0.00043397 | -0.011787 | 34.1521 | 34.1437 | 0.00049942 | -0.0084152
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 500.164 | 500.164 | 0.48616 | 0.00027466 | 34.0737 | 34.057 | 0.00045502 | -0.016743 | 34.0737 | 34.0607 | 0.00030891 | -0.012966
2021-036-0056.mrg | 2021-036-0056.sam | 211 | 399.695 | 399.544 | 0.61932 | -0.15131 | 33.9627 | 33.9455 | 0.0010964 | -0.017166 | 33.9627 | 33.9504 | 0.0011882 | -0.012291
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 301.047 | 301.047 | 0.37718 | -6.1035e-05 | 33.8679 | 33.8522 | 0.00038926 | -0.015697 | 33.8679 | 33.856 | 0.00038201 | -0.011909
2021-036-0056.mrg | 2021-036-0056.sam | 241 | 10.592 | 10.5925 | 0.31689 | 0.00046921 | 32.2484 | 32.2322 | 0.0029185 | -0.016243 | 32.2484 | 32.2347 | 0.00205 | -0.013664
2021-036-0062.mrg | 2021-036-0062.sam | 241 | 2249.74 | 2249.74 | 0.5872 | 0.00024414 | 34.6205 | 34.6047 | 0.00028137 | -0.015785 | 34.6205 | 34.6091 | 0.00024958 | -0.011402
2021-036-0067.mrg | 2021-036-0067.sam | 240 | 1656.37 | 1656.38 | 0.59132 | 0.0037842 | 34.538 | 34.5237 | 0.0005367 | -0.014297 | 34.538 | 34.5279 | 0.00044922 | -0.010056
2021-036-0071.mrg | 2021-036-0071.sam | 241 | 2523.32 | 2523.32 | 0.47387 | -0.00048828 | 34.6292 | 34.6131 | 0.00031025 | -0.016064 | 34.6292 | 34.6175 | 0.0002414 | -0.011677
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 12
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2021-036-0004.mrg | 2021-036-0004.sam | 12 | 12 | 2021-036-0004.mrg & 2021-036-0004.sam
2 | 2021-036-0008.mrg | 2021-036-0008.sam | 14 | 14 | 2021-036-0008.mrg & 2021-036-0008.sam
3 | 2021-036-0015.mrg | 2021-036-0015.sam | 16 | 16 | 2021-036-0015.mrg & 2021-036-0015.sam
4 | 2021-036-0026.mrg | 2021-036-0026.sam | 18 | 18 | 2021-036-0026.mrg & 2021-036-0026.sam
5 | 2021-036-0032.mrg | 2021-036-0032.sam | 20 | 20 | 2021-036-0032.mrg & 2021-036-0032.sam
6 | 2021-036-0035.mrg | 2021-036-0035.sam | 22 | 22 | 2021-036-0035.mrg & 2021-036-0035.sam
7 | 2021-036-0037.mrg | 2021-036-0037.sam | 24 | 24 | 2021-036-0037.mrg & 2021-036-0037.sam
8 | 2021-036-0052.mrg | 2021-036-0052.sam | 26 | 26 | 2021-036-0052.mrg & 2021-036-0052.sam
9 | 2021-036-0056.mrg | 2021-036-0056.sam | 28 | 49 | 2021-036-0056.mrg & 2021-036-0056.sam
10 | 2021-036-0062.mrg | 2021-036-0062.sam | 51 | 51 | 2021-036-0062.mrg & 2021-036-0062.sam
11 | 2021-036-0067.mrg | 2021-036-0067.sam | 53 | 53 | 2021-036-0067.mrg & 2021-036-0067.sam
12 | 2021-036-0071.mrg | 2021-036-0071.sam | 55 | 55 | 2021-036-0071.mrg & 2021-036-0071.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 12
2 | (Sal1 - Sal_Bot) | Pressure | 12
3 | (Sal1 - Sal_Bot) | Pressure | 12
4 | (Sal0 - Sal_Bot) | Pressure | 12
5 | (Sal0 - Sal_Bot) | File Pair # | 12
6 | (Sal1 - Sal_Bot) | Pressure | 12
7 | (Sal0 - Sal_Bot) | Pressure | 1
8 | (Sal0 - Sal_Bot) | Pressure | 12

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.13796 |  | X | CALC Y
m= | 9.4388e-07 | 4.4588e-07 |  | Std Dev of Resid |  | 0.0020864 |  | 301.047 | -0.015189
b= | -0.015473 | 0.00085546 |  | Resid Sum of Squares |  | 0.00012188 |  | 3214.57 | -0.012439
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 30
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1721.19 | 34.5523 | 0.00027374 | -0.016705 | -0.013849 | 0.0028559 | 34.569 | 1721.19 | -0.016705
2 | YES | 2579.04 | 34.6283 | 0.00026853 | -0.010059 | -0.013039 | -0.0029796 | 34.6384 | 2579.04 | -0.010059
3 | NO | 552.774 | 34.1449 | 0.0060964 | -0.0076675 | -0.014952 | -0.007284 | 34.1526 | 1351.16 | -0.012726
4 | YES | 1351.16 | 34.4564 | 0.00068006 | -0.012726 | -0.014198 | -0.0014721 | 34.4691 | 1333.6 | -0.016109
5 | YES | 1333.6 | 34.4482 | 0.00027835 | -0.016109 | -0.014215 | 0.001895 | 34.4643 | 1110.28 | -0.011448
6 | YES | 1110.28 | 34.3413 | 0.00028621 | -0.011448 | -0.014425 | -0.0029774 | 34.3527 | 2161.36 | -0.010029
7 | YES | 2161.36 | 34.5949 | 0.00035066 | -0.010029 | -0.013433 | -0.0034044 | 34.6049 | 3185.41 | -0.013329
8 | YES | 3185.41 | 34.6383 | 0.00025351 | -0.013329 | -0.012467 | 0.00086198 | 34.6516 | 3214.57 | -0.013176
9 | YES | 3214.57 | 34.6469 | 0.00031385 | -0.013176 | -0.012439 | 0.0007369 | 34.6601 | 3197.69 | -0.009861
9 | YES | 3197.69 | 34.6446 | 0.00036327 | -0.009861 | -0.012455 | -0.002594 | 34.6545 | 3000.29 | -0.012909
9 | YES | 3000.29 | 34.6377 | 0.00030456 | -0.012909 | -0.012641 | 0.00026755 | 34.6506 | 2750.59 | -0.012875
9 | YES | 2750.59 | 34.6281 | 0.00029127 | -0.012875 | -0.012877 | -2.442e-06 | 34.641 | 2498.66 | -0.012524
9 | YES | 2498.66 | 34.613 | 0.00037396 | -0.012524 | -0.013115 | -0.00059113 | 34.6255 | 2250.13 | -0.014759
9 | YES | 2250.13 | 34.5943 | 0.00030734 | -0.014759 | -0.013349 | 0.0014097 | 34.6091 | 1997.92 | -0.01302
9 | YES | 1997.92 | 34.5687 | 0.00032206 | -0.01302 | -0.013587 | -0.00056788 | 34.5817 | 1751.71 | -0.013443
9 | YES | 1751.71 | 34.5259 | 0.00030082 | -0.013443 | -0.01382 | -0.00037687 | 34.5393 | 1499.91 | -0.011948
9 | YES | 1499.91 | 34.4791 | 0.00038575 | -0.011948 | -0.014058 | -0.0021099 | 34.491 | 1350.08 | -0.014091
9 | YES | 1350.08 | 34.44 | 0.00041443 | -0.014091 | -0.014199 | -0.00010746 | 34.4541 | 1250.08 | -0.019421
9 | YES | 1250.08 | 34.4158 | 0.00025853 | -0.019421 | -0.014293 | 0.0051273 | 34.4352 | 1101.41 | -0.013245
9 | YES | 1101.41 | 34.3759 | 0.00035828 | -0.013245 | -0.014434 | -0.0011891 | 34.3891 | 1048.51 | -0.015476
9 | YES | 1048.51 | 34.3559 | 0.00032745 | -0.015476 | -0.014484 | 0.00099259 | 34.3714 | 950.634 | -0.016151
9 | YES | 950.634 | 34.321 | 0.00027289 | -0.016151 | -0.014576 | 0.0015754 | 34.3372 | 900.295 | -0.013809
9 | YES | 900.295 | 34.2969 | 0.00034131 | -0.013809 | -0.014624 | -0.00081431 | 34.3107 | 799.562 | -0.015095
9 | YES | 799.562 | 34.2633 | 0.00030765 | -0.015095 | -0.014719 | 0.00037621 | 34.2784 | 700.391 | -0.012974
9 | YES | 700.391 | 34.2083 | 0.00023683 | -0.012974 | -0.014812 | -0.0018384 | 34.2213 | 601.094 | -0.011787
9 | YES | 601.094 | 34.1403 | 0.00043397 | -0.011787 | -0.014906 | -0.0031185 | 34.1521 | 500.164 | -0.016743
9 | YES | 500.164 | 34.057 | 0.00045502 | -0.016743 | -0.015001 | 0.0017415 | 34.0737 | 301.047 | -0.015697
9 | NO | 399.544 | 33.9455 | 0.0010964 | -0.017166 | -0.015096 | 0.0020699 | 33.9627 | 2249.74 | -0.015785
9 | YES | 301.047 | 33.8522 | 0.00038926 | -0.015697 | -0.015189 | 0.00050837 | 33.8679 | 1656.38 | -0.014297 |  |  | excluding 3 outliers |  | also excluding 4 flagged 2
9 | NO | 10.5925 | 32.2322 | 0.0029185 | -0.016243 | -0.015463 | 0.00077972 | 32.2484 | 2523.32 | -0.016064 | -0.013852 | 0.0022081 | -0.013852 | 0.0022081 | -0.1375 | 0.002311
10 | YES | 2249.74 | 34.6047 | 0.00028137 | -0.015785 | -0.01335 | 0.0024354 | 34.6205
11 | YES | 1656.38 | 34.5237 | 0.0005367 | -0.014297 | -0.01391 | 0.00038765 | 34.538
12 | YES | 2523.32 | 34.6131 | 0.00031025 | -0.016064 | -0.013092 | 0.0029721 | 34.6292
33
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.22838 |  | X | CALC Y
m= | 1.2886e-06 | 4.4761e-07 |  | Std Dev of Resid |  | 0.0020945 |  | 301.047 | -0.011444
b= | -0.011832 | 0.00085878 |  | Resid Sum of Squares |  | 0.00012283 |  | 3214.57 | -0.0076895
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 30
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y |  |  | Pressure | Sec-bot | Pri-bot | Sec-Pri
1 | YES | 1721.19 | 34.5568 | 0.00023047 | -0.012192 | -0.0096139 | 0.0025779 |  | 1721.19 | -0.012192 |  |  | 1721.19 | -0.012192 | -0.016705 | 0.0045128
2 | YES | 2579.04 | 34.6334 | 0.00025222 | -0.005024 | -0.0085084 | -0.0034845 |  | 2579.04 | -0.005024 |  |  | 2579.04 | -0.005024 | -0.010059 | 0.0050354
3 | NO | 552.774 | 34.151 | 0.0050416 | -0.0015717 | -0.011119 | -0.0095478 |  | 1351.16 | -0.0082512 |  |  | 552.774 | -0.0082512 | -0.012726 | 0.0044746
4 | YES | 1351.16 | 34.4608 | 0.00079143 | -0.0082512 | -0.010091 | -0.0018395 |  | 1333.6 | -0.012081 |  |  | 1351.16 | -0.012081 | -0.016109 | 0.0040284
5 | YES | 1333.6 | 34.4522 | 0.00024542 | -0.012081 | -0.010113 | 0.0019678 |  | 1110.28 | -0.0077324 |  |  | 1333.6 | -0.0077324 | -0.011448 | 0.0037155
6 | YES | 1110.28 | 34.345 | 0.00023379 | -0.0077324 | -0.010401 | -0.0026687 |  | 2161.36 | -0.0053787 |  |  | 1110.28 | -0.0053787 | -0.010029 | 0.0046501
7 | YES | 2161.36 | 34.5995 | 0.0002549 | -0.0053787 | -0.0090467 | -0.0036679 |  | 3185.41 | -0.0089111 |  |  | 2161.36 | -0.0089111 | -0.013329 | 0.0044175
8 | YES | 3185.41 | 34.6427 | 0.00025519 | -0.0089111 | -0.0077271 | 0.001184 |  | 3214.57 | -0.0084915 |  |  | 3185.41 | -0.0084915 | -0.013176 | 0.0046845
9 | YES | 3214.57 | 34.6516 | 0.00025195 | -0.0084915 | -0.0076895 | 0.000802 | 34.6601 | 3197.69 | -0.0053139 |  |  | 3214.57 | -0.0053139 | -0.009861 | 0.0045471
9 | YES | 3197.69 | 34.6492 | 0.00024773 | -0.0053139 | -0.0077113 | -0.0023974 | 34.6545 | 3000.29 | -0.0083389 |  |  | 3197.69 | -0.0083389 | -0.012909 | 0.00457
9 | YES | 3000.29 | 34.6423 | 0.00023787 | -0.0083389 | -0.0079656 | 0.00037329 | 34.6506 | 2750.59 | -0.0081978 |  |  | 3000.29 | -0.0081978 | -0.012875 | 0.0046768
9 | YES | 2750.59 | 34.6328 | 0.00023803 | -0.0081978 | -0.0082874 | -8.9615e-05 | 34.641 | 2498.66 | -0.0081024 |  |  | 2750.59 | -0.0081024 | -0.012524 | 0.0044213
9 | YES | 2498.66 | 34.6174 | 0.0002332 | -0.0081024 | -0.008612 | -0.0005096 | 34.6255 | 2250.13 | -0.010307 |  |  | 2498.66 | -0.010307 | -0.014759 | 0.0044518
9 | YES | 2250.13 | 34.5988 | 0.00026409 | -0.010307 | -0.0089323 | 0.001375 | 34.6091 | 1997.92 | -0.0087128 |  |  | 2250.13 | -0.0087128 | -0.01302 | 0.0043068
9 | YES | 1997.92 | 34.573 | 0.00023205 | -0.0087128 | -0.0092573 | -0.00054449 | 34.5817 | 1751.71 | -0.0091095 |  |  | 1997.92 | -0.0091095 | -0.013443 | 0.0043335
9 | YES | 1751.71 | 34.5302 | 0.00025208 | -0.0091095 | -0.0095745 | -0.00046502 | 34.5393 | 1499.91 | -0.0077972 |  |  | 1751.71 | -0.0077972 | -0.011948 | 0.0041504
9 | YES | 1499.91 | 34.4832 | 0.00031847 | -0.0077972 | -0.009899 | -0.0021017 | 34.491 | 1350.08 | -0.0102 |  |  | 1499.91 | -0.0102 | -0.014091 | 0.003891
9 | YES | 1350.08 | 34.4439 | 0.00042921 | -0.0102 | -0.010092 | 0.00010845 | 34.4541 | 1250.08 | -0.015293 |  |  | 1350.08 | -0.015293 | -0.019421 | 0.0041275
9 | YES | 1250.08 | 34.4199 | 0.00021976 | -0.015293 | -0.010221 | 0.0050722 | 34.4352 | 1101.41 | -0.0089188 |  |  | 1250.08 | -0.0089188 | -0.013245 | 0.0043258
9 | YES | 1101.41 | 34.3802 | 0.00033137 | -0.0089188 | -0.010412 | -0.0014937 | 34.3891 | 1048.51 | -0.011112 |  |  | 1101.41 | -0.011112 | -0.015476 | 0.004364
9 | YES | 1048.51 | 34.3603 | 0.0003974 | -0.011112 | -0.010481 | 0.00063155 | 34.3714 | 950.634 | -0.012219 |  |  | 1048.51 | -0.012219 | -0.016151 | 0.0039329
9 | YES | 950.634 | 34.325 | 0.00024858 | -0.012219 | -0.010607 | 0.0016117 |  | 900.295 | -0.0099945 |  |  | 950.634 | -0.0099945 | -0.013809 | 0.0038147
9 | YES | 900.295 | 34.3007 | 0.0003129 | -0.0099945 | -0.010672 | -0.00067712 |  | 799.562 | -0.01133 |  |  | 900.295 | -0.01133 | -0.015095 | 0.0037651
9 | YES | 799.562 | 34.2671 | 0.00032455 | -0.01133 | -0.010801 | 0.00052827 |  | 700.391 | -0.0091095 |  |  | 799.562 | -0.0091095 | -0.012974 | 0.0038643
9 | YES | 700.391 | 34.2122 | 0.00023032 | -0.0091095 | -0.010929 | -0.0018197 |  | 601.094 | -0.0084152 |  |  | 700.391 | -0.0084152 | -0.011787 | 0.0033722
9 | YES | 601.094 | 34.1437 | 0.00049942 | -0.0084152 | -0.011057 | -0.002642 |  | 500.164 | -0.012966 |  |  | 601.094 | -0.012966 | -0.016743 | 0.0037765
9 | YES | 500.164 | 34.0607 | 0.00030891 | -0.012966 | -0.011187 | 0.001779 |  | 301.047 | -0.011909 |  |  | 500.164 | -0.011909 | -0.015697 | 0.003788
9 | NO | 399.544 | 33.9504 | 0.0011882 | -0.012291 | -0.011317 | 0.00097411 |  | 2249.74 | -0.011402 |  |  | 399.544 | -0.011402 | -0.015785 | 0.0043831
9 | YES | 301.047 | 33.856 | 0.00038201 | -0.011909 | -0.011444 | 0.00046569 |  | 1656.38 | -0.010056 |  |  | 301.047 | -0.010056 | -0.014297 | 0.004242 | avg | stdev
9 | NO | 10.5925 | 32.2347 | 0.00205 | -0.013664 | -0.011818 | 0.0018461 |  | 2523.32 | -0.011677 | -0.0096181 | 0.0023429 | 10.5925 | -0.011677 | -0.016064 | 0.0043869 | 0.0042337 | 0.00037067
10 | YES | 2249.74 | 34.6091 | 0.00024958 | -0.011402 | -0.0089328 | 0.0024693
11 | YES | 1656.38 | 34.5279 | 0.00044922 | -0.010056 | -0.0096974 | 0.00035814
12 | YES | 2523.32 | 34.6175 | 0.0002414 | -0.011677 | -0.0085802 | 0.0030966
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.24442 |  | X | CALC Y
m= | 1.0959e-06 | 4.5416e-07 |  | Std Dev of Resid |  | 0.0018645 |  | 301.047 | -0.015355
b= | -0.015685 | 0.00083116 |  | Resid Sum of Squares |  | 6.2573e-05 |  | 3214.57 | -0.012162
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 20
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 1721.19 | 34.5523 | 0.00027374 | -0.016705 | -0.013799 | 0.0029056 |  | 3214.57 | -0.013176
2 | NO | 2579.04 | 34.6283 | 0.00026853 | -0.010059 | -0.012859 | -0.0027994 |  | 3197.69 | -0.009861
3 | NO | 552.774 | 34.1449 | 0.0060964 | -0.0076675 | -0.015079 | -0.0074119 |  | 3000.29 | -0.012909
4 | NO | 1351.16 | 34.4564 | 0.00068006 | -0.012726 | -0.014204 | -0.0014787 |  | 2750.59 | -0.012875
5 | NO | 1333.6 | 34.4482 | 0.00027835 | -0.016109 | -0.014224 | 0.0018858 |  | 2498.66 | -0.012524
6 | NO | 1110.28 | 34.3413 | 0.00028621 | -0.011448 | -0.014468 | -0.0030206 |  | 2250.13 | -0.014759
7 | NO | 2161.36 | 34.5949 | 0.00035066 | -0.010029 | -0.013317 | -0.0032878 |  | 1997.92 | -0.01302
8 | NO | 3185.41 | 34.6383 | 0.00025351 | -0.013329 | -0.012194 | 0.0011343 |  | 1751.71 | -0.013443
9 | YES | 3214.57 | 34.6469 | 0.00031385 | -0.013176 | -0.012162 | 0.0010136 |  | 1499.91 | -0.011948
9 | YES | 3197.69 | 34.6446 | 0.00036327 | -0.009861 | -0.012181 | -0.0023199 |  | 1350.08 | -0.014091
9 | YES | 3000.29 | 34.6377 | 0.00030456 | -0.012909 | -0.012397 | 0.0005117 |  | 1250.08 | -0.019421
9 | YES | 2750.59 | 34.6281 | 0.00029127 | -0.012875 | -0.012671 | 0.00020375 |  | 1101.41 | -0.013245
9 | YES | 2498.66 | 34.613 | 0.00037396 | -0.012524 | -0.012947 | -0.00042324 |  | 1048.51 | -0.015476
9 | YES | 2250.13 | 34.5943 | 0.00030734 | -0.014759 | -0.013219 | 0.0015398 |  | 950.634 | -0.016151
9 | YES | 1997.92 | 34.5687 | 0.00032206 | -0.01302 | -0.013496 | -0.0004761 |  | 900.295 | -0.013809
9 | YES | 1751.71 | 34.5259 | 0.00030082 | -0.013443 | -0.013766 | -0.00032253 |  | 799.562 | -0.015095
9 | YES | 1499.91 | 34.4791 | 0.00038575 | -0.011948 | -0.014041 | -0.0020939 |  | 700.391 | -0.012974
9 | YES | 1350.08 | 34.44 | 0.00041443 | -0.014091 | -0.014206 | -0.00011418 |  | 601.094 | -0.011787
9 | YES | 1250.08 | 34.4158 | 0.00025853 | -0.019421 | -0.014315 | 0.0051053 |  | 500.164 | -0.016743
9 | YES | 1101.41 | 34.3759 | 0.00035828 | -0.013245 | -0.014478 | -0.0012336 |  | 301.047 | -0.015697
9 | YES | 1048.51 | 34.3559 | 0.00032745 | -0.015476 | -0.014536 | 0.00094003
9 | YES | 950.634 | 34.321 | 0.00027289 | -0.016151 | -0.014643 | 0.001508
9 | YES | 900.295 | 34.2969 | 0.00034131 | -0.013809 | -0.014699 | -0.0008894
9 | YES | 799.562 | 34.2633 | 0.00030765 | -0.015095 | -0.014809 | 0.0002858
9 | YES | 700.391 | 34.2083 | 0.00023683 | -0.012974 | -0.014918 | -0.0019439
9 | YES | 601.094 | 34.1403 | 0.00043397 | -0.011787 | -0.015026 | -0.0032391
9 | YES | 500.164 | 34.057 | 0.00045502 | -0.016743 | -0.015137 | 0.0016056
9 | NO | 399.544 | 33.9455 | 0.0010964 | -0.017166 | -0.015247 | 0.0019187
9 | YES | 301.047 | 33.8522 | 0.00038926 | -0.015697 | -0.015355 | 0.00034218
9 | NO | 10.5925 | 32.2322 | 0.0029185 | -0.016243 | -0.015674 | 0.00056937
10 | NO | 2249.74 | 34.6047 | 0.00028137 | -0.015785 | -0.01322 | 0.0025655
11 | NO | 1656.38 | 34.5237 | 0.0005367 | -0.014297 | -0.01387 | 0.0004275
12 | NO | 2523.32 | 34.6131 | 0.00031025 | -0.016064 | -0.01292 | 0.0031438
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.37131 |  | X | CALC Y
m= | 1.4335e-06 | 4.3966e-07 |  | Std Dev of Resid |  | 0.001805 |  | 301.047 | -0.01163
b= | -0.012062 | 0.00080463 |  | Resid Sum of Squares |  | 5.8642e-05 |  | 3214.57 | -0.0074534
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 20
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 1721.19 | 34.5568 | 0.00023047 | -0.012192 | -0.0095942 | 0.0025976 |  | 3214.57 | -0.0084915
2 | NO | 2579.04 | 34.6334 | 0.00025222 | -0.005024 | -0.0083644 | -0.0033405 |  | 3197.69 | -0.0053139
3 | NO | 552.774 | 34.151 | 0.0050416 | -0.0015717 | -0.011269 | -0.0096975 |  | 3000.29 | -0.0083389
4 | NO | 1351.16 | 34.4608 | 0.00079143 | -0.0082512 | -0.010125 | -0.0018735 |  | 2750.59 | -0.0081978
5 | NO | 1333.6 | 34.4522 | 0.00024542 | -0.012081 | -0.01015 | 0.0019313 |  | 2498.66 | -0.0081024
6 | NO | 1110.28 | 34.345 | 0.00023379 | -0.0077324 | -0.01047 | -0.0027376 |  | 2250.13 | -0.010307
7 | NO | 2161.36 | 34.5995 | 0.0002549 | -0.0053787 | -0.0089632 | -0.0035845 |  | 1997.92 | -0.0087128
8 | NO | 3185.41 | 34.6427 | 0.00025519 | -0.0089111 | -0.0074952 | 0.0014159 |  | 1751.71 | -0.0091095
9 | YES | 3214.57 | 34.6516 | 0.00025195 | -0.0084915 | -0.0074534 | 0.0010381 |  | 1499.91 | -0.0077972
9 | YES | 3197.69 | 34.6492 | 0.00024773 | -0.0053139 | -0.0074776 | -0.0021637 |  | 1350.08 | -0.0102
9 | YES | 3000.29 | 34.6423 | 0.00023787 | -0.0083389 | -0.0077606 | 0.00057837 |  | 1250.08 | -0.015293
9 | YES | 2750.59 | 34.6328 | 0.00023803 | -0.0081978 | -0.0081185 | 7.9267e-05 |  | 1101.41 | -0.0089188
9 | YES | 2498.66 | 34.6174 | 0.0002332 | -0.0081024 | -0.0084797 | -0.00037724 |  | 1048.51 | -0.011112
9 | YES | 2250.13 | 34.5988 | 0.00026409 | -0.010307 | -0.0088359 | 0.0014714 |  | 950.634 | -0.012219
9 | YES | 1997.92 | 34.573 | 0.00023205 | -0.0087128 | -0.0091975 | -0.00048472 |  | 900.295 | -0.0099945
9 | YES | 1751.71 | 34.5302 | 0.00025208 | -0.0091095 | -0.0095504 | -0.00044094 |  | 799.562 | -0.01133
9 | YES | 1499.91 | 34.4832 | 0.00031847 | -0.0077972 | -0.0099114 | -0.0021142 |  | 700.391 | -0.0091095
9 | YES | 1350.08 | 34.4439 | 0.00042921 | -0.0102 | -0.010126 | 7.4311e-05 |  | 601.094 | -0.0084152
9 | YES | 1250.08 | 34.4199 | 0.00021976 | -0.015293 | -0.01027 | 0.0050236 |  | 500.164 | -0.012966
9 | YES | 1101.41 | 34.3802 | 0.00033137 | -0.0089188 | -0.010483 | -0.0015639 |  | 301.047 | -0.011909
9 | YES | 1048.51 | 34.3603 | 0.0003974 | -0.011112 | -0.010558 | 0.0005537
9 | YES | 950.634 | 34.325 | 0.00024858 | -0.012219 | -0.010699 | 0.0015197
9 | YES | 900.295 | 34.3007 | 0.0003129 | -0.0099945 | -0.010771 | -0.00077646
9 | YES | 799.562 | 34.2671 | 0.00032455 | -0.01133 | -0.010915 | 0.00041433
9 | YES | 700.391 | 34.2122 | 0.00023032 | -0.0091095 | -0.011058 | -0.001948
9 | YES | 601.094 | 34.1437 | 0.00049942 | -0.0084152 | -0.0112 | -0.0027847
9 | YES | 500.164 | 34.0607 | 0.00030891 | -0.012966 | -0.011345 | 0.0016216
9 | NO | 399.544 | 33.9504 | 0.0011882 | -0.012291 | -0.011489 | 0.00080219
9 | YES | 301.047 | 33.856 | 0.00038201 | -0.011909 | -0.01163 | 0.00027949
9 | NO | 10.5925 | 32.2347 | 0.00205 | -0.013664 | -0.012046 | 0.0016178
10 | NO | 2249.74 | 34.6091 | 0.00024958 | -0.011402 | -0.0088365 | 0.0025656
11 | NO | 1656.38 | 34.5279 | 0.00044922 | -0.010056 | -0.0096871 | 0.0003684
12 | NO | 2523.32 | 34.6175 | 0.0002414 | -0.011677 | -0.0084443 | 0.0032325
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_5
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.096959 |  | X | CALC Y
m= | 0.0002216 | 0.00033814 |  | Std Dev of Resid |  | 0.0025191 |  | 1 | -0.015397
b= | -0.015619 | 0.0027366 |  | Resid Sum of Squares |  | 2.5383e-05 |  | 10 | -0.013403
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 6
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1721.19 | 34.5523 | 0.00027374 | -0.016705 | -0.015397 | 0.0013075 |  | 1 | -0.016705
2 | NO | 2579.04 | 34.6283 | 0.00026853 | -0.010059 | -0.015176 | -0.0051161 |  | 7 | -0.010029
3 | NO | 552.774 | 34.1449 | 0.0060964 | -0.0076675 | -0.014954 | -0.0072864 |  | 9 | -0.014759
4 | NO | 1351.16 | 34.4564 | 0.00068006 | -0.012726 | -0.014732 | -0.0020065 |  | 9 | -0.01302
5 | NO | 1333.6 | 34.4482 | 0.00027835 | -0.016109 | -0.014511 | 0.0015988 |  | 9 | -0.013443
6 | NO | 1110.28 | 34.3413 | 0.00028621 | -0.011448 | -0.014289 | -0.0028412 |  | 10 | -0.015785
7 | YES | 2161.36 | 34.5949 | 0.00035066 | -0.010029 | -0.014068 | -0.0040387
8 | NO | 3185.41 | 34.6383 | 0.00025351 | -0.013329 | -0.013846 | -0.00051732
9 | NO | 3214.57 | 34.6469 | 0.00031385 | -0.013176 | -0.013624 | -0.00044832
9 | NO | 3197.69 | 34.6446 | 0.00036327 | -0.009861 | -0.013624 | -0.0037633
9 | NO | 3000.29 | 34.6377 | 0.00030456 | -0.012909 | -0.013624 | -0.00071542 |  |  |  | -0.013957
9 | NO | 2750.59 | 34.6281 | 0.00029127 | -0.012875 | -0.013624 | -0.00074972
9 | NO | 2498.66 | 34.613 | 0.00037396 | -0.012524 | -0.013624 | -0.0011006
9 | YES | 2250.13 | 34.5943 | 0.00030734 | -0.014759 | -0.013624 | 0.0011348
9 | YES | 1997.92 | 34.5687 | 0.00032206 | -0.01302 | -0.013624 | -0.00060472
9 | YES | 1751.71 | 34.5259 | 0.00030082 | -0.013443 | -0.013624 | -0.00018132
9 | NO | 1499.91 | 34.4791 | 0.00038575 | -0.011948 | -0.013624 | -0.0016767
9 | NO | 1350.08 | 34.44 | 0.00041443 | -0.014091 | -0.013624 | 0.00046718
9 | NO | 1250.08 | 34.4158 | 0.00025853 | -0.019421 | -0.013624 | 0.0057963
9 | NO | 1101.41 | 34.3759 | 0.00035828 | -0.013245 | -0.013624 | -0.00037972
9 | NO | 1048.51 | 34.3559 | 0.00032745 | -0.015476 | -0.013624 | 0.0018519
9 | NO | 950.634 | 34.321 | 0.00027289 | -0.016151 | -0.013624 | 0.0025271
9 | NO | 900.295 | 34.2969 | 0.00034131 | -0.013809 | -0.013624 | 0.00018488
9 | NO | 799.562 | 34.2633 | 0.00030765 | -0.015095 | -0.013624 | 0.0014705
9 | NO | 700.391 | 34.2083 | 0.00023683 | -0.012974 | -0.013624 | -0.00065052
9 | NO | 601.094 | 34.1403 | 0.00043397 | -0.011787 | -0.013624 | -0.0018369
9 | NO | 500.164 | 34.057 | 0.00045502 | -0.016743 | -0.013624 | 0.0031184
9 | NO | 399.544 | 33.9455 | 0.0010964 | -0.017166 | -0.013624 | 0.0035418
9 | NO | 301.047 | 33.8522 | 0.00038926 | -0.015697 | -0.013624 | 0.0020732
9 | NO | 10.5925 | 32.2322 | 0.0029185 | -0.016243 | -0.013624 | 0.0026187
10 | YES | 2249.74 | 34.6047 | 0.00028137 | -0.015785 | -0.013403 | 0.0023825
11 | NO | 1656.38 | 34.5237 | 0.0005367 | -0.014297 | -0.013181 | 0.0011164
12 | NO | 2523.32 | 34.6131 | 0.00031025 | -0.016064 | -0.01296 | 0.0031042
YES
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_6
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.077719 |  | X | CALC Y
m= | 0.00020246 | 0.00034872 |  | Std Dev of Resid |  | 0.0025979 |  | 1 | -0.010833
b= | -0.011035 | 0.0028222 |  | Resid Sum of Squares |  | 2.6996e-05 |  | 10 | -0.0090109
# of Values Extracted = |  |  | 33
# of Values Used in Fit = |  |  | 6
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1721.19 | 34.5568 | 0.00023047 | -0.012192 | -0.010833 | 0.0013588 |  | 1 | -0.012192
2 | NO | 2579.04 | 34.6334 | 0.00025222 | -0.005024 | -0.010631 | -0.0056066 |  | 7 | -0.0053787
3 | NO | 552.774 | 34.151 | 0.0050416 | -0.0015717 | -0.010428 | -0.0088564 |  | 9 | -0.010307
4 | NO | 1351.16 | 34.4608 | 0.00079143 | -0.0082512 | -0.010226 | -0.0019744 |  | 9 | -0.0087128
5 | NO | 1333.6 | 34.4522 | 0.00024542 | -0.012081 | -0.010023 | 0.0020579 |  | 9 | -0.0091095
6 | NO | 1110.28 | 34.345 | 0.00023379 | -0.0077324 | -0.0098207 | -0.0020883 |  | 10 | -0.011402
7 | YES | 2161.36 | 34.5995 | 0.0002549 | -0.0053787 | -0.0096183 | -0.0042395
8 | NO | 3185.41 | 34.6427 | 0.00025519 | -0.0089111 | -0.0094158 | -0.00050467 |  |  |  | -0.009517
9 | NO | 3214.57 | 34.6516 | 0.00025195 | -0.0084915 | -0.0092133 | -0.00072182
9 | NO | 3197.69 | 34.6492 | 0.00024773 | -0.0053139 | -0.0092133 | -0.0038995
9 | NO | 3000.29 | 34.6423 | 0.00023787 | -0.0083389 | -0.0092133 | -0.00087441
9 | NO | 2750.59 | 34.6328 | 0.00023803 | -0.0081978 | -0.0092133 | -0.0010156
9 | NO | 2498.66 | 34.6174 | 0.0002332 | -0.0081024 | -0.0092133 | -0.0011109
9 | YES | 2250.13 | 34.5988 | 0.00026409 | -0.010307 | -0.0092133 | 0.001094
9 | YES | 1997.92 | 34.573 | 0.00023205 | -0.0087128 | -0.0092133 | -0.00050057
9 | YES | 1751.71 | 34.5302 | 0.00025208 | -0.0091095 | -0.0092133 | -0.00010384
9 | NO | 1499.91 | 34.4832 | 0.00031847 | -0.0077972 | -0.0092133 | -0.0014161
9 | NO | 1350.08 | 34.4439 | 0.00042921 | -0.0102 | -0.0092133 | 0.00098716
9 | NO | 1250.08 | 34.4199 | 0.00021976 | -0.015293 | -0.0092133 | 0.0060798
9 | NO | 1101.41 | 34.3802 | 0.00033137 | -0.0089188 | -0.0092133 | -0.00029458
9 | NO | 1048.51 | 34.3603 | 0.0003974 | -0.011112 | -0.0092133 | 0.0018989
9 | NO | 950.634 | 34.325 | 0.00024858 | -0.012219 | -0.0092133 | 0.0030052
9 | NO | 900.295 | 34.3007 | 0.0003129 | -0.0099945 | -0.0092133 | 0.00078117
9 | NO | 799.562 | 34.2671 | 0.00032455 | -0.01133 | -0.0092133 | 0.0021164
9 | NO | 700.391 | 34.2122 | 0.00023032 | -0.0091095 | -0.0092133 | -0.00010384
9 | NO | 601.094 | 34.1437 | 0.00049942 | -0.0084152 | -0.0092133 | -0.00079812
9 | NO | 500.164 | 34.0607 | 0.00030891 | -0.012966 | -0.0092133 | 0.0037529
9 | NO | 399.544 | 33.9504 | 0.0011882 | -0.012291 | -0.0092133 | 0.0030777
9 | NO | 301.047 | 33.856 | 0.00038201 | -0.011909 | -0.0092133 | 0.0026962
9 | NO | 10.5925 | 32.2347 | 0.00205 | -0.013664 | -0.0092133 | 0.0044509
10 | YES | 2249.74 | 34.6091 | 0.00024958 | -0.011402 | -0.0090109 | 0.0023912
11 | NO | 1656.38 | 34.5279 | 0.00044922 | -0.010056 | -0.0088084 | 0.0012471
12 | NO | 2523.32 | 34.6175 | 0.0002414 | -0.011677 | -0.008606 | 0.0030708
YES
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE