## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2025 | RBR Oxygen:Dissolved:Concentration derived and added to 6 RBR files. SH |

## PROCESSING NOTES

Cruise: 2020-038

Agency: IOS, Ocean Sciences Division, Sidney, BC; Freshwater Institute, Central & Arctic Region, Winnipeg, MB.

Location:  Beaufort Sea

Project: Canadian Beaufort Sea Marine Ecosystem Assessment

Chief Scientist: Eert J.

Platform: Frosti

Date: 31 July 2021 – 12 September 2021

Processed by: Germaine Gatien and Samantha Huntington

Date of Processing: 13 January 2021  -  29 January 2022

Number of original HEX files: 1 SBE911+; 52 SBE19+;  6 RBR

Number of CTD files:  58	Number of bottle casts: 30

# INSTRUMENT SUMMARY

Three CTDs were used for this cruise:

1. For event #2 was a SeaBird Model SBE 911+ CTD (Arctic #1189) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#SCST1050DR), a SBE 43 DO sensor (#637), a SeaPoint Fluorometer (#SCF3651), a Biospherical PAR sensor (#70501), a surface PAR sensors (#20279), a SeaPoint turbidity meter (#11074), an Optode Dissolved Oxygen sensor (16) and an altimeter (#1161).

The data acquisition program was SeaSave V 7.26.7.107.

2. For most events a SeaBird Model SBE 19+ CTD (6009) was mounted in a rosette and attached were an SBE 43 DO sensor (#615), a SeaPoint Fluorometer (#SCF2745), a SeaPoint Turbidity meter (#11432) and an altimeter (#1186).

The data acquisition program was SeaSave 2.1.0.

3. For 5 events an RBR Concerto (65579) was deployed from a Hurricane Zodiac skiff using a 100m long  marked handline measuring Salinity, Temperature, Depth, Chlorophyll a fluorescence (Turner fluorometer), and dissolved oxygen concentration (Rinko III Optode oxygen sensor)

# SUMMARY OF QUALITY AND CONCERNS

No Daily Science Log book was available. There were scanned rosette sheets, an RBR log, a bridge log and an electronic sampling log with sample numbers, Niskin #s, nominal closure depth and comments. Some other documents provided were incomplete. There was no list of equipment.

The raw file names had a small format error that was fixed before processing.

Deployment included a subsurface soak at around 5db for the SBE911+ cast; for other casts there was a long soak at the surface, but no 5m soak.

There was no dissolved oxygen sampling.  Oxygen saturation values at 2m were ~103% for the SBE911+ cast and ranged from 100% to 102% for the RBR casts. There was a wide range of saturation values at 2m from the SBE19+ with most between 103% and 107%. We expect calibration drift in the SBE sensors to be towards lower values, so this is likely not evidence of drift. No recalibration was applied.

There was salinity sampling to compare with 1 RBR cast and many SBE19+ casts. Comparisons with bottles indicate that both were producing reliable salinity data. The SBE19+ appears to be good to ±0.003psu. No estimate is possible for the RBR but the 1 point of comparison looked good.

Data from all 3 CTDs were combined for archival. For a summary of equipment used see document 2021-038-ctd-final.rpt.

# PROCESSING SUMMARY

Data from each of 3 CTDs were processed separately.

SECTION A – SBE911+  (Event # 2 only)

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Only 1 cast was run using this CTD was run and there was no rosette sampling. Equipment malfunction prevented any further use of the CTD.

##### Preliminary Steps

The cruise summary sheet was completed.

File 2021-038_0002.xmlcon was used to convert the only HEX file.

There was no history available for the pressure, temperature, conductivity and DO sensors since their last factory calibration and this was likely the first use for this equipment since the factory visit.

There were no post-cruise calibrations available.

##### Conversion of Full Files from Raw Data

The deployment method was to soak at the surface, then drop to 5db for about 1 minute. The CTD was then raised to about 2db and after a 2 minute wait the main cast was run. A malfunction occurred at the bottom so no bottle sampling was done. All expected channels are present and the profiles looked reasonable but the Surface PAR is extremely noisy.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

There were quite a few padded values in pressure and depth. This later produced led to padded values in salinity and dissolved oxygen channels.

##### ALIGN DO

Given the lack of upcast data it is impossible to apply the usual test of alignment by comparing DO offset between downcast and upcast with that of temperature. The best setting for this type of sensor is usually about 2.5s to 4s and in previous cruises in the Arctic +3s usually looked best overall. Applying that setting led to a good match in features for T and DO in the downcast.

ALIGNCTD was run using a setting of +3s to the dissolved oxygen channel.

##### CELLTM

The usual tests for the best settings for this routine are difficult to interpret because of the noisy upcast data, so the default settings was used.

CELLTM was run using (α = 0.0245, β = 9.5) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The differences between sensor pairs were:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2021-038-0002 | 180 | +0.0004 | -0.0001 | -0.0020 | High, F.Steady |

Such small differences suggest that all sensors were working well.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header entries were checked against document CTD-summary.xlsx and no errors were found.

The initial pressure readings are very close to 0db; pumps were not on so conductivity does not help in determining if these readings are from right at the surface but transmissivity does suggest the CTD was in water. So there is no apparent error in pressure.

##### CLIP & Header Altimetry Check

##### To remove the records from the initial drop to 5m CLIP was run to remove the first 15580 records.

The header altimeter reading was 7.8db and the maximum depth sampled 194.4. The header entry for water depth was 220m, but in the bridge log it is given as 207m. The sum of maximum depth sampled plus altimetry suggests a value of 202m. While there will be variations during the cast, the 207m entry looks more accurate than the header entry so the header was changed to 207m.

##### Shift

Fluorescence

There is no way to check the alignment of fluorescence since there are no upcast data, but for this type of sensor the usual shift is +24 records.

SHIFT was run on the fluorescence channel using the usual advance of +24 records.

Dissolved Oxygen

Channel Oxygen:Voltage:SBE was aligned earlier. Alignment with features in temperature looks good.

Conductivity

Tests were run using a variety of shifts and the best results were found with the settings -1.3s for the primary and -0.7s for the secondary. SHIFT was run twice using those settings.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors – None available since last factory calibrations.

Historic ranges – Local climatology was not available.

Repeat Casts – None

Post-Cruise Calibration – None available.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited since they looked smoother overall in T-S space than the primary channels.

CTDEDIT was used to clean a few small spikes in salinity that appear to be due to misalignment of T and C and a few records probably corrupted by shed wakes near the top and bottom of the cast.

The output file was named *.EDT.  A T-S plot was examined and the results looked ok.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the EDT file. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Initial Recalibration

Pressure appears to be accurate.

There was no Salinity or Dissolved Oxygen calibration sampling.

CHANGE UNITS was used to derive dissolved oxygen in mass units.

DERIVED QUANTITIES was used to calculate DO surface saturation.

A plot was made of DO surface saturation and the value was about 103%. This is within the range of values found in 2019 in this area. This type of sensor tends to read low rather than high, so there is no indication of large drift in calibration. No recalibration was applied.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run to remove the following channels: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity:T0:C0, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag from all casts

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add some header comments.

The Standards Check routine was run and no problems were found.

A header check was run on the final files and no problems were found.

Plots of each file were examined and no further problems were found.

SECTION B – SBE19+  (Events #23-200, 209, 216-332, 347-391)

##### Seasave

This step was completed at sea. The file names had small error that was corrected.

##### Conversion of Raw Data

The configuration file used at sea was saved as 2021-038-sbe19plus.xmlcon; no errors were found. It was then used to convert the hex files into CNV files and ROS files.

Plots were made of a few CNV files and the data look reasonable. There was a soak of about 2 minutes at 2m after the pumps came on after which the full cast was run.

##### WILDEDIT

This step was skipped as the only spikes are at the surface and will be removed later.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the temperature and conductivity to smooth the high-frequency noise.

The pressure looks smooth and does not require a filter.

##### ALIGNCTD

There was run to advance the DO signal by +3.0 seconds which brought the DO traces into better alignment with the temperature traces.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

Plots show the filters and alignment worked well.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

A spreadsheet was prepared with file names, latitude, longitude, station names and event #s.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output *.MRH.

##### Checking Headers

A preliminary header check was also run and no errors were found.

A track plot was produced and looked fine.

The surface check is not very useful for this type of sampling as there may be records that are in-water but with the pump not on. The average surface pressure was -0.3db. The pressure looks close to zero based on both downcasts and upcasts.

##### SHIFT

Conductivity

Tests were run to find the best shift  to apply to the conductivity channel to reduce unstable features in T-S space. Settings from -2.5 to +1 records were tried; good results were achieved with -2.0 records.

SHIFT was run to apply an advance of -2.0 records to the conductivity channel for all files. Salinity was recalculated.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

Fluorescence

The fluorescence data are very noisy but appear to be aligned with temperature well.

##### DELETE, CLIP and CTDEDIT

CLIP was run to remove the first 360 records; all are in the soak period with pumps likely off.

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: All warnings came from the ends of casts.

The DELETE logs for the 3 CTDs were combined in file delete.log.

All DEL files were examined in CTDEDIT and some editing was applied to 33 casts. This involved removing records corrupted by shed wakes and light cleaning of salinity.

- The files for event #23, 73, 166  had too much data removed in DELETE because the descent rate was very low on average. They were put through DELETE again with the low-drop-rate feature turned off with extension DELX.

- Cast #296 appears to have a problem with plumbing that cleared at about 8m, so salinity points only were removed from that section of the profile. Other channels looked ok.

- There were other cases of a lot of data being removed using the 0.3m/s cut-off but turning off the low-drop-rate feature left many corrupted data, so the cut-off was maintained.

- Cast #332 had 2 pressure spikes that were removed.

- Of the 52 casts, 34 were edited and 18 required no editing.

The edited files were copied to *.EDT.

T-S plots were examined and most looked fine with no noteworthy unstable features.

##### Bottle Comparison

There were salinity samples available for comparison with the SBE911+ salinity.

The ROS files were converted to IOS Header format.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output *.BOT.

All MRH files were plotted to look for outliers in temperature and salinity. There were no major outliers. While there is considerable scatter for some bottle stops, these are shallow cases where variability is likely high.

File ADDSAMP.csv was created and sample numbers were added to it. That file was converted to CST files.

ADD SAMPLE NUMBER was used to add sample numbers to the BOT files to create SAM files, which were averaged on bottle number to create SAMAVG files.

The salinity spreadsheet from the analyst was simplified, saved as a CSV file and converted to individual SAL files for each cast.

MERGE was used to add the salinity samples to the CST files and those files were put through CLEAN and SORT to get them in bottle number order. (MRGCLN1s). Those file were then merged with the SAMAVAG files.

COMPARE was run to compare the bottle salinity in the MRG files with the CTD salinity in the SAM files. Initial runs turned up many problems due to errors in the ADDSAMP file. One bottle had been fired inadvertently and not sampled during event 166. Other samples that were planned but had not been taken were not included in the bottle analysis report, so that was updated by entering pad values for such bottles with flag 9 and an explanation. The Electronic sample log provided good information.

There is a lot of scatter in plots of (CTD Sal – Bottle Sal) versus pressure. Most outliers are in the top 100db with the majority having the CTD reading lower than bottles. This is likely due the incomplete flushing of Niskin bottles in depths with high vertical salinity gradients.

There are a few large outliers with CTD salinity reading high:

- Event 82 – Sample 34 - This outlier is the largest at +1psu and is at the bottom of a shallow cast. Any flushing error would tend to produce low bottle values. Water from about 5m higher has similar values to the sample salinity. There is no evidence of a problem with the collection or analysis so no flag is justified.

- Event 138 – Sample 83 - This outlier is from 81db and the looks like CTD values at 45db. The bottle value is out of line in profile while the full CTD profile shows no significant reversals. The analyst noted problems with losing the connection twice during analysis and having to re-connect the salinometer. Flag 4 was recommended as the difference is 0.53psu, large given no local reversals or large gradients. The analyst agreed to that change.

- Event 322 – Sample 242 – This outlier is from 40db and has a value that looks like CTD salinity at about 20db. No problems were noted in analysis. The bottle is out of line in profile and there is no evidence of mis-sampling  as it does not look like it fits at any other firing level. The CTD data during the 10s window around firing time was quiet. Flag 4 was recommended as the difference is 0.6psu and the value is out of line in profile. The analyst agreed to that change.

The scatter above 100db is large and indicative of incomplete flushing in most cases.

- When data are excluded if above 100m or with standard deviation in the CTD salinity during the 10s window  >0.001psu, the CTD salinity is lower than bottle salinity by an average of 0.0030psu (std dev 0.0057). There were 44 bottles included in the comparison.

- Flushing errors will tend to make the CTD look lower than it really is during upcasts, but the opposite is likely for the bottles fired at the bottom, so the data were separated into 2 groups.

- Using all the same data except excluding bottles fired at the bottom the CTD is low by an average of 0.0038psu (std dev 0.0053psu); that included 30 bottles. Allowing for some error due to incomplete flushing would suggest the CTD is reading low by 0.002psu to 0.003psu.

- Using ONLY bottles fired at the bottom the CTD salinity is low by 0.0013psu (std dev 0.0064psu) with 14 bottles in the comparison. Any flushing error at the bottom would make the CTD look higher than it really is, but the local gradients are quite low for those bottles, so this shouldn’t be a large effect. An estimate that the CTD is reading low by about 0.002psu looks reasonable.

- There may also be a small error due to evaporation and desorption, but this should be small since the analysis was done within 3 to 6 weeks of collection. That could raise bottle salinity by about 0.002psu which then brings the CTD within 0.001 of bottles, an excellent correspondence.

- Given the precision study result of Sp=0.0032psu and the scatter in all the fits, there is no justification for recalibration of the CTD salinity. It is well within expectations for this equipment and errors are likely within ±0.003psu.

##### Other Calibration Checks

Sensor History – There is no history available of this CTD being used since it was last calibrated.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Local climatology was not available.

Post-cruise calibrations – None available.

Dissolved Oxygen Saturation –

After editing a second dissolved oxygen channel was added with mass units.

Dissolved Oxygen saturation was then derived and plotted. Values near the surface ranged from 100% to 115% with most casts between 103% and 107%. These values are higher than observed in 2019 and much higher than in 2018. High dissolved oxygen values are not likely to be caused by calibration drift; these sensors tend to drift towards lower values. The high values and large variability in DO are likely due to local conditions in the region.

##### CALIBRATE

CTD pressure and salinity values appear to be reliable.

Given a large range in dissolved oxygen saturation and no calibration sampling, no recalibration is justified for that variable.

No recalibration was applied.

##### Bin Average and REMOVE

The files were bin averaged using 1db bins.

At this point it was discovered that water depths were missing from the files so the header merge file prepared earlier was adjusted by adding those readings and MERGE CSV file to Headers was rerun.

Checks were then made to see if depths and/or altimetry header entries were in error. A calculation was made of the sum of the maximum depth sampled plus the altimetry header reading minus the water depth minus 1 since altimetry is averaged over the top 2m. There were no altimetry header readings for some casts since the CTD did not get within 15m of the bottom.

Where that “check value” was >4 the result was investigated. For 5 casts altimetry plots and header entries were investigated:

- Event 117 was off by 20m. There was quite a lot of drift in altimetry at the bottom so this is likely a matter of shoaling and the depth may represent the depth at the beginning of the cast. No change was made. The bridge log also has large differences in the 3 readings.

- Events 252 and 261 are only slightly high; the altimetry looks ok. No change was made.

- Event 270 is slightly high but the altimetry is very noisy. No change was made.

- Event 311 had the water depth entered as 161m but sampling went to 275m, so this is clearly wrong. An estimate was made that it should be 285m based on altimetry and sampling depth. The bridge log confirmed this value.

T-S plots were examined and no problems were found.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate, Altimeter and Flag channels from all casts.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add a note to the headers:

CLEAN was run to update the channel limits in the headers.

A header check was run on the CTD files and no errors were found.

##### Final Bottle Files

The MRG files were put through CLEAN to add 0s to empty flag channels and to update headers.

SORT was run to put files in pressure order.

REMOVE was run to remove channels Scan_Number, Oxygen:Voltage, Descent Rate, Altimeter and Flag channels from all casts. This step was run twice because Oxygen;voltage had been converted twice, by accident.

REORDER was run to put channels in usual order.

Header Edit was run to add header comments and fix formats.

Standards check and header checks were run and no problems noted.

SECTION C – RBR (Events #205, 211, 338, 339, 343)

##### Conversion to IOS Headers

Conversion of RBR files to IOS Header format was done by S. Huntington.

Data from 6 casts were found in files:

- 065579_20210812_1421.rsk

- 065579_20210822_2119.rsk

- 065579_20210831_1555.rsk

File “2021-038_CSV_DATA-6Linehdr.csv” was created including header formats and CTD data.

CONVERT Spreadsheet Files was run to produce files with IOS Header format. Header entries of “Administration”, “File” and “Instrument” were filled in this step.

File “2021-038_header-merge.csv” was created, based on the information provided by the chief scientist.

- Column “File_Name”: entries were derived from the event number.

- Column “LOC:LATITUDE”

- Column “LOC:LONGITUDE.

- Column “LOC: Event Number

- Colmun “LOC: STATIONS.

The routine “Merge:CSV Files to headers” was run to add location headers to the IOS files.

CLEAN was run to add a start time to the headers.

All data were plotted and examined. There are spikes in temperature and salinity that should disappear after DELETE is run.

The track plot looks reasonable.

ADTIMECH was run to add a records number to enable plots to determine pressure accuracy. At the end of the record conductivity dropped to 0 when pressure was between 0 and 0.4db, well within the accuracy expected from this instrument.

A gull-wing filter, size 7, was run on pressure, temperature and conductivity.

SHIFT was run on the conductivity channel by -2 records. Salinity was recalculated.

SHIFT was run on the oxygen saturation channel with an advance of +13 records.

DELETE was run to remove swells greater than 2m and records with descent rate lower than 0.3m/s over 5 points. This was not applied in the top or bottom 5m to avoid loss of records as the CTD began and ended its descent.

Profile plots look good. T-S plots show a few small problems best fixed in CTDEDIT.

Dissolved Oxygen saturation values at 2m range from 100% to 102%. During cast #204, which is not to be archived, the value was 105% right at the surface and about 102.5% at 2m.

CTDEDIT was used to clean salinity in most of the files, mostly at the top and bottom of casts.

REMOVE was run to remove channels Record #, Event #, Date and Time:UTC.

REORDER was used to reorganize the channel order so that pressure is first.

CALIBRATE was run to correct the units in conductivity.

There was a single salinity sample from 2m above the CTD during cast #205. The CTD hit bottom during the cast. The sample value was 31.2942psu. The salinity at 37db (2m above bottom) was 31.3034psu and a value close to that of the sample can be found between 32 and 33db. The water in the Niskin bottle likely came from a little higher in the water column due to incomplete flushing so the value does not necessarily indicate poor calibration.

During a previous cast at the same site the CTD salinity at 36db was 31.2948.

There is not enough sampling to justify recalibration and what evidence is available does not suggest a large error.

Casts #204 and 205 are not both needed. Plots were compared and they only differ significantly right at the surface. Cast #205 goes deeper and has a salinity sample to go with it, so it was selected.

Header Edit  was used to fix channel names and formats.

CLEAN was run to reset channel header limits.

Standards check was run and turned up a few problems. Header Edit was rerun until all problems were resolved.

Header Check was run and no problems were found.

FINAL STEPS

Once the data from the 3 CTDs were ready they were all copied to folder “2021-038\Archive\CTD”. Then checks were run on the combined set of files:

- Header Check was run and a few inconsistencies in header entries were found. Changes were made to the RBR headers until all were consistent.

- Track plots were produced and added to the end of the report.

- A cross-reference list was prepared.

Bottle files prepared for the SBE19+ casts were copied to the Archive folder.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1189 |  | No |  | Yes |  |  |  |
| 2 | SEABIRD | 19+ | 6009 |  | Yes |  | Yes |  |  |  |
| 3 | RBR | Concerto | 065579 |  | No |  | Partly |  |  |  |
| Calibration Information - 1189 |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  | 5830 |  | 30Oct2019 |  | Factory |  |  |  |  |
| Conductivity |  | 4327 |  | 26Nov2019 |  | Factory |  |  |  |  |
| Secondary Temp. |  | 5831 |  | 30Oct2019 |  | Factory |  |  |  |  |
| Secondary Cond. |  | 4339 |  | 26Nov2019 |  | Factory |  |  |  |  |
| Transmissometer |  | 993 |  | 20May2020 |  | Laurier |  |  |  |  |
| SBE 43 DO sensor |  | 637 |  | 31Oct2019 |  | Factory |  |  |  |  |
| Optode DO |  | 16 |  | 2018 |  | Factory |  |  |  |  |
| PAR |  | 70501 |  | 4Apr2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  | 2859 |  | 31May2108 |  | Factory |  |  |  |  |
| Surface PAR |  | 20279 |  | 4Apr2016 |  | Factory |  |  |  |  |
| Pressure Sensor |  | 0130015 |  | 23Dec2013 |  | Factory |  |  |  |  |
| Altimeter |  | 41098 |  | Dec2007 |  | Factory |  |  |  |  |
| Turbidity Meter |  | 11074 |  | 23Mar2007 |  | Factory |  |  |  |  |

| Calibration Information - 6009 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 6009 | 22Feb2020 | Factory |  |  |
| Conductivity | 6009 | 22Feb2020 | Factory |  |  |
| SBE 43 DO sensor | 615 | ? | Factory |  |  |
| PAR | 70501 | 4Apr2016 | Factory |  |  |
| SeaPoint Fluor. | 2745 | 8Apr2021 | Factory |  |  |
| Pressure Sensor | 6009 | 22Feb2020 | Factory |  |  |
| Altimeter | 1186 | 28May2014 | Factory |  |  |
| Turbidity Meter | 11432 | Oct2012 | Factory |  |  |

| Calibration Information – RBR Concerto |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 65578 | 13April2021 | Factory |  |  |
| Conductivity | 65578 | 13April2021 | Factory |  |  |
| Fluorometer | ? | ? |  |  |  |
| Rinko DO | ? | ? |  |  |  |
| Pressure Sensor | 65578 | 13April2021 | Factory |  |  |