## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2021-071

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B.C.

Party Chief: Boldt J.				Location: Strait of Georgia

Project: Juvenile Herring			Platform: Walker Rock

Date: 8 September 2021 – 24 September 2021

Processed by: Germaine Gatien

Date of Processing: 12 May 2022 – 24 May 2022

Number of original HEX files: 20			Number of CTD files: 12

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 8118) was used for all casts. There were no external sensors attached.

# SUMMARY OF QUALITY AND CONCERNS

A spreadsheet with header information was provided which included start times and positions.

The file names were non-standard, but included information that made it easy to change them to standard format.

8 of the 10 casts in the second leg had no useful data. The likely cause was a low battery.

The CTD was soaked at about 1db. The recommended protocol calls for the CTD to be lowered to 10db to soak, and then raised to the surface. After a short wait at the surface to allow the disturbance caused by the rising CTD to dissipate,  the complete cast can be run. This is done because the pressure at 10db forces bubbles to clear from the plumbing and for sensors to then equilibrate. If that method is not practical, then a wait at the surface or at least 30 to 60s after the pump comes on is recommended. The pump for this cruise was set to come on after 60s and after a minimum conductivity was reached.

There was no calibration sampling for salinity. The calibrations were from February 2021, so little calibration drift is expected.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were non-standard.

The raw files were copied to another folder where they were renamed in standard format.

##### Preliminary Steps

A spreadsheet with positions and comments for casts was provided.

The event numbers go from 1 to 20.

The cruise summary sheet was completed.

A single configuration file used for all casts.

There were no parameters on file at IOS for the temperature, conductivity and pressure calibrations, so the values could not be checked.

There was no history available for this CTD.

The configuration file was saved as file 2021-071-ctd.xmlcon.

##### Conversion of Raw Data

All casts were converted using file 2021-071-ctd.xmlcon.

The CTD was usually soaked at about 1m. The CTD pump was set to turn on 60s after starting and only if conductivity was above a minimum frequency. The conductivity sensor requires some time to equilibrate after that. Full equilibration may not be achieved before the cast starts if there is no 10m soak as increased pressure clears the plumbing.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the temperature and conductivity to force it to have the same response as the temperature. Pressure is very smooth, so no filter was applied..

##### ALIGNCTD

Temperature and conductivity were advanced by +0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Plots of temperature and salinity versus pressure and temperature versus salinity were examined to see that data look reasonable and they do for the first 10 files. For the second group, files are incomplete, with many containing only surface data. Cast #12 contains data to 11db while #16 appears to have the full downcast.

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth. Spreadsheet “Boldt_JuvHerringSOG_CTD_2021.xlsx” provided that information.

The file was saved as 2021-071-header-mrg.csv and position formats were adjusted to ones suitable for IOS SHELL.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH including the start time for each cast.

The start times from the files differ from those in the spreadsheet by up to 23 minutes; the differences seem too random to be caused by the CTD clock being off, though in most cases the log time is earlier and when it is later it is by only a few minutes. The times in the spreadsheet are likely based on arrival on site, rather than the beginning of the CTD cast, so the times from the instrument will be used.

Times are local (evening sampling for all casts), so 7 hours need to be added to get UTC.

A cross-reference listing was produced and one problem was found and corrected.

A preliminary header check was also run and no further errors were found.

Plots show that the CTD obviously moved from air to water between pressures of -0.08db and +0.05db. The ends of casts had a minimum pressure of -0.09db. These are excellent results especially given that the initial accuracy quoted for this type of pressure gauge is ±1db.

Track plots look good and were added to the end of this report. They include the sites with no useful CTD data as they may be useful for other types of sampling done.

##### CLIP

There are reversals in salinity at the top of some casts as the pump comes on and equilibration starts. There are also a few pressure reversals.

DELETE can remove many such unstable features of instrumental origin, but not all.

Running CLIP to remove the top 1m after running DELETE is often a good approach to handle such data, but in this case it didn’t work well on some casts where downcasts started before the system equilibrated and sometimes a pressure reversal complicated the picture.

So CLIP was run before DELETE. This required some fine-tuning to pick the number of records to remove for each cast; the data files had to be examined in detail for a few to pick an appropriate starting point. For cast #8 the downcast began before the pump came on, so more data than usual were removed.

##### SHIFT

Conductivity

Small shifts to conductivity were tested to see if they improved the salinity, as judged by T-S plots. A setting of -0.5 was applied to conductivity and salinity was recalculated.

##### DELETE and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

All DEL files were copied to *.EDT

T-S plots were examined and graphical editing was found to be necessary for only cast #16. The output was copied to *.EDT.

##### Inter-comparisons

Sensor History – None available – new CTD.

Comparison of repeat casts –The sites are too shallow for repeats to be helpful.

Historic Ranges – For a few casts there were excursions from the climatology with higher temperature and lower salinity around 10 to 15m. Near-surface outliers are common in this area. The climatology is not well suited to near-shore sampling. These excursions are not systematic and are not considered indicative of a calibration problem.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling and no useful history to enable estimates of calibration drift. No recalibration was applied.

##### Bin Average and REMOVE

The files were bin-averaged using 1db bins.

T-S plots were examined. No problems were found.

REMOVE was run to remove Scan_Number, Descent Rate and Flag channels.

ADD TIME CHANNEL was used to get the time in UTC by adding 7 hours.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add some comments to the headers:

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2021-071 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 8 September 2021 End: 24 September 2021 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Chief Scientist: Boldt J. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 8118 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE19+/8118  Cruise ID#:		2021-071

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 8118 | 18Feb2021 | Factory |  |  |
| Conductivity | 8118 | 18Feb2021 | Factory |  |  |
| Pressure | 8118 | 21Feb2021 | Factory |  |  |

Cross-reference list – does not include casts with only surface data