## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 10 June 2022 | Recalibrated CTD Fluorescence, entered corrected silicate values and corrected nutrient formats. See end of report for details. |

## PROCESSING NOTES

Cruise: 2021-153

Agency: ESD, Ecosystem Sciences Division, West Vancouver BC

Chief Scientist: Sutherland T.  		Platform: Vector

Location: Jervis Inlet			Project: HAB - Jervis Inlet

Date: 31August 2021 –6 September 2021

Processed by: Germaine Gatien

Date of Processing: 27 April 2022 –  11 May 2022

Number of original HEX files: 40

Number of CTD files: 40			Number of BOT files: 30

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 1091) was used with temperature sensor #6376, conductivity sensor #4799, Wetlabs ECO Fluorometer #2216, dissolved oxygen sensor #1483 and pressure sensor 1091.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order with comments about problems encountered.

Header information (station names and positions) was not entered in the raw files, but an electronic log was provided that contained the missing details. This was very helpful.

Samples were collected using 3 different methods. During some of the CTD casts a Niskin bottle was mounted 2.5m above the CTD. At the same sites surface samples were taken using a bucket. There was also another cast with bottles fastened to a wire to collect samples from 5m, 10m, 15m, 25m and 50m.

The exception was event #7 when all bottles were attached during the CTD cast in order to save time. This involved many stops during the downcast, so the CTD data from event #7 are considered of lower quality than for other casts.

BOT files were prepared that include salinity and nutrient samples plus CTD data extracted from downcast CTD files averaged over 0.5m at the nominal sampling depths. The event numbers for these files come from the CTD casts, but include the surface bottle and hydro wire samples as well.

No recalibration was applied to the CTD data. Pressure looked excellent. There was no dissolved oxygen sampling. There were salinity samples but analysis was delayed and many bottles had salt on lids and inserts. Without a rosette, matching CTD data to bottles is also uncertain. The comparison was extremely noisy. The CTD had been serviced recently so calibration errors are expected to be small. The fluorometer had not been serviced recently and deep values seem high, ~0.6ug/L, but inlet waters do often have quite high fluorescence values at depth.

Extracted chlorophyll sample data were not available at the time of processing.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea.

2. Preliminary Steps

The Daily Log and a spreadsheet of sampling done were obtained.

The cruise summary sheet was completed.

Typical sampling scheme: The CTD was soaked briefly at the surface and then lowered to 10m for a soak of between 1 and 2minutes. It was bought back to surface to sit briefly. Then the full cast was run. Typically this started at about 2 to 2.5minutes after acquisition began.

For casts with bottom sampling a 1.7L Niskin bottle was mounted ~2.5m above the CTD.

Next, a second cast was run with bottles on a wire and a surface bucket sample was taken.

To save time during cast #7 all bottles were mounted during the downcast so that all sampling could be done when the CTD was at the bottom. This leads to frequent stops so the quality of the CTD data is lower than for other casts during this cruise.

## 3. Conversion of Raw Data

The configuration file used at sea was correct. It was saved as 2021-153-ctd.xmlcon and used to convert all HEX files.

An initial conversion showed that surface pressure looks good, with conductivity falling to ~0 when pressure is between -0.2db and +0.2db. The deep fluorescence values (~600db) are about 0.6ug/L; this seems high, but few casts were found in the archive for comparison, especially ones from this time of year. Cast 2021-153-0058 is fairly close to cast 2019-053-0109 from September 2019. The earlier cast had a fluorescence value of 0.33ug/L at 200db compared to 0.6ug/L for the current cruise, but it also had significantly lower dissolved oxygen values than found in 2021. Values in inlets are often higher at depth than seen offshore.

All expected variables are present.

There is the usual steppy pressure which leads to unusual profiles.

4. WILDEDIT

WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 32

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

The routine removed spikes effectively.

5. FILTER

Usually data from SBE25 CTDs are put through the FILTER routine at this stage, rather than leaving it until the DELETE stage as is done for SBE911 data. However, for this cruise the sampling rate was higher than usual, at 0.0625s, so this step did not appear necessary.

Tests were run to see if the usual approach of applying a cosine filter to temperature and conductivity using routine WFILTER did a good job of removing small reversals. Filter size 5 to 21 were tested and size 19 removed most problems. This is a larger size than usual, but the sampling rate is 16Hz compared to 8Hz from most cruises. WFILTER was run on all casts using size 19.

6. ALIGNCTD

Based on tests run for other cruises in this project using the same DO sensor, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results look excellent.

7. CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). Plots were examined and confirmed that steps 5, 6 and 7 had improved the data.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values. Fluorescence values <0 were replaced with pad values; while most were in the soak period, some may be spikes in deep water.

Positions and station names were not entered in the headers, but were in a digital log. The data were reformatted in spreadsheet 2021-153-Header-merge.csv. The start times were in the converted files, but for cast #1 the time was changed to match the digital log since it is noted in the Daily Science log that the CTD was turned on long before the cast started. Other times were close to those in the digital log.

Merge CSV File to Headers was run.

The next step is to remove the data collected during soaks at 10m. A list was made of event numbers and the number of records that need to be removed based on plots of pressure versus scan #. CLIP was run using that list and plots were again examined to ensure the best choice was made. Values were fine-tuned until all plots showed good results.

10. Checking Headers

A cross-reference list was produced and all looked fine.

Track plots looked ok so were added to the end of this report.

Header Check was run. No problems were found.

The surface check shows an average of 0.004db before CLIP was run. The range was from -0.2 to +0.2db.

Plots were made to find the pressure at which conductivity dropped to near 0 at the ends of casts and values varied from -0.2db to +0.2db, with most close to 0. Downcast data also showed a sudden rise in conductivity near 0, typically at about 0.1db.

The file for event #113 is very big with a lot of data at the end at the surface. It is likely the CTD was either on the surface or on the boat. In either case the pressures are very close to zero (-0.1db to +0.1db)  for almost 1 hour, indicating that the pressure was very accurate.

##### 11. SHIFT

Conductivity

Tests were run to see what shift to conductivity made the best improvement to stability in T-S space. Shifts of +1.2 records looked best and that was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is ok.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. No further adjustment was made.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width:    11

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 9 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

All DEL files were copied to *.EDT so there will be a complete set of files even if some need no editing.

CTDEDIT was used to remove records corrupted by shed wakes and to clean salinity where unstable features looked likely to be caused by misalignment of T and C.

24 casts were edited, mostly near the surface and bottom.

Notes of editing details were made in the headers.

T-S plots were examined after this step. No significant unstable features were found.

##### 14. Initial Bottle Data Steps

There was no rosette available for this cruise.

Bottles samples were taken at the bottom during CTD casts, then a separate cast was done to get bottles at 5, 10, 15, 25, 50m and bucket sample was taken for the surface value. 3 separate event numbers were used for the 3 types of sampling. For event #4 all bottles were fired during the CTD cast except for the bucket sample.

A spreadsheet was prepared with the nutrient and salinity samples. CHL was not yet available. A single event number based on the CTD cast was entered for all sampling. 2021-153_bottle_data_6linehdr.csv

The only depths available for samples are nominal. To get bottom depths the maximum depth sampled was obtained. The bottles were closed at ~2.5m above the bottom.

Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2021-153-bot-hdr.txt which will be updated as needed during processing.

- Extracted CHL sampling was done but the data were not yet available.

- Salinity analysis was obtained in spreadsheet QF 2021-153 SAL*.xlsx.

- Nutrient analysis was obtained in spreadsheet QF2021-153 nuts*.xlsx.

A workbook was created. Data from the QF worksheets were added to one sheet of the workbook. A 6-line header was prepared using entries from the event log relevant to Niskin sampling plus bottle data. Space was included for the addition of CTD data.

For this cruise samples came from 0, 5, 10, 15, 25, 50 and bottom – 2.5m.

There are no CTD data right at the surface. For most casts data were found within 1.5m of the surface; those were included in the files. For the bottom samples data were found at the maximum depth – 2.5m. To obtain those data the EDT files were binned to 0.5db and then thinned to 0, 5, 10, 15, 25, 50m and the levels of the bottom samples (in order of maximum depth). They were extracted to a spreadsheet and the inappropriate deep entries were removed to get a match with the bottle depths. Those data were added to the 6-line header file and saved as file 2021-153-CTD_plus_Bottle_data_6linehdr.csv.

The same data were also saved in file 2021-153-Bottle_CTD_comparison.xlsx..

Some data files (events 38-41)are missing because they were run from a small boat with no CTD sampling.

##### 15. Compare

Salinity Comparison

Salinity data from bottles were compared with CTD salinity from bin-averaged files using estimates of sampling depth.

It was discovered that samples 83B and 89 had been reversed. This was fixed by the analyst in the salinity spreadsheet. It was then fixed in the MRH file.

When all data are included except for one record with flag 4, the CTD is higher than the bottles by an average of 0.141psu (std dev 0.352psu). When only bottles fired below 50m are included the CTD is low by an average of 0.014psu (std. dev. 0.008psu).

A plot of differences versus nominal depth looks random for the deep samples.

There are several potential sources of error:

- For all samples, analysis was delayed by 11-12 weeks, which would be expected to raise salinity in the bottles by about 0.002psu due to desorption of glass.

- Evaporation is also expected due to the delay and especially so in this case since the analyst noted some screw tops were loose and some samples had salty caps and inserts. Some bottle samples are likely reading significantly high, making CTD data look much lower than it really is.

- Errors due to incomplete flushing of Niskin bottles are likely to mean the salinity in the bottles would be lower than ambient waters. This may partly offset the errors due to evaporation.

- For the surface samples the CTD data came from lower in the water column that the bottle samples, so we expect the CTD to read higher than bottles if calibration is perfect.

- For all samples the depths are nominal – CTD data were selected based on nominal depths and may not be a great match, though the errors are likely random in sign.

The issues of evaporation and desorption and possibly having salt fall into samples would explain the CTD reading lower than bottles. The randomness in the fit could be more evidence of evaporation, but it could also be due to errors in depth estimates.

16. Other calibration checks

Sensor History – There is no history of use of this CTD since it was last serviced at the factory.

Historic Ranges – There was local climatology available for only 1 cast. Both salinity and temperature fell within the climatology for that cast.

Post-cruise calibrations – None were available.

17. CALIBRATE

Pressure was found to be very accurate.

The comparison of CTD salinity with bottles was not considered reliable enough to justify recalibration.

There was no dissolved oxygen sampling. Surface saturation rates ranged from 80% to 130% with only 2 casts <96%. This type of sensor tends to drift low, but there is no evidence to suggest a large error.

No recalibration was applied.

18. Fluorescence Filter

The fluorescence data was not filtered as it is not very spiky.

19. Bin Average, Remove, Derive DO in mass units, Reorder

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show a range of 80% to 130%.

REORDER was used to get the 2 dissolved oxygen channels together.

20. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add comments about processing.

A cross-reference listing was produced and may be found at the end of this report..

A header check and standards check were run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

T-S plots from Jervis Inlet fall into a tight group.

##### 20. Final BOT file preparation

File 2021-153_bottle_data_6linehdr was converted into individual IOS files.

Some elements were missing (Start Time, Water Depth) from the headers, so were exported from the CTD files. The Merge CSV files to Headers routine was used to add them to the IOS files (Output:MRG).

CLEAN was run to enter 0 flags where the flag channels are empty.

Change Units was used to add CTD dissolved oxygen in mass units.

Reorder was used to get the 2 CTD DO channels together.

HEADEDIT was used to add comments.

The final files have extensions BOT.

Profile plots were examined to ensure no data were lost or out of place. A problem was found in event #70 due to an overflow value. That was corrected and the resulting file was ok.

T-S plots were examined and no problems were found.

The standards check was run and no errors were found.

The track plot looks fine.

A cross-reference list and header check were run on the BOT files and no problems were found.

CORRECTIONS – June 10, 2022

- The nutrient formats were corrected in the 6-line header.

- Silicate corrections where salinity <25psu were received from analyst (QF2021-153_Nuts_9Jun2022.xlsx). New values were entered in the 6-line header.

- The bottle files were reassembled and header comments updated.

- Fluorescence values recalibrated in CTD and BOT files since the instrument was found to have been set to 12, not 24, so the scale factor and offset were incorrect.

FL (corrected) = FL(uncorrected)*0.5 – 0.192

PARTICULARS –

General – problems with sounder.

- CTD turned on long before cast began.

7.  CTD and bottles combined on wire to save time.

12. Second bucket taken because SAL not taken.

38-41. Samples taken from small boat.

79-101. Time series JI2.

102-125. Time series JI6C.

105. Consistent current at ~10m

113/114. CTD left running between casts.

CRUISE SUMMARY

| Cruise ID#: 2021-153 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 31 August 2021 End: 6 September 2021 |  |  |  |  |  |  |
| Location: Jervis Inlet |  |  |  |  |  |  |
| Chief Scientist: Sutherland T. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 1091 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/1091	Cruise ID#:		2021-153

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 6376 | 1Mar2021 | Factory |  |  |
| Conductivity | 4799 | 5Mar2021 | Factory |  |  |
| ECO Fluorometer | 2216 | 8Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1483 | 5Mar2021 | Factory |  |  |
| Pressure | 1091 | 10Mar2021 | Factory |  |  |