## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2022/10/03 15:13:53 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2022-008-0018.mrg | 2022-008-0018.sam | 242 | 1249.45 | 1249.41 | 0.2245 | -0.047607 | 34.4606 | 34.4592 | 0.0002493 | -0.0014496 | 34.4606 | 34.4595 | 0.00023282 | -0.0011063
2022-008-0018.mrg | 2022-008-0018.sam | 242 | 1249.36 | 1249.41 | 0.2245 | 0.047363 | 34.4606 | 34.4592 | 0.0002493 | -0.0014496 | 34.4606 | 34.4595 | 0.00023282 | -0.0011063
2022-008-0018.mrg | 2022-008-0018.sam | 121 | 1000.39 | 1000.39 | 0.087997 | -0.00012207 | 34.3733 | 34.3692 | 0.00026172 | -0.0041122 | 34.3733 | 34.3692 | 0.0002916 | -0.0040703
2022-008-0018.mrg | 2022-008-0018.sam | 121 | 800.385 | 800.385 | 0.045273 | 0.00036621 | 34.2732 | 34.27 | 0.00025105 | -0.0032005 | 34.2732 | 34.27 | 0.00027095 | -0.0032158
2022-008-0018.mrg | 2022-008-0018.sam | 121 | 4.943 | 4.9435 | 0.088551 | 0.0004878 | 31.7869 | 31.7774 | 0.00029579 | -0.0095119 | 31.7869 | 31.7775 | 0.00030083 | -0.0094013
2022-008-0027.mrg | 2022-008-0027.sam | 121 | 5.417 | 5.4165 | 0.071005 | -0.00048733 | 31.8928 | 31.8866 | 0.0079538 | -0.0062408 | 31.8928 | 31.8869 | 0.0084602 | -0.0059318
2022-008-0029.mrg | 2022-008-0029.sam | 121 | 5.087 | 5.0867 | 0.085704 | -0.00026417 | 31.9155 | 31.899 | 0.0017521 | -0.01651 | 31.9155 | 31.8994 | 0.0023966 | -0.01609
2022-008-0033.mrg | 2022-008-0033.sam | 121 | 2000.61 | 2000.61 | 0.30794 | -0.00036621 | 34.5859 | 34.5842 | 0.00027139 | -0.0017281 | 34.5859 | 34.5848 | 0.00022745 | -0.0011444
2022-008-0033.mrg | 2022-008-0033.sam | 121 | 1500.34 | 1500.35 | 0.14315 | 0.00048828 | 34.5125 | 34.5107 | 0.00034639 | -0.0017929 | 34.5125 | 34.511 | 0.00025369 | -0.0015106
2022-008-0033.mrg | 2022-008-0033.sam | 121 | 1250.43 | 1250.43 | 0.15121 | -0.00048828 | 34.4505 | 34.4474 | 0.00025759 | -0.0030632 | 34.4505 | 34.4477 | 0.00027648 | -0.0028496
2022-008-0033.mrg | 2022-008-0033.sam | 121 | 1000.36 | 1000.36 | 0.044304 | -0.00024414 | 34.4038 | 34.3983 | 0.00026999 | -0.0054588 | 34.4038 | 34.3984 | 0.00028179 | -0.0053825
2022-008-0033.mrg | 2022-008-0033.sam | 121 | 800.791 | 800.791 | 0.10637 | 6.1035e-05 | 34.305 | 34.3025 | 0.00037756 | -0.0025368 | 34.305 | 34.3022 | 0.00035346 | -0.0027504
2022-008-0034.mrg | 2022-008-0034.sam | 121 | 5.165 | 5.1646 | 0.21953 | -0.00040483 | 31.8277 | 31.8274 | 0.00043663 | -0.00034714 | 31.8277 | 31.8274 | 0.00049417 | -0.00027084
2022-008-0038.mrg | 2022-008-0038.sam | 124 | 2998.4 | 2998.4 | 0.5898 | 0.00024414 | 34.6462 | 34.6454 | 0.00027236 | -0.00078964 | 34.6462 | 34.6461 | 0.00027567 | -0.00014496
2022-008-0038.mrg | 2022-008-0038.sam | 145 | 2999.44 | 2999.49 | 0.47938 | 0.049805 | 34.6461 | 34.6454 | 0.00025291 | -0.00067902 | 34.6461 | 34.646 | 0.0002608 | -0.00010681
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 2500.32 | 2500.32 | 0.53152 | -0.00024414 | 34.6238 | 34.6235 | 0.00028436 | -0.00034714 | 34.6238 | 34.6241 | 0.00024004 | 0.00028229
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 1999.72 | 1999.72 | 0.2528 | 0.00036621 | 34.5887 | 34.588 | 0.00025896 | -0.0007019 | 34.5887 | 34.5886 | 0.00024436 | -0.00014877
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 1500.84 | 1500.84 | 0.16205 | -0.00012207 | 34.5112 | 34.5091 | 0.000263 | -0.0020676 | 34.5112 | 34.5095 | 0.000232 | -0.0017471
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 1251.02 | 1251.02 | 0.51945 | -0.00024414 | 34.4545 | 34.4529 | 0.00025925 | -0.0015984 | 34.4545 | 34.4531 | 0.00025644 | -0.0014305
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 999.451 | 999.451 | 0.40722 | -0.00036621 | 34.3685 | 34.3676 | 0.00023208 | -0.00087738 | 34.3685 | 34.3677 | 0.00025206 | -0.0007782
2022-008-0038.mrg | 2022-008-0038.sam | 121 | 5.201 | 5.2009 | 0.2842 | -0.00012398 | 32.1456 | 32.1462 | 0.00023974 | 0.00055695 | 32.1456 | 32.1464 | 0.00020709 | 0.00080109
2022-008-0048.mrg | 2022-008-0048.sam | 121 | 4.804 | 4.8037 | 0.13168 | -0.00031376 | 32.2205 | 32.2203 | 0.00023888 | -0.00019074 | 32.2205 | 32.2199 | 0.00034732 | -0.00056076
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 3359.02 | 3359.02 | 0.32148 | 0 | 34.6563 | 34.6553 | 0.00027164 | -0.0010223 | 34.6563 | 34.656 | 0.00023636 | -0.00032043
2022-008-0049.mrg | 2022-008-0049.sam | 105 | 2999.53 | 2999.35 | 0.50689 | -0.17578 | 34.649 | 34.65 | 0.0002739 | 0.00099564 | 34.649 | 34.6508 | 0.00025676 | 0.0017662
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 2500.22 | 2500.22 | 0.42423 | 0.00024414 | 34.6302 | 34.6292 | 0.00029204 | -0.0010376 | 34.6302 | 34.6298 | 0.00026352 | -0.00038528
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 2249.12 | 2249.12 | 0.35358 | 0 | 34.6083 | 34.6088 | 0.00025415 | 0.00048828 | 34.6083 | 34.6094 | 0.00025247 | 0.0011177
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 2000.64 | 2000.63 | 0.32565 | -0.00036621 | 34.5856 | 34.584 | 0.00026102 | -0.0016022 | 34.5856 | 34.5846 | 0.00022898 | -0.00099564
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 1749.83 | 1749.83 | 0.23974 | 0.00024414 | 34.5551 | 34.5524 | 0.00026912 | -0.0026932 | 34.5551 | 34.5529 | 0.00024041 | -0.0022469
2022-008-0049.mrg | 2022-008-0049.sam | 86 | 1499.7 | 1499.36 | 0.59493 | -0.3407 | 34.5078 | 34.5063 | 0.00038571 | -0.0014954 | 34.5078 | 34.5065 | 0.00028308 | -0.0012665
2022-008-0049.mrg | 2022-008-0049.sam | 98 | 1249.59 | 1249.27 | 0.57865 | -0.31946 | 34.4518 | 34.4506 | 0.00033206 | -0.0012093 | 34.4518 | 34.4507 | 0.0002659 | -0.0010643
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 1000.65 | 1000.65 | 0.22555 | 0.00012207 | 34.3704 | 34.3688 | 0.00020316 | -0.0015793 | 34.3704 | 34.369 | 0.00022357 | -0.0013771
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 799.545 | 799.545 | 0.16498 | -0.00036621 | 34.2821 | 34.2813 | 0.00026137 | -0.0007515 | 34.2821 | 34.2814 | 0.0002601 | -0.00071716
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 600.327 | 600.327 | 0.26653 | 0.00018311 | 34.1459 | 34.1431 | 0.00028963 | -0.002842 | 34.1459 | 34.1428 | 0.0002843 | -0.0031052
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 400.437 | 400.437 | 0.33491 | -0.00012207 | 33.9404 | 33.9404 | 0.00025309 | 2.6703e-05 | 33.9404 | 33.9404 | 0.0002818 | -1.5259e-05
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 299.392 | 299.392 | 0.30551 | 0.00021362 | 33.8287 | 33.8284 | 0.00047768 | -0.0002861 | 33.8287 | 33.8284 | 0.00056987 | -0.00029755
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 249.804 | 249.804 | 0.40507 | 0.00041199 | 33.8068 | 33.8061 | 0.00041418 | -0.00071335 | 33.8068 | 33.806 | 0.00026868 | -0.00081635
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 200.768 | 200.768 | 0.17197 | -0.00041199 | 33.771 | 33.7699 | 0.00047069 | -0.0010529 | 33.771 | 33.7698 | 0.00037694 | -0.0012169
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 174.785 | 174.785 | 0.15757 | -0.00039673 | 33.7313 | 33.7305 | 0.00065778 | -0.00081253 | 33.7313 | 33.7307 | 0.0007299 | -0.00057983
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 150.431 | 150.431 | 0.23862 | -0.00016785 | 33.6075 | 33.5993 | 0.00094877 | -0.0081558 | 33.6075 | 33.5991 | 0.0012477 | -0.0084229
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 125.193 | 125.193 | 0.22542 | -0.00026703 | 33.1641 | 33.1446 | 0.0015435 | -0.019547 | 33.1641 | 33.1429 | 0.0015026 | -0.021229
2022-008-0049.mrg | 2022-008-0049.sam | 120 | 99.57 | 99.5611 | 0.49643 | -0.0089264 | 32.631 | 32.6081 | 0.0013818 | -0.022873 | 32.631 | 32.608 | 0.0014828 | -0.023014
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 75.415 | 75.4152 | 0.24145 | 0.00017548 | 32.3471 | 32.3448 | 0.00044928 | -0.0023155 | 32.3471 | 32.3446 | 0.00037288 | -0.0025291
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 50 | 49.9996 | 0.31063 | -0.00040436 | 32.3001 | 32.301 | 0.0032722 | 0.00089264 | 32.3001 | 32.3012 | 0.0033732 | 0.0011482
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 24.464 | 24.4641 | 0.29817 | 0.00012398 | 32.2206 | 32.2205 | 0.00091053 | -8.7738e-05 | 32.2206 | 32.2203 | 0.0010984 | -0.00033951
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 10.667 | 10.6675 | 0.12751 | 0.0004549 | 32.2161 | 32.2158 | 0.00018787 | -0.00030899 | 32.2161 | 32.2158 | 0.00019545 | -0.00027466
2022-008-0049.mrg | 2022-008-0049.sam | 121 | 4.831 | 4.8307 | 0.28724 | -0.00026417 | 32.2174 | 32.2161 | 0.0015046 | -0.0013008 | 32.2174 | 32.2165 | 0.0028287 | -0.00087357
2022-008-0051.mrg | 2022-008-0051.sam | 121 | 5.214 | 5.2141 | 0.16979 | 0.0001154 | 32.148 | 32.1473 | 0.0015834 | -0.0006752 | 32.148 | 32.1472 | 0.0015668 | -0.00079727
2022-008-0057.mrg | 2022-008-0057.sam | 242 | 3500.14 | 3500.17 | 0.2377 | 0.037109 | 34.6651 | 34.6639 | 0.00023403 | -0.0011749 | 34.6651 | 34.6645 | 0.00025192 | -0.00060654
2022-008-0057.mrg | 2022-008-0057.sam | 242 | 3500.21 | 3500.17 | 0.2377 | -0.036865 | 34.6651 | 34.6639 | 0.00023403 | -0.0011749 | 34.6651 | 34.6645 | 0.00025192 | -0.00060654
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 3001.09 | 3001.09 | 0.27807 | 0 | 34.6522 | 34.6524 | 0.00027319 | 0.00022888 | 34.6522 | 34.6531 | 0.00023742 | 0.00094223
2022-008-0057.mrg | 2022-008-0057.sam | 119 | 2500.35 | 2500.37 | 0.54344 | 0.016846 | 34.6246 | 34.623 | 0.00031571 | -0.0015602 | 34.6246 | 34.6236 | 0.00024065 | -0.0010338
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 1999.65 | 1999.65 | 0.12449 | -0.00024414 | 34.5778 | 34.5747 | 0.00022288 | -0.0030785 | 34.5778 | 34.5752 | 0.0002517 | -0.0025864
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 1499.41 | 1499.41 | 0.33076 | -0.00036621 | 34.4935 | 34.4919 | 0.00022583 | -0.0015526 | 34.4935 | 34.4923 | 0.00028475 | -0.0012093
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 1250.27 | 1250.27 | 0.2577 | 0.00036621 | 34.4362 | 34.4323 | 0.00023651 | -0.0038948 | 34.4362 | 34.4325 | 0.00024183 | -0.0036926
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 1000.01 | 1000.01 | 0.23737 | 0.00012207 | 34.3504 | 34.3488 | 0.00028203 | -0.0016365 | 34.3504 | 34.3489 | 0.00026351 | -0.0014572
2022-008-0057.mrg | 2022-008-0057.sam | 121 | 800.236 | 800.236 | 0.25743 | 0.00030518 | 34.2664 | 34.2653 | 0.00025735 | -0.0011215 | 34.2664 | 34.2654 | 0.00023217 | -0.00098801
2022-008-0061.mrg | 2022-008-0061.sam | 121 | 5.11 | 5.1104 | 0.29657 | 0.00039673 | 32.3047 | 32.3047 | 0.00022224 | 0 | 32.3047 | 32.3045 | 0.00022018 | -0.00022888
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 3500.36 | 3500.36 | 0.31707 | -0.00024414 | 34.6659 | 34.6663 | 0.00026718 | 0.00039673 | 34.6659 | 34.6668 | 0.0002296 | 0.00086975
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 3001.04 | 3001.04 | 0.29299 | 0.00024414 | 34.6539 | 34.6531 | 0.00024095 | -0.00079346 | 34.6539 | 34.6537 | 0.00026184 | -0.00019074
2022-008-0062.mrg | 2022-008-0062.sam | 69 | 2499.91 | 2499.79 | 0.65889 | -0.11475 | 34.6302 | 34.6265 | 0.00026978 | -0.0037079 | 34.6302 | 34.6269 | 0.00023168 | -0.0032844
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 2000.06 | 2000.06 | 0.13585 | -0.00024414 | 34.5815 | 34.5801 | 0.00028982 | -0.0013695 | 34.5815 | 34.5806 | 0.00027733 | -0.00092316
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 1750.73 | 1750.73 | 0.33847 | 0.00036621 | 34.5406 | 34.5387 | 0.00024204 | -0.0018997 | 34.5406 | 34.539 | 0.00024554 | -0.0016327
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 1500.78 | 1500.78 | 0.33273 | -0.00012207 | 34.4941 | 34.493 | 0.0002886 | -0.001049 | 34.4941 | 34.4934 | 0.00024538 | -0.0006752
2022-008-0062.mrg | 2022-008-0062.sam | 96 | 1250.68 | 1250.34 | 0.59128 | -0.33752 | 34.4366 | 34.4358 | 0.00032793 | -0.00078583 | 34.4366 | 34.436 | 0.00029685 | -0.0006218
2022-008-0062.mrg | 2022-008-0062.sam | 121 | 1000.73 | 1000.73 | 0.5504 | 0.00042725 | 34.3605 | 34.3599 | 0.00023286 | -0.0006218 | 34.3605 | 34.36 | 0.00021891 | -0.0005455
2022-008-0064.mrg | 2022-008-0064.sam | 121 | 5.285 | 5.2854 | 0.30296 | 0.00038052 | 32.2891 | 32.288 | 0.00028753 | -0.0010872 | 32.2891 | 32.2877 | 0.00028585 | -0.0014267
2022-008-0067.mrg | 2022-008-0067.sam | 239 | 3499.02 | 3498.78 | 0.27308 | -0.24023 | 34.6689 | 34.6668 | 0.00026182 | -0.0021439 | 34.6689 | 34.6672 | 0.00023424 | -0.0017242
2022-008-0067.mrg | 2022-008-0067.sam | 239 | 3498.55 | 3498.78 | 0.27308 | 0.23364 | 34.6676 | 34.6668 | 0.00026182 | -0.00084305 | 34.6676 | 34.6672 | 0.00023424 | -0.00042343
2022-008-0067.mrg | 2022-008-0067.sam | 121 | 2999.64 | 2999.64 | 0.53419 | 0.00024414 | 34.6523 | 34.6522 | 0.00023523 | -6.485e-05 | 34.6523 | 34.6527 | 0.00022954 | 0.00040054
2022-008-0067.mrg | 2022-008-0067.sam | 121 | 2500.15 | 2500.15 | 0.4311 | -0.00024414 | 34.6242 | 34.6231 | 0.00027794 | -0.0010567 | 34.6242 | 34.6235 | 0.0002254 | -0.00068283
2022-008-0067.mrg | 2022-008-0067.sam | 116 | 1998.33 | 1998.29 | 0.50196 | -0.046387 | 34.583 | 34.5798 | 0.00029967 | -0.0031586 | 34.583 | 34.5801 | 0.00022996 | -0.0028839
2022-008-0067.mrg | 2022-008-0067.sam | 121 | 1500.45 | 1500.45 | 0.25012 | 0.00012207 | 34.4919 | 34.4905 | 0.00026744 | -0.0013695 | 34.4919 | 34.4908 | 0.00025156 | -0.0010605
2022-008-0067.mrg | 2022-008-0067.sam | 38 | 1249.51 | 1249.39 | 0.63015 | -0.11377 | 34.4375 | 34.4361 | 0.00021765 | -0.0014267 | 34.4375 | 34.436 | 0.00028428 | -0.001461
2022-008-0067.mrg | 2022-008-0067.sam | 121 | 1000.91 | 1000.91 | 0.35224 | -0.00012207 | 34.3589 | 34.3579 | 0.00026336 | -0.0010071 | 34.3589 | 34.3581 | 0.00026429 | -0.00080872
2022-008-0075.mrg | 2022-008-0075.sam | 121 | 5.979 | 5.9786 | 0.42535 | -0.00040483 | 32.2592 | 32.2582 | 0.00052614 | -0.00098801 | 32.2592 | 32.2575 | 0.00073425 | -0.0017052
2022-008-0077.mrg | 2022-008-0077.sam | 121 | 3500.88 | 3500.88 | 0.53124 | 0.00024414 | 34.6673 | 34.6665 | 0.00021474 | -0.00082016 | 34.6673 | 34.6669 | 0.00022743 | -0.00043106
2022-008-0077.mrg | 2022-008-0077.sam | 121 | 3000.38 | 3000.38 | 0.25583 | 0 | 34.6506 | 34.652 | 0.00025523 | 0.0013657 | 34.6506 | 34.6525 | 0.00024883 | 0.0018616
2022-008-0077.mrg | 2022-008-0077.sam | 121 | 2498.99 | 2498.99 | 0.42057 | -0.00024414 | 34.6252 | 34.6251 | 0.00024547 | -6.485e-05 | 34.6252 | 34.6256 | 0.00023044 | 0.00038528
2022-008-0077.mrg | 2022-008-0077.sam | 102 | 2000.83 | 2001.02 | 0.68757 | 0.18372 | 34.5825 | 34.5803 | 0.00026163 | -0.0022202 | 34.5825 | 34.5805 | 0.00027694 | -0.0019531
2022-008-0077.mrg | 2022-008-0077.sam | 121 | 1749.07 | 1749.07 | 0.31636 | 0.00012207 | 34.5463 | 34.5447 | 0.0002976 | -0.0016365 | 34.5463 | 34.545 | 0.00027885 | -0.0013351
2022-008-0077.mrg | 2022-008-0077.sam | 116 | 1502.91 | 1502.96 | 0.59323 | 0.048462 | 34.4968 | 34.4949 | 0.00029607 | -0.0018883 | 34.4968 | 34.4951 | 0.00025505 | -0.0017013
2022-008-0077.mrg | 2022-008-0077.sam | 121 | 1249.36 | 1249.36 | 0.30777 | 0 | 34.4389 | 34.4372 | 0.00023103 | -0.0017281 | 34.4389 | 34.4372 | 0.00025587 | -0.0016594
2022-008-0077.mrg | 2022-008-0077.sam | 115 | 1000.19 | 1000.24 | 0.44766 | 0.054688 | 34.3682 | 34.3665 | 0.0002581 | -0.001709 | 34.3682 | 34.3666 | 0.00021767 | -0.0015717
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 3500.22 | 3500.22 | 0.077005 | -0.00024414 | 34.6674 | 34.6689 | 0.00022846 | 0.0014915 | 34.6674 | 34.6693 | 0.00029317 | 0.0018806
2022-008-0080.mrg | 2022-008-0080.sam | 118 | 3001.26 | 3001.23 | 0.49702 | -0.026367 | 34.6537 | 34.6527 | 0.00024813 | -0.0009613 | 34.6537 | 34.6531 | 0.00022428 | -0.00060272
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 2500.87 | 2500.87 | 0.13918 | 0 | 34.6256 | 34.6239 | 0.0002918 | -0.0016975 | 34.6256 | 34.6243 | 0.00022354 | -0.0012779
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 2000.84 | 2000.84 | 0.49721 | -0.00012207 | 34.5798 | 34.5812 | 0.00027682 | 0.0013885 | 34.5798 | 34.5815 | 0.00026157 | 0.0016556
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 1750.7 | 1750.7 | 0.24055 | -0.00036621 | 34.5472 | 34.5448 | 0.00029086 | -0.0024338 | 34.5472 | 34.5449 | 0.00024578 | -0.0022736
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 1500.09 | 1500.09 | 0.16545 | -0.00024414 | 34.4989 | 34.4983 | 0.00028111 | -0.00061035 | 34.4989 | 34.4983 | 0.00024923 | -0.00058746
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 1250.5 | 1250.5 | 0.3088 | 0.00024414 | 34.4426 | 34.4397 | 0.00030287 | -0.002903 | 34.4426 | 34.4398 | 0.00032597 | -0.0028229
2022-008-0080.mrg | 2022-008-0080.sam | 121 | 999.365 | 999.365 | 0.11451 | 0.00036621 | 34.3657 | 34.3641 | 0.00036639 | -0.0015526 | 34.3657 | 34.3642 | 0.00031888 | -0.0015297
2022-008-0081.mrg | 2022-008-0081.sam | 121 | 4.965 | 4.9652 | 0.30088 | 0.00022316 | 32.2835 | 32.2814 | 0.0013545 | -0.002121 | 32.2835 | 32.2814 | 0.0015537 | -0.0021324
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 4000.71 | 4000.71 | 0.44038 | 0.00048828 | 34.6773 | 34.6777 | 0.00024064 | 0.00042343 | 34.6773 | 34.6782 | 0.00021232 | 0.00085068
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 3462.47 | 3462.47 | 0.21713 | 0.00048828 | 34.6695 | 34.668 | 0.00024565 | -0.0014992 | 34.6695 | 34.6682 | 0.00022521 | -0.001339
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 3000.08 | 3000.08 | 0.22971 | -0.00024414 | 34.6533 | 34.6524 | 0.00025231 | -0.00086594 | 34.6533 | 34.6527 | 0.00024903 | -0.00059128
2022-008-0092.mrg | 2022-008-0092.sam | 111 | 2498.88 | 2498.78 | 0.56144 | -0.099609 | 34.6281 | 34.6242 | 0.00024188 | -0.0038528 | 34.6281 | 34.6244 | 0.00022564 | -0.0036888
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 1999.85 | 1999.85 | 0.29425 | 0.00024414 | 34.5806 | 34.5771 | 0.00029267 | -0.0035477 | 34.5806 | 34.5771 | 0.00023575 | -0.0035438
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 1499.3 | 1499.3 | 0.5502 | 0.00012207 | 34.4984 | 34.4961 | 0.00028586 | -0.0023308 | 34.4984 | 34.4961 | 0.00024763 | -0.0022964
2022-008-0092.mrg | 2022-008-0092.sam | 121 | 1250.26 | 1250.26 | 0.3783 | -0.00012207 | 34.4404 | 34.4375 | 0.00026741 | -0.002903 | 34.4404 | 34.4375 | 0.00023933 | -0.0028915
2022-008-0092.mrg | 2022-008-0092.sam | 95 | 1000.76 | 1000.96 | 0.69654 | 0.20001 | 34.3622 | 34.3627 | 0.00040632 | 0.00049973 | 34.3622 | 34.3626 | 0.00041871 | 0.00036621
2022-008-0102.mrg | 2022-008-0102.sam | 121 | 1500.87 | 1500.87 | 0.11639 | 0.00024414 | 34.5015 | 34.4959 | 0.00028478 | -0.0056076 | 34.5015 | 34.495 | 0.00024533 | -0.0064621
2022-008-0102.mrg | 2022-008-0102.sam | 121 | 1248.71 | 1248.71 | 0.44679 | 0 | 34.4305 | 34.4301 | 0.00026286 | -0.00035858 | 34.4305 | 34.429 | 0.00026454 | -0.0015068
2022-008-0102.mrg | 2022-008-0102.sam | 96 | 1001.04 | 1000.69 | 0.62203 | -0.3454 | 34.3602 | 34.3572 | 0.00050126 | -0.0029945 | 34.3602 | 34.356 | 0.00059579 | -0.0041771
2022-008-0102.mrg | 2022-008-0102.sam | 121 | 4.903 | 4.9028 | 0.20591 | -0.00023937 | 31.1663 | 31.1627 | 0.00066387 | -0.0035725 | 31.1663 | 31.1615 | 0.00047108 | -0.0048237
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2022-008-0018.mrg | 2022-008-0018.sam | 12 | 16 | 2022-008-0018.mrg & 2022-008-0018.sam
2 | 2022-008-0027.mrg | 2022-008-0027.sam | 18 | 18 | 2022-008-0027.mrg & 2022-008-0027.sam
3 | 2022-008-0029.mrg | 2022-008-0029.sam | 20 | 20 | 2022-008-0029.mrg & 2022-008-0029.sam
4 | 2022-008-0033.mrg | 2022-008-0033.sam | 22 | 26 | 2022-008-0033.mrg & 2022-008-0033.sam
5 | 2022-008-0034.mrg | 2022-008-0034.sam | 28 | 28 | 2022-008-0034.mrg & 2022-008-0034.sam
6 | 2022-008-0038.mrg | 2022-008-0038.sam | 30 | 37 | 2022-008-0038.mrg & 2022-008-0038.sam
7 | 2022-008-0048.mrg | 2022-008-0048.sam | 39 | 39 | 2022-008-0048.mrg & 2022-008-0048.sam
8 | 2022-008-0049.mrg | 2022-008-0049.sam | 41 | 64 | 2022-008-0049.mrg & 2022-008-0049.sam
9 | 2022-008-0051.mrg | 2022-008-0051.sam | 66 | 66 | 2022-008-0051.mrg & 2022-008-0051.sam
10 | 2022-008-0057.mrg | 2022-008-0057.sam | 68 | 76 | 2022-008-0057.mrg & 2022-008-0057.sam
11 | 2022-008-0061.mrg | 2022-008-0061.sam | 78 | 78 | 2022-008-0061.mrg & 2022-008-0061.sam
12 | 2022-008-0062.mrg | 2022-008-0062.sam | 80 | 87 | 2022-008-0062.mrg & 2022-008-0062.sam
13 | 2022-008-0064.mrg | 2022-008-0064.sam | 89 | 89 | 2022-008-0064.mrg & 2022-008-0064.sam
14 | 2022-008-0067.mrg | 2022-008-0067.sam | 91 | 98 | 2022-008-0067.mrg & 2022-008-0067.sam
15 | 2022-008-0075.mrg | 2022-008-0075.sam | 100 | 100 | 2022-008-0075.mrg & 2022-008-0075.sam
16 | 2022-008-0077.mrg | 2022-008-0077.sam | 102 | 109 | 2022-008-0077.mrg & 2022-008-0077.sam
17 | 2022-008-0080.mrg | 2022-008-0080.sam | 111 | 118 | 2022-008-0080.mrg & 2022-008-0080.sam
18 | 2022-008-0081.mrg | 2022-008-0081.sam | 120 | 120 | 2022-008-0081.mrg & 2022-008-0081.sam
19 | 2022-008-0092.mrg | 2022-008-0092.sam | 122 | 129 | 2022-008-0092.mrg & 2022-008-0092.sam
20 | 2022-008-0102.mrg | 2022-008-0102.sam | 131 | 134 | 2022-008-0102.mrg & 2022-008-0102.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 20
2 | (Sal1 - Sal_Bot) | Pressure | 20
3 | (Sal0 - Sal_Bot) | File Pair # | 20
4 | (Sal0 - Sal_Bot) | Pressure | 1
5 | (Sal1 - Sal_Bot) | File Pair # | 20
6 | (Sal0 - Sal_Bot) | Pressure | 20

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.031274 |  | X | CALC Y
m= | 2.1544e-07 | 1.2639e-07 |  | Std Dev of Resid |  | 0.0012849 |  | 4.8037 | -0.0017215
b= | -0.0017225 | 0.00025138 |  | Resid Sum of Squares |  | 0.0001486 |  | 4000.71 | -0.00086061
# of Values Extracted = |  |  | 104
# of Values Used in Fit = |  |  | 92
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid | Bottle | X | Y | [7 empty] | Sal0 | Sal1 | Sal1-Sal0
1 | YES | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.0014534 | -3.773e-06 | 34.4606 | 1249.41 | -0.0014496 | [7 empty] | 34.4592 | 34.4595 | 0.0003
1 | YES | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.0014534 | -3.773e-06 | 34.4606 | 1249.41 | -0.0014496 | [7 empty] | 34.4592 | 34.4595 | 0.0003
1 | YES | 1000.39 | 34.3692 | 0.00026172 | -0.0041122 | -0.001507 | 0.0026052 | 34.3733 | 1000.39 | -0.0041122 | [7 empty] | 34.3692 | 34.3692 | 0
1 | YES | 800.385 | 34.27 | 0.00025105 | -0.0032005 | -0.0015501 | 0.0016504 | 34.2732 | 800.385 | -0.0032005 | [7 empty] | 34.27 | 34.27 | 0
1 | NO | 4.9435 | 31.7774 | 0.00029579 | -0.0095119 | -0.0017215 | 0.0077905 | 31.7869 | 2000.61 | -0.0017281 | [7 empty] | 31.7774 | 31.7775 | 0.0001
2 | NO | 5.4165 | 31.8866 | 0.0079538 | -0.0062408 | -0.0017214 | 0.0045195 | 31.8928 | 1500.35 | -0.0017929 | [7 empty] | 31.8866 | 31.8869 | 0.0003
3 | NO | 5.0867 | 31.899 | 0.0017521 | -0.01651 | -0.0017214 | 0.014789 | 31.9155 | 1250.43 | -0.0030632 | [7 empty] | 31.899 | 31.8994 | 0.0004
4 | YES | 2000.61 | 34.5842 | 0.00027139 | -0.0017281 | -0.0012915 | 0.00043655 | 34.5859 | 800.791 | -0.0025368 | [7 empty] | 34.5842 | 34.5848 | 0.0006
4 | YES | 1500.35 | 34.5107 | 0.00034639 | -0.0017929 | -0.0013993 | 0.00039362 | 34.5125 | 5.1646 | -0.00034714 | [7 empty] | 34.5107 | 34.511 | 0.0003
4 | YES | 1250.43 | 34.4474 | 0.00025759 | -0.0030632 | -0.0014531 | 0.0016101 | 34.4505 | 2998.4 | -0.00078964 | [7 empty] | 34.4474 | 34.4477 | 0.0003
4 | NO | 1000.36 | 34.3983 | 0.00026999 | -0.0054588 | -0.001507 | 0.0039518 | 34.4038 | 2999.49 | -0.00067902 | 46 |  | 34.4005 | 34.4071 | -0.0022 | -0.0088 |  | 34.3983 | 34.3984 | 0.0001
4 | YES | 800.791 | 34.3025 | 0.00037756 | -0.0025368 | -0.00155 | 0.00098677 | 34.305 | 2500.32 | -0.00034714 | [7 empty] | 34.3025 | 34.3022 | -0.0003
5 | YES | 5.1646 | 31.8274 | 0.00043663 | -0.00034714 | -0.0017214 | -0.0013743 | 31.8277 | 1999.72 | -0.0007019 | [7 empty] | 31.8274 | 31.8274 | 0
6 | YES | 2998.4 | 34.6454 | 0.00027236 | -0.00078964 | -0.0010765 | -0.0002869 | 34.6462 | 1500.84 | -0.0020676 | [7 empty] | 34.6454 | 34.6461 | 0.0007
6 | YES | 2999.49 | 34.6454 | 0.00025291 | -0.00067902 | -0.0010763 | -0.0003973 | 34.6461 | 1251.02 | -0.0015984 | [7 empty] | 34.6454 | 34.646 | 0.0006
6 | YES | 2500.32 | 34.6235 | 0.00028436 | -0.00034714 | -0.0011839 | -0.00083672 | 34.6238 | 999.451 | -0.00087738 | [7 empty] | 34.6235 | 34.6241 | 0.0006
6 | YES | 1999.72 | 34.588 | 0.00025896 | -0.0007019 | -0.0012917 | -0.0005898 | 34.5887 | 5.2009 | 0.00055695 | [7 empty] | 34.588 | 34.5886 | 0.0006
6 | YES | 1500.84 | 34.5091 | 0.000263 | -0.0020676 | -0.0013992 | 0.00066839 | 34.5112 | 4.8037 | -0.00019074 | [7 empty] | 34.5091 | 34.5095 | 0.0004
6 | YES | 1251.02 | 34.4529 | 0.00025925 | -0.0015984 | -0.001453 | 0.00014535 | 34.4545 | 3359.02 | -0.0010223 | [7 empty] | 34.4529 | 34.4531 | 0.0002
6 | YES | 999.451 | 34.3676 | 0.00023208 | -0.00087738 | -0.0015072 | -0.00062982 | 34.3685 | 2999.35 | 0.00099564 | [7 empty] | 34.3676 | 34.3677 | 0.0001
6 | YES | 5.2009 | 32.1462 | 0.00023974 | 0.00055695 | -0.0017214 | -0.0022784 | 32.1456 | 2500.22 | -0.0010376 |  |  | Surface comparison - well-mixed casts |  |  |  |  | 32.1462 | 32.1464 | 0.0002
7 | YES | 4.8037 | 32.2203 | 0.00023888 | -0.00019074 | -0.0017215 | -0.0015308 | 32.2205 | 2249.12 | 0.00048828 |  |  | File Pair | CTD | Bottle | CTD-Bot |  | 32.2203 | 32.2199 | -0.0004
8 | YES | 3359.02 | 34.6553 | 0.00027164 | -0.0010223 | -0.00099885 | 2.3486e-05 | 34.6563 | 2000.63 | -0.0016022 |  |  | 6 | 32.1462 | 32.1456 | 0.00055695 |  | 34.6553 | 34.656 | 0.0007
8 | YES | 2999.35 | 34.65 | 0.0002739 | 0.00099564 | -0.0010763 | -0.002072 | 34.649 | 1749.83 | -0.0026932 |  |  | 7 | 32.2203 | 32.2205 | -0.00019074 |  | 34.65 | 34.6508 | 0.0008
8 | YES | 2500.22 | 34.6292 | 0.00029204 | -0.0010376 | -0.0011839 | -0.00014628 | 34.6302 | 1499.36 | -0.0014954 |  | 32.3045 | 11 | 32.3047 | 32.3047 | 0 |  | 34.6292 | 34.6298 | 0.0006
8 | YES | 2249.12 | 34.6088 | 0.00025415 | 0.00048828 | -0.001238 | -0.0017263 | 34.6083 | 1249.27 | -0.0012093 |  |  | 13 | 32.288 | 32.2891 | -0.0010872 |  | 34.6088 | 34.6094 | 0.0006
8 | YES | 2000.63 | 34.584 | 0.00026102 | -0.0016022 | -0.0012915 | 0.00031066 | 34.5856 | 1000.65 | -0.0015793 |  |  | 18 | 32.2814 | 32.2835 | -0.002121 | x | 34.584 | 34.5846 | 0.0006
8 | YES | 1749.83 | 34.5524 | 0.00026912 | -0.0026932 | -0.0013455 | 0.0013476 | 34.5551 | 799.545 | -0.0007515 | [7 empty] | 34.5524 | 34.5529 | 0.0005
8 | YES | 1499.36 | 34.5063 | 0.00038571 | -0.0014954 | -0.0013995 | 9.5857e-05 | 34.5078 | 600.327 | -0.002842 | [7 empty] | 34.5063 | 34.5065 | 0.0002
8 | YES | 1249.27 | 34.4506 | 0.00033206 | -0.0012093 | -0.0014534 | -0.00024412 | 34.4518 | 400.437 | 2.6703e-05 | [7 empty] | 34.4506 | 34.4507 | 0.0001
8 | YES | 1000.65 | 34.3688 | 0.00020316 | -0.0015793 | -0.0015069 | 7.2334e-05 | 34.3704 | 299.392 | -0.0002861 | [7 empty] | 34.3688 | 34.369 | 0.0002
8 | YES | 799.545 | 34.2813 | 0.00026137 | -0.0007515 | -0.0015503 | -0.00079878 | 34.2821 | 249.804 | -0.00071335 | [7 empty] | 34.2813 | 34.2814 | 0.0001
8 | YES | 600.327 | 34.1431 | 0.00028963 | -0.002842 | -0.0015932 | 0.0012488 | 34.1459 | 200.768 | -0.0010529 | [7 empty] | 34.1431 | 34.1428 | -0.0003
8 | YES | 400.437 | 33.9404 | 0.00025309 | 2.6703e-05 | -0.0016363 | -0.001663 | 33.9404 | 174.785 | -0.00081253 | [7 empty] | 33.9404 | 33.9404 | 0
8 | YES | 299.392 | 33.8284 | 0.00047768 | -0.0002861 | -0.001658 | -0.0013719 | 33.8287 | 75.4152 | -0.0023155 | [7 empty] | 33.8284 | 33.8284 | 0
8 | YES | 249.804 | 33.8061 | 0.00041418 | -0.00071335 | -0.0016687 | -0.00095536 | 33.8068 | 10.6675 | -0.00030899 | [7 empty] | 33.8061 | 33.806 | -0.0001
8 | YES | 200.768 | 33.7699 | 0.00047069 | -0.0010529 | -0.0016793 | -0.00062641 | 33.771 | 3500.17 | -0.0011749 | [7 empty] | 33.7699 | 33.7698 | -0.0001
8 | YES | 174.785 | 33.7305 | 0.00065778 | -0.00081253 | -0.0016849 | -0.00087234 | 33.7313 | 3500.17 | -0.0011749 | [7 empty] | 33.7305 | 33.7307 | 0.0002
8 | NO | 150.431 | 33.5993 | 0.00094877 | -0.0081558 | -0.0016901 | 0.0064657 | 33.6075 | 3001.09 | 0.00022888 | [7 empty] | 33.5993 | 33.5991 | -0.0002
8 | NO | 125.193 | 33.1446 | 0.0015435 | -0.019547 | -0.0016956 | 0.017851 | 33.1641 | 2500.37 | -0.0015602 | 4 | [6 empty] | 33.1446 | 33.1429 | -0.0017
8 | NO | 99.5611 | 32.6081 | 0.0013818 | -0.022873 | -0.0017011 | 0.021172 | 32.631 | 1999.65 | -0.0030785 | 4 | [6 empty] | 32.6081 | 32.608 | -0.0001
8 | YES | 75.4152 | 32.3448 | 0.00044928 | -0.0023155 | -0.0017063 | 0.00060924 | 32.3471 | 1499.41 | -0.0015526 | [7 empty] | 32.3448 | 32.3446 | -0.0002
8 | NO | 49.9996 | 32.301 | 0.0032722 | 0.00089264 | -0.0017118 | -0.0026044 | 32.3001 | 1250.27 | -0.0038948 | [7 empty] | 32.301 | 32.3012 | 0.0002
8 | NO | 24.4641 | 32.2205 | 0.00091053 | -8.7738e-05 | -0.0017173 | -0.0016295 | 32.2206 | 1000.01 | -0.0016365 | [7 empty] | 32.2205 | 32.2203 | -0.0002
8 | YES | 10.6675 | 32.2158 | 0.00018787 | -0.00030899 | -0.0017202 | -0.0014112 | 32.2161 | 800.236 | -0.0011215 | [7 empty] | 32.2158 | 32.2158 | 0
8 | NO | 4.8307 | 32.2161 | 0.0015046 | -0.0013008 | -0.0017215 | -0.00042068 | 32.2174 | 5.1104 | 0 | [7 empty] | 32.2161 | 32.2165 | 0.0004
9 | NO | 5.2141 | 32.1473 | 0.0015834 | -0.0006752 | -0.0017214 | -0.0010462 | 32.148 | 3500.36 | 0.00039673 | [7 empty] | 32.1473 | 32.1472 | -0.0001
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.00096844 | 0.00020649 | 34.6651 | 3001.04 | -0.00079346 | [7 empty] | 34.6639 | 34.6645 | 0.0006
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.00096844 | 0.00020649 | 34.6651 | 2499.79 | -0.0037079 | [7 empty] | 34.6639 | 34.6645 | 0.0006
10 | YES | 3001.09 | 34.6524 | 0.00027319 | 0.00022888 | -0.001076 | -0.0013048 | 34.6522 | 2000.06 | -0.0013695 | [7 empty] | 34.6524 | 34.6531 | 0.0007
10 | YES | 2500.37 | 34.623 | 0.00031571 | -0.0015602 | -0.0011838 | 0.00037637 | 34.6246 | 1750.73 | -0.0018997 | [7 empty] | 34.623 | 34.6236 | 0.0006
10 | YES | 1999.65 | 34.5747 | 0.00022288 | -0.0030785 | -0.0012917 | 0.0017867 | 34.5778 | 1500.78 | -0.001049 | [7 empty] | 34.5747 | 34.5752 | 0.0005
10 | YES | 1499.41 | 34.4919 | 0.00022583 | -0.0015526 | -0.0013995 | 0.00015309 | 34.4935 | 1250.34 | -0.00078583 | [7 empty] | 34.4919 | 34.4923 | 0.0004
10 | YES | 1250.27 | 34.4323 | 0.00023651 | -0.0038948 | -0.0014532 | 0.0024416 | 34.4362 | 1000.73 | -0.0006218 | [7 empty] | 34.4323 | 34.4325 | 0.0002
10 | YES | 1000.01 | 34.3488 | 0.00028203 | -0.0016365 | -0.0015071 | 0.00012943 | 34.3504 | 5.2854 | -0.0010872 | [7 empty] | 34.3488 | 34.3489 | 0.0001
10 | YES | 800.236 | 34.2653 | 0.00025735 | -0.0011215 | -0.0015501 | -0.0004286 | 34.2664 | 3498.78 | -0.0021439 | [7 empty] | 34.2653 | 34.2654 | 0.0001
11 | YES | 5.1104 | 32.3047 | 0.00022224 | 0 | -0.0017214 | -0.0017214 | 32.3047 | 3498.78 | -0.00084305 | [7 empty] | 32.3047 | 32.3045 | -0.0002
12 | YES | 3500.36 | 34.6663 | 0.00026718 | 0.00039673 | -0.0009684 | -0.0013651 | 34.6659 | 2999.64 | -6.485e-05 | [7 empty] | 34.6663 | 34.6668 | 0.0005
12 | YES | 3001.04 | 34.6531 | 0.00024095 | -0.00079346 | -0.001076 | -0.00028252 | 34.6539 | 2500.15 | -0.0010567 | [7 empty] | 34.6531 | 34.6537 | 0.0006
12 | YES | 2499.79 | 34.6265 | 0.00026978 | -0.0037079 | -0.001184 | 0.0025239 | 34.6302 | 1998.29 | -0.0031586 | [7 empty] | 34.6265 | 34.6269 | 0.0004
12 | YES | 2000.06 | 34.5801 | 0.00028982 | -0.0013695 | -0.0012916 | 7.7849e-05 | 34.5815 | 1500.45 | -0.0013695 | [7 empty] | 34.5801 | 34.5806 | 0.0005
12 | YES | 1750.73 | 34.5387 | 0.00024204 | -0.0018997 | -0.0013453 | 0.00055437 | 34.5406 | 1249.39 | -0.0014267 | [7 empty] | 34.5387 | 34.539 | 0.0003
12 | YES | 1500.78 | 34.493 | 0.0002886 | -0.001049 | -0.0013992 | -0.00035016 | 34.494 | 1000.91 | -0.0010071 | [7 empty] | 34.493 | 34.4934 | 0.0004
12 | YES | 1250.34 | 34.4358 | 0.00032793 | -0.00078583 | -0.0014532 | -0.00066732 | 34.4366 | 5.9786 | -0.00098801 | [7 empty] | 34.4358 | 34.436 | 0.0002
12 | YES | 1000.73 | 34.3599 | 0.00023286 | -0.0006218 | -0.0015069 | -0.00088513 | 34.3605 | 3500.88 | -0.00082016 | [7 empty] | 34.3599 | 34.36 | 0.0001
13 | YES | 5.2854 | 32.288 | 0.00028753 | -0.0010872 | -0.0017214 | -0.0006342 | 32.2891 | 3000.38 | 0.0013657 | [7 empty] | 32.288 | 32.2877 | -0.0003
14 | YES | 3498.78 | 34.6668 | 0.00026182 | -0.0021439 | -0.00096874 | 0.0011751 | 34.6689 | 2498.99 | -6.485e-05 | [7 empty] | 34.6668 | 34.6672 | 0.0004
14 | YES | 3498.78 | 34.6668 | 0.00026182 | -0.00084305 | -0.00096874 | -0.0001257 | 34.6676 | 2001.02 | -0.0022202 | [7 empty] | 34.6668 | 34.6672 | 0.0004
14 | YES | 2999.64 | 34.6522 | 0.00023523 | -6.485e-05 | -0.0010763 | -0.0010114 | 34.6523 | 1749.07 | -0.0016365 | [7 empty] | 34.6522 | 34.6527 | 0.0005
14 | YES | 2500.15 | 34.6231 | 0.00027794 | -0.0010567 | -0.0011839 | -0.00012722 | 34.6242 | 1502.96 | -0.0018883 | [7 empty] | 34.6231 | 34.6235 | 0.0004
14 | YES | 1998.29 | 34.5798 | 0.00029967 | -0.0031586 | -0.001292 | 0.0018666 | 34.583 | 1249.36 | -0.0017281 | [7 empty] | 34.5798 | 34.5801 | 0.0003
14 | YES | 1500.45 | 34.4905 | 0.00026744 | -0.0013695 | -0.0013993 | -2.9788e-05 | 34.4919 | 1000.24 | -0.001709 | [7 empty] | 34.4905 | 34.4908 | 0.0003
14 | YES | 1249.39 | 34.4361 | 0.00021765 | -0.0014267 | -0.0014534 | -2.6657e-05 | 34.4375 | 3500.22 | 0.0014915 | [7 empty] | 34.4361 | 34.436 | -0.0001
14 | YES | 1000.91 | 34.3579 | 0.00026336 | -0.0010071 | -0.0015069 | -0.00049981 | 34.3589 | 3001.23 | -0.0009613 | [7 empty] | 34.3579 | 34.3581 | 0.0002
15 | YES | 5.9786 | 32.2582 | 0.00052614 | -0.00098801 | -0.0017212 | -0.00073323 | 32.2592 | 2500.87 | -0.0016975 | [7 empty] | 32.2582 | 32.2575 | -0.0007
16 | YES | 3500.88 | 34.6665 | 0.00021474 | -0.00082016 | -0.00096829 | -0.00014813 | 34.6673 | 2000.84 | 0.0013885 | [7 empty] | 34.6665 | 34.6669 | 0.0004
16 | YES | 3000.38 | 34.652 | 0.00025523 | 0.0013657 | -0.0010761 | -0.0024418 | 34.6506 | 1750.7 | -0.0024338 | [7 empty] | 34.652 | 34.6525 | 0.0005
16 | YES | 2498.99 | 34.6251 | 0.00024547 | -6.485e-05 | -0.0011841 | -0.0011193 | 34.6252 | 1500.09 | -0.00061035 | [7 empty] | 34.6251 | 34.6256 | 0.0005
16 | YES | 2001.02 | 34.5803 | 0.00026163 | -0.0022202 | -0.0012914 | 0.00092873 | 34.5825 | 1250.5 | -0.002903 | [7 empty] | 34.5803 | 34.5805 | 0.0002
16 | YES | 1749.07 | 34.5447 | 0.0002976 | -0.0016365 | -0.0013457 | 0.0002908 | 34.5463 | 999.365 | -0.0015526 | [7 empty] | 34.5447 | 34.545 | 0.0003
16 | YES | 1502.96 | 34.4949 | 0.00029607 | -0.0018883 | -0.0013987 | 0.00048955 | 34.4968 | 4000.71 | 0.00042343 | [7 empty] | 34.4949 | 34.4951 | 0.0002
16 | YES | 1249.36 | 34.4372 | 0.00023103 | -0.0017281 | -0.0014534 | 0.0002747 | 34.4389 | 3462.47 | -0.0014992 | [7 empty] | 34.4372 | 34.4372 | 0
16 | YES | 1000.24 | 34.3665 | 0.0002581 | -0.001709 | -0.001507 | 0.00020195 | 34.3682 | 3000.08 | -0.00086594 | [7 empty] | 34.3665 | 34.3666 | 0.0001
17 | YES | 3500.22 | 34.6689 | 0.00022846 | 0.0014915 | -0.00096843 | -0.00246 | 34.6674 | 2498.78 | -0.0038528 | [7 empty] | 34.6689 | 34.6693 | 0.0004
17 | YES | 3001.23 | 34.6527 | 0.00024813 | -0.0009613 | -0.0010759 | -0.00011463 | 34.6537 | 1999.85 | -0.0035477 | [7 empty] | 34.6527 | 34.6531 | 0.0004
17 | YES | 2500.87 | 34.6239 | 0.0002918 | -0.0016975 | -0.0011837 | 0.0005138 | 34.6256 | 1499.3 | -0.0023308 | [7 empty] | 34.6239 | 34.6243 | 0.0004
17 | YES | 2000.84 | 34.5812 | 0.00027682 | 0.0013885 | -0.0012915 | -0.00268 | 34.5798 | 1250.26 | -0.002903 | [7 empty] | 34.5812 | 34.5815 | 0.0003
17 | YES | 1750.7 | 34.5448 | 0.00029086 | -0.0024338 | -0.0013454 | 0.0010884 | 34.5472 | 1000.96 | 0.00049973 | [7 empty] | 34.5448 | 34.5449 | 0.0001
17 | YES | 1500.09 | 34.4983 | 0.00028111 | -0.00061035 | -0.0013993 | -0.00078899 | 34.4989 | 1500.87 | -0.0056076 | [7 empty] | 34.4983 | 34.4983 | 0
17 | YES | 1250.5 | 34.4397 | 0.00030287 | -0.002903 | -0.0014531 | 0.0014499 | 34.4426 | 1248.71 | -0.00035858 | [7 empty] | 34.4397 | 34.4398 | 0.0001
17 | YES | 999.365 | 34.3641 | 0.00036639 | -0.0015526 | -0.0015072 | 4.5357e-05 | 34.3657 | 1000.69 | -0.0029945 | [7 empty] | 34.3641 | 34.3642 | 0.0001
18 | NO | 4.9652 | 32.2814 | 0.0013545 | -0.002121 | -0.0017215 | 0.00039951 | 32.2835 | 4.9028 | -0.0035725 | [7 empty] | 32.2814 | 32.2814 | 0
19 | YES | 4000.71 | 34.6777 | 0.00024064 | 0.00042343 | -0.00086061 | -0.001284 | 34.6773 |  |  | -0.00136 | 0.0012983 | [5 empty] | 34.6777 | 34.6782 | 0.0005
19 | YES | 3462.47 | 34.668 | 0.00024565 | -0.0014992 | -0.00097657 | 0.00052261 | 34.6695 | [9 empty] | 34.668 | 34.6682 | 0.0002
19 | YES | 3000.08 | 34.6524 | 0.00025231 | -0.00086594 | -0.0010762 | -0.00021025 | 34.6533 | [9 empty] | 34.6524 | 34.6527 | 0.0003
19 | YES | 2498.78 | 34.6242 | 0.00024188 | -0.0038528 | -0.0011842 | 0.0026687 | 34.6281 | [9 empty] | 34.6242 | 34.6244 | 0.0002
19 | YES | 1999.85 | 34.5771 | 0.00029267 | -0.0035477 | -0.0012917 | 0.002256 | 34.5806 | [9 empty] | 34.5771 | 34.5771 | 0
19 | YES | 1499.3 | 34.4961 | 0.00028586 | -0.0023308 | -0.0013995 | 0.00093126 | 34.4984 | [9 empty] | 34.4961 | 34.4961 | 0
19 | YES | 1250.26 | 34.4375 | 0.00026741 | -0.002903 | -0.0014532 | 0.0014498 | 34.4404 | [9 empty] | 34.4375 | 34.4375 | 0
19 | YES | 1000.96 | 34.3627 | 0.00040632 | 0.00049973 | -0.0015069 | -0.0020066 | 34.3622 | [9 empty] | 34.3627 | 34.3626 | -0.0001
20 | YES | 1500.87 | 34.4959 | 0.00028478 | -0.0056076 | -0.0013992 | 0.0042084 | 34.5015 | [9 empty] | 34.4959 | 34.495 | -0.0009
20 | YES | 1248.71 | 34.4301 | 0.00026286 | -0.00035858 | -0.0014535 | -0.0010949 | 34.4305 | [9 empty] | 34.4301 | 34.429 | -0.0011
20 | YES | 1000.69 | 34.3572 | 0.00050126 | -0.0029945 | -0.0015069 | 0.0014876 | 34.3602 | [9 empty] | 34.3572 | 34.356 | -0.0012
20 | YES | 4.9028 | 31.1627 | 0.00066387 | -0.0035725 | -0.0017215 | 0.001851 | 31.1663 | [9 empty] | 31.1627 | 31.1615 | -0.0012
|  | average std dev |  | 0.00047722
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.10627 |  | X | CALC Y
m= | 4.4895e-07 | 1.3724e-07 |  | Std Dev of Resid |  | 0.0013952 |  | 4.8037 | -0.0019327
b= | -0.0019348 | 0.00027295 |  | Resid Sum of Squares |  | 0.00017519 |  | 4000.71 | -0.00013872
# of Values Extracted = |  |  | 104
# of Values Used in Fit = |  |  | 92
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1249.41 | 34.4595 | 0.00023282 | -0.0011063 | -0.0013739 | -0.00026766 |  | 1249.41 | -0.0011063 |  |  |  | Surface comparison - 5 surface well-mixed casts - one rejected due to high std dev in CTD Sal
1 | YES | 1249.41 | 34.4595 | 0.00023282 | -0.0011063 | -0.0013739 | -0.00026766 |  | 1249.41 | -0.0011063 |  | Primary |  | File Pair | CTD | Bottle | CTD-Bot
1 | YES | 1000.39 | 34.3692 | 0.0002916 | -0.0040703 | -0.0014857 | 0.0025846 |  | 1000.39 | -0.0040703 |  |  |  | 6 | 32.1462 | 32.1456 | 0.00055695
1 | YES | 800.385 | 34.27 | 0.00027095 | -0.0032158 | -0.0015755 | 0.0016403 |  | 800.385 | -0.0032158 |  |  |  | 7 | 32.2203 | 32.2205 | -0.00019074
1 | NO | 4.9435 | 31.7775 | 0.00030083 | -0.0094013 | -0.0019326 | 0.0074687 |  | 2000.61 | -0.0011444 |  |  | 32.3045 | 11 | 32.3047 | 32.3047 | 0
2 | NO | 5.4165 | 31.8869 | 0.0084602 | -0.0059318 | -0.0019324 | 0.0039994 |  | 1500.35 | -0.0015106 |  |  |  | 13 | 32.288 | 32.2891 | -0.0010872
3 | NO | 5.0867 | 31.8994 | 0.0023966 | -0.01609 | -0.0019326 | 0.014158 |  | 1250.43 | -0.0028496 |  |  |  | 18 | 32.2814 | 32.2835 | -0.002121 | high std dev in CTD Sal
4 | YES | 2000.61 | 34.5848 | 0.00022745 | -0.0011444 | -0.0010367 | 0.00010774 |  | 800.791 | -0.0027504 |  |  |  |  | average |  | -0.00018024
4 | YES | 1500.35 | 34.511 | 0.00025369 | -0.0015106 | -0.0012613 | 0.00024936 |  | 5.1646 | -0.00027084
4 | YES | 1250.43 | 34.4477 | 0.00027648 | -0.0028496 | -0.0013735 | 0.0014761 |  | 2998.4 | -0.00014496 |  |  |  | Surface comparison - well-mixed casts - one rejected due to high std dev in CTD Sal
4 | NO | 1000.36 | 34.3984 | 0.00028179 | -0.0053825 | -0.0014857 | 0.0038968 |  | 2999.49 | -0.00010681 |  | Secondary |  | File Pair | CTD | Bottle | CTD-Bot
4 | YES | 800.791 | 34.3022 | 0.00035346 | -0.0027504 | -0.0015753 | 0.0011751 |  | 2500.32 | 0.00028229 |  |  |  | 6 | 32.1464 | 32.1456 | 0.00075695
5 | YES | 5.1646 | 31.8274 | 0.00049417 | -0.00027084 | -0.0019325 | -0.0016617 |  | 1999.72 | -0.00014877 |  |  |  | 7 | 32.2199 | 32.2205 | -0.00059073
6 | YES | 2998.4 | 34.6461 | 0.00027567 | -0.00014496 | -0.00058871 | -0.00044375 |  | 1500.84 | -0.0017471 |  |  |  | 11 | 32.3045 | 32.3047 | -0.0002
6 | YES | 2999.49 | 34.646 | 0.0002608 | -0.00010681 | -0.00058822 | -0.0004814 |  | 1251.02 | -0.0014305 |  |  |  | 13 | 32.2877 | 32.2891 | -0.0013872
6 | YES | 2500.32 | 34.6241 | 0.00024004 | 0.00028229 | -0.00081232 | -0.0010946 |  | 999.451 | -0.0007782 |  |  |  | 18 | 32.2814 | 32.2835 | -0.002121 | high std dev in CTD Sal
6 | YES | 1999.72 | 34.5886 | 0.00024436 | -0.00014877 | -0.0010371 | -0.00088829 |  | 5.2009 | 0.00080109 |  |  |  |  | average |  | -0.00035524
6 | YES | 1500.84 | 34.5095 | 0.000232 | -0.0017471 | -0.001261 | 0.00048609 |  | 4.8037 | -0.00056076
6 | YES | 1251.02 | 34.4531 | 0.00025644 | -0.0014305 | -0.0013732 | 5.7314e-05 |  | 3359.02 | -0.00032043
6 | YES | 999.451 | 34.3677 | 0.00025206 | -0.0007782 | -0.0014861 | -0.00070794 |  | 2999.35 | 0.0017662
6 | YES | 5.2009 | 32.1464 | 0.00020709 | 0.00080109 | -0.0019325 | -0.0027336 |  | 2500.22 | -0.00038528
7 | YES | 4.8037 | 32.2199 | 0.00034732 | -0.00056076 | -0.0019327 | -0.0013719 |  | 2249.12 | 0.0011177
8 | YES | 3359.02 | 34.656 | 0.00023636 | -0.00032043 | -0.0004268 | -0.00010637 |  | 2000.63 | -0.00099564
8 | YES | 2999.35 | 34.6508 | 0.00025676 | 0.0017662 | -0.00058828 | -0.0023545 |  | 1749.83 | -0.0022469
8 | YES | 2500.22 | 34.6298 | 0.00026352 | -0.00038528 | -0.00081236 | -0.00042708 |  | 1499.36 | -0.0012665
8 | YES | 2249.12 | 34.6094 | 0.00025247 | 0.0011177 | -0.0009251 | -0.0020428 |  | 1249.27 | -0.0010643
8 | YES | 2000.63 | 34.5846 | 0.00022898 | -0.00099564 | -0.0010367 | -4.102e-05 |  | 1000.65 | -0.0013771
8 | YES | 1749.83 | 34.5529 | 0.00024041 | -0.0022469 | -0.0011493 | 0.0010976 |  | 799.545 | -0.00071716
8 | YES | 1499.36 | 34.5065 | 0.00028308 | -0.0012665 | -0.0012617 | 4.7772e-06 |  | 600.327 | -0.0031052
8 | YES | 1249.27 | 34.4507 | 0.0002659 | -0.0010643 | -0.001374 | -0.00030968 |  | 400.437 | -1.5259e-05
8 | YES | 1000.65 | 34.369 | 0.00022357 | -0.0013771 | -0.0014856 | -0.00010849 |  | 299.392 | -0.00029755
8 | YES | 799.545 | 34.2814 | 0.0002601 | -0.00071716 | -0.0015759 | -0.00085872 |  | 249.804 | -0.00081635
8 | YES | 600.327 | 34.1428 | 0.0002843 | -0.0031052 | -0.0016653 | 0.0014398 |  | 200.768 | -0.0012169
8 | YES | 400.437 | 33.9404 | 0.0002818 | -1.5259e-05 | -0.0017551 | -0.0017398 |  | 174.785 | -0.00057983
8 | YES | 299.392 | 33.8284 | 0.00056987 | -0.00029755 | -0.0018004 | -0.0015029 |  | 75.4152 | -0.0025291
8 | YES | 249.804 | 33.806 | 0.00026868 | -0.00081635 | -0.0018227 | -0.0010063 |  | 10.6675 | -0.00027466
8 | YES | 200.768 | 33.7698 | 0.00037694 | -0.0012169 | -0.0018447 | -0.00062782 |  | 3500.17 | -0.00060654
8 | YES | 174.785 | 33.7307 | 0.0007299 | -0.00057983 | -0.0018564 | -0.0012765 |  | 3500.17 | -0.00060654
8 | NO | 150.431 | 33.5991 | 0.0012477 | -0.0084229 | -0.0018673 | 0.0065555 |  | 3001.09 | 0.00094223
8 | NO | 125.193 | 33.1429 | 0.0015026 | -0.021229 | -0.0018786 | 0.01935 |  | 2500.37 | -0.0010338
8 | NO | 99.5611 | 32.608 | 0.0014828 | -0.023014 | -0.0018901 | 0.021124 |  | 1999.65 | -0.0025864
8 | YES | 75.4152 | 32.3446 | 0.00037288 | -0.0025291 | -0.001901 | 0.00062815 |  | 1499.41 | -0.0012093
8 | NO | 49.9996 | 32.3012 | 0.0033732 | 0.0011482 | -0.0019124 | -0.0030606 |  | 1250.27 | -0.0036926
8 | NO | 24.4641 | 32.2203 | 0.0010984 | -0.00033951 | -0.0019239 | -0.0015844 |  | 1000.01 | -0.0014572
8 | YES | 10.6675 | 32.2158 | 0.00019545 | -0.00027466 | -0.0019301 | -0.0016554 |  | 800.236 | -0.00098801
8 | NO | 4.8307 | 32.2165 | 0.0028287 | -0.00087357 | -0.0019327 | -0.0010591 |  | 5.1104 | -0.00022888
9 | NO | 5.2141 | 32.1472 | 0.0015668 | -0.00079727 | -0.0019325 | -0.0011352 |  | 3500.36 | 0.00086975
10 | YES | 3500.17 | 34.6645 | 0.00025192 | -0.00060654 | -0.00036343 | 0.0002431 |  | 3001.04 | -0.00019074
10 | YES | 3500.17 | 34.6645 | 0.00025192 | -0.00060654 | -0.00036343 | 0.0002431 |  | 2499.79 | -0.0032844
10 | YES | 3001.09 | 34.6531 | 0.00023742 | 0.00094223 | -0.0005875 | -0.0015297 |  | 2000.06 | -0.00092316
10 | YES | 2500.37 | 34.6236 | 0.00024065 | -0.0010338 | -0.0008123 | 0.00022148 |  | 1750.73 | -0.0016327
10 | YES | 1999.65 | 34.5752 | 0.0002517 | -0.0025864 | -0.0010371 | 0.0015493 |  | 1500.78 | -0.0006752
10 | YES | 1499.41 | 34.4923 | 0.00028475 | -0.0012093 | -0.0012617 | -5.242e-05 |  | 1250.34 | -0.0006218
10 | YES | 1250.27 | 34.4325 | 0.00024183 | -0.0036926 | -0.0013735 | 0.0023191 |  | 1000.73 | -0.0005455
10 | YES | 1000.01 | 34.3489 | 0.00026351 | -0.0014572 | -0.0014859 | -2.8677e-05 |  | 5.2854 | -0.0014267
10 | YES | 800.236 | 34.2654 | 0.00023217 | -0.00098801 | -0.0015756 | -0.00058757 |  | 3498.78 | -0.0017242
11 | YES | 5.1104 | 32.3045 | 0.00022018 | -0.00022888 | -0.0019325 | -0.0017037 |  | 3498.78 | -0.00042343
12 | YES | 3500.36 | 34.6668 | 0.0002296 | 0.00086975 | -0.00036335 | -0.0012331 |  | 2999.64 | 0.00040054
12 | YES | 3001.04 | 34.6537 | 0.00026184 | -0.00019074 | -0.00058752 | -0.00039679 |  | 2500.15 | -0.00068283
12 | YES | 2499.79 | 34.6269 | 0.00023168 | -0.0032844 | -0.00081256 | 0.0024719 |  | 1998.29 | -0.0028839
12 | YES | 2000.06 | 34.5806 | 0.00027733 | -0.00092316 | -0.0010369 | -0.00011376 |  | 1500.45 | -0.0010605
12 | YES | 1750.73 | 34.539 | 0.00024554 | -0.0016327 | -0.0011488 | 0.00048384 |  | 1249.39 | -0.001461
12 | YES | 1500.78 | 34.4934 | 0.00024538 | -0.0006752 | -0.0012611 | -0.00058586 |  | 1000.91 | -0.00080872
12 | YES | 1250.34 | 34.436 | 0.00029685 | -0.0006218 | -0.0013735 | -0.0007517 |  | 5.9786 | -0.0017052
12 | YES | 1000.73 | 34.36 | 0.00021891 | -0.0005455 | -0.0014856 | -0.00094006 |  | 3500.88 | -0.00043106
13 | YES | 5.2854 | 32.2877 | 0.00028585 | -0.0014267 | -0.0019325 | -0.00050577 |  | 3000.38 | 0.0018616
14 | YES | 3498.78 | 34.6672 | 0.00023424 | -0.0017242 | -0.00036406 | 0.0013602 |  | 2498.99 | 0.00038528
14 | YES | 3498.78 | 34.6672 | 0.00023424 | -0.00042343 | -0.00036406 | 5.9372e-05 |  | 2001.02 | -0.0019531
14 | YES | 2999.64 | 34.6527 | 0.00022954 | 0.00040054 | -0.00058815 | -0.00098869 |  | 1749.07 | -0.0013351
14 | YES | 2500.15 | 34.6235 | 0.0002254 | -0.00068283 | -0.0008124 | -0.00012956 |  | 1502.96 | -0.0017013
14 | YES | 1998.29 | 34.5801 | 0.00022996 | -0.0028839 | -0.0010377 | 0.0018462 |  | 1249.36 | -0.0016594
14 | YES | 1500.45 | 34.4908 | 0.00025156 | -0.0010605 | -0.0012612 | -0.00020072 |  | 1000.24 | -0.0015717
14 | YES | 1249.39 | 34.436 | 0.00028428 | -0.001461 | -0.0013739 | 8.7103e-05 |  | 3500.22 | 0.0018806
14 | YES | 1000.91 | 34.3581 | 0.00026429 | -0.00080872 | -0.0014855 | -0.00067677 |  | 3001.23 | -0.00060272
15 | YES | 5.9786 | 32.2575 | 0.00073425 | -0.0017052 | -0.0019322 | -0.00022699 |  | 2500.87 | -0.0012779
16 | YES | 3500.88 | 34.6669 | 0.00022743 | -0.00043106 | -0.00036312 | 6.7945e-05 |  | 2000.84 | 0.0016556
16 | YES | 3000.38 | 34.6525 | 0.00024883 | 0.0018616 | -0.00058782 | -0.0024494 |  | 1750.7 | -0.0022736
16 | YES | 2498.99 | 34.6256 | 0.00023044 | 0.00038528 | -0.00081292 | -0.0011982 |  | 1500.09 | -0.00058746
16 | YES | 2001.02 | 34.5805 | 0.00027694 | -0.0019531 | -0.0010365 | 0.00091664 |  | 1250.5 | -0.0028229
16 | YES | 1749.07 | 34.545 | 0.00027885 | -0.0013351 | -0.0011496 | 0.00018555 |  | 999.365 | -0.0015297
16 | YES | 1502.96 | 34.4951 | 0.00025505 | -0.0017013 | -0.0012601 | 0.00044126 |  | 4000.71 | 0.00085068
16 | YES | 1249.36 | 34.4372 | 0.00025587 | -0.0016594 | -0.0013739 | 0.00028545 |  | 3462.47 | -0.001339
16 | YES | 1000.24 | 34.3666 | 0.00021767 | -0.0015717 | -0.0014858 | 8.5876e-05 |  | 3000.08 | -0.00059128
17 | YES | 3500.22 | 34.6693 | 0.00029317 | 0.0018806 | -0.00036341 | -0.0022441 |  | 2498.78 | -0.0036888
17 | YES | 3001.23 | 34.6531 | 0.00022428 | -0.00060272 | -0.00058743 | 1.5287e-05 |  | 1999.85 | -0.0035438
17 | YES | 2500.87 | 34.6243 | 0.00022354 | -0.0012779 | -0.00081207 | 0.00046585 |  | 1499.3 | -0.0022964
17 | YES | 2000.84 | 34.5815 | 0.00026157 | 0.0016556 | -0.0010366 | -0.0026921 |  | 1250.26 | -0.0028915
17 | YES | 1750.7 | 34.5449 | 0.00024578 | -0.0022736 | -0.0011489 | 0.0011247 |  | 1000.96 | 0.00036621
17 | YES | 1500.09 | 34.4983 | 0.00024923 | -0.00058746 | -0.0012614 | -0.00067391 |  | 1500.87 | -0.0064621
17 | YES | 1250.5 | 34.4398 | 0.00032597 | -0.0028229 | -0.0013734 | 0.0014495 |  | 1248.71 | -0.0015068
17 | YES | 999.365 | 34.3642 | 0.00031888 | -0.0015297 | -0.0014862 | 4.3513e-05 |  | 1000.69 | -0.0041771
18 | NO | 4.9652 | 32.2814 | 0.0015537 | -0.0021324 | -0.0019326 | 0.00019981 |  | 4.9028 | -0.0048237
19 | YES | 4000.71 | 34.6782 | 0.00021232 | 0.00085068 | -0.00013872 | -0.00098939 |  |  |  | -0.0011793 | 0.0014677
19 | YES | 3462.47 | 34.6682 | 0.00022521 | -0.001339 | -0.00038036 | 0.0009586
19 | YES | 3000.08 | 34.6527 | 0.00024903 | -0.00059128 | -0.00058795 | 3.3268e-06
19 | YES | 2498.78 | 34.6244 | 0.00022564 | -0.0036888 | -0.00081301 | 0.0028758
19 | YES | 1999.85 | 34.5771 | 0.00023575 | -0.0035438 | -0.001037 | 0.0025068
19 | YES | 1499.3 | 34.4961 | 0.00024763 | -0.0022964 | -0.0012617 | 0.0010347
19 | YES | 1250.26 | 34.4375 | 0.00023933 | -0.0028915 | -0.0013735 | 0.001518
19 | YES | 1000.96 | 34.3626 | 0.00041871 | 0.00036621 | -0.0014855 | -0.0018517
20 | YES | 1500.87 | 34.495 | 0.00024533 | -0.0064621 | -0.001261 | 0.0052011
20 | YES | 1248.71 | 34.429 | 0.00026454 | -0.0015068 | -0.0013742 | 0.00013258
20 | YES | 1000.69 | 34.356 | 0.00059579 | -0.0041771 | -0.0014856 | 0.0026915
20 | YES | 4.9028 | 31.1615 | 0.00047108 | -0.0048237 | -0.0019326 | 0.002891
|  | avg std dev |  | 0.0004986
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0054494 |  | X | CALC Y
m= | 2.1726e-05 | 3.0768e-05 |  | Std Dev of Resid |  | 0.0015476 |  | 1 | -0.0016731
b= | -0.0016948 | 0.00038506 |  | Resid Sum of Squares |  | 0.00021794 |  | 20 | -0.0012603
# of Values Extracted = |  |  | 104
# of Values Used in Fit = |  |  | 93
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.0016731 | -0.00022348 |  | 1 | -0.0014496
1 | YES | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.0016731 | -0.00022348 |  | 1 | -0.0014496
1 | YES | 1000.39 | 34.3692 | 0.00026172 | -0.0041122 | -0.0016731 | 0.0024392 |  | 1 | -0.0041122
1 | YES | 800.385 | 34.27 | 0.00025105 | -0.0032005 | -0.0016731 | 0.0015275 |  | 1 | -0.0032005
1 | YES | 4.9435 | 31.7774 | 0.00029579 | -0.0095119 | -0.0016731 | 0.0078389 |  | 1 | -0.0095119
2 | NO | 5.4165 | 31.8866 | 0.0079538 | -0.0062408 | -0.0016513 | 0.0045895 |  | 4 | -0.0017281
3 | NO | 5.0867 | 31.899 | 0.0017521 | -0.01651 | -0.0016296 | 0.01488 |  | 4 | -0.0017929
4 | YES | 2000.61 | 34.5842 | 0.00027139 | -0.0017281 | -0.0016079 | 0.00012018 |  | 4 | -0.0030632
4 | YES | 1500.35 | 34.5107 | 0.00034639 | -0.0017929 | -0.0016079 | 0.00018503 |  | 4 | -0.0025368
4 | YES | 1250.43 | 34.4474 | 0.00025759 | -0.0030632 | -0.0016079 | 0.0014553 |  | 5 | -0.00034714
4 | NO | 1000.36 | 34.3983 | 0.00026999 | -0.0054588 | -0.0016079 | 0.003851 |  | 6 | -0.00078964
4 | YES | 800.791 | 34.3025 | 0.00037756 | -0.0025368 | -0.0016079 | 0.00092889 |  | 6 | -0.00067902
5 | YES | 5.1646 | 31.8274 | 0.00043663 | -0.00034714 | -0.0015862 | -0.001239 |  | 6 | -0.00034714
6 | YES | 2998.4 | 34.6454 | 0.00027236 | -0.00078964 | -0.0015644 | -0.00077478 |  | 6 | -0.0007019
6 | YES | 2999.49 | 34.6454 | 0.00025291 | -0.00067902 | -0.0015644 | -0.00088541 |  | 6 | -0.0020676
6 | YES | 2500.32 | 34.6235 | 0.00028436 | -0.00034714 | -0.0015644 | -0.0012173 |  | 6 | -0.0015984
6 | YES | 1999.72 | 34.588 | 0.00025896 | -0.0007019 | -0.0015644 | -0.00086252 |  | 6 | -0.00087738
6 | YES | 1500.84 | 34.5091 | 0.000263 | -0.0020676 | -0.0015644 | 0.00050315 |  | 6 | 0.00055695
6 | YES | 1251.02 | 34.4529 | 0.00025925 | -0.0015984 | -0.0015644 | 3.3936e-05 |  | 7 | -0.00019074
6 | YES | 999.451 | 34.3676 | 0.00023208 | -0.00087738 | -0.0015644 | -0.00068704 |  | 8 | -0.0010223
6 | YES | 5.2009 | 32.1462 | 0.00023974 | 0.00055695 | -0.0015644 | -0.0021214 |  | 8 | 0.00099564
7 | YES | 4.8037 | 32.2203 | 0.00023888 | -0.00019074 | -0.0015427 | -0.001352 |  | 8 | -0.0010376
8 | YES | 3359.02 | 34.6553 | 0.00027164 | -0.0010223 | -0.001521 | -0.00049863 |  | 8 | 0.00048828
8 | YES | 2999.35 | 34.65 | 0.0002739 | 0.00099564 | -0.001521 | -0.0025166 |  | 8 | -0.0016022
8 | YES | 2500.22 | 34.6292 | 0.00029204 | -0.0010376 | -0.001521 | -0.00048337 |  | 8 | -0.0026932
8 | YES | 2249.12 | 34.6088 | 0.00025415 | 0.00048828 | -0.001521 | -0.0020093 |  | 8 | -0.0014954
8 | YES | 2000.63 | 34.584 | 0.00026102 | -0.0016022 | -0.001521 | 8.1198e-05 |  | 8 | -0.0012093
8 | YES | 1749.83 | 34.5524 | 0.00026912 | -0.0026932 | -0.001521 | 0.0011722 |  | 8 | -0.0015793
8 | YES | 1499.36 | 34.5063 | 0.00038571 | -0.0014954 | -0.001521 | -2.5612e-05 |  | 8 | -0.0007515
8 | YES | 1249.27 | 34.4506 | 0.00033206 | -0.0012093 | -0.001521 | -0.00031171 |  | 8 | -0.002842
8 | YES | 1000.65 | 34.3688 | 0.00020316 | -0.0015793 | -0.001521 | 5.8308e-05 |  | 8 | 2.6703e-05
8 | YES | 799.545 | 34.2813 | 0.00026137 | -0.0007515 | -0.001521 | -0.00076948 |  | 8 | -0.0002861
8 | YES | 600.327 | 34.1431 | 0.00028963 | -0.002842 | -0.001521 | 0.001321 |  | 8 | -0.00071335
8 | YES | 400.437 | 33.9404 | 0.00025309 | 2.6703e-05 | -0.001521 | -0.0015477 |  | 8 | -0.0010529
8 | YES | 299.392 | 33.8284 | 0.00047768 | -0.0002861 | -0.001521 | -0.0012349 |  | 8 | -0.00081253
8 | YES | 249.804 | 33.8061 | 0.00041418 | -0.00071335 | -0.001521 | -0.00080762 |  | 8 | -0.0023155
8 | YES | 200.768 | 33.7699 | 0.00047069 | -0.0010529 | -0.001521 | -0.00046811 |  | 8 | -0.00030899
8 | YES | 174.785 | 33.7305 | 0.00065778 | -0.00081253 | -0.001521 | -0.00070844 |  | 10 | -0.0011749
8 | NO | 150.431 | 33.5993 | 0.00094877 | -0.0081558 | -0.001521 | 0.0066348 |  | 10 | -0.0011749
8 | NO | 125.193 | 33.1446 | 0.0015435 | -0.019547 | -0.001521 | 0.018026 |  | 10 | 0.00022888
8 | NO | 99.5611 | 32.6081 | 0.0013818 | -0.022873 | -0.001521 | 0.021352 |  | 10 | -0.0015602
8 | YES | 75.4152 | 32.3448 | 0.00044928 | -0.0023155 | -0.001521 | 0.00079455 |  | 10 | -0.0030785
8 | NO | 49.9996 | 32.301 | 0.0032722 | 0.00089264 | -0.001521 | -0.0024136 |  | 10 | -0.0015526
8 | NO | 24.4641 | 32.2205 | 0.00091053 | -8.7738e-05 | -0.001521 | -0.0014332 |  | 10 | -0.0038948
8 | YES | 10.6675 | 32.2158 | 0.00018787 | -0.00030899 | -0.001521 | -0.001212 |  | 10 | -0.0016365
8 | NO | 4.8307 | 32.2161 | 0.0015046 | -0.0013008 | -0.001521 | -0.00022016 |  | 10 | -0.0011215
9 | NO | 5.2141 | 32.1473 | 0.0015834 | -0.0006752 | -0.0014992 | -0.00082404 |  | 11 | 0
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.0014775 | -0.00030259 |  | 12 | 0.00039673
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.0014775 | -0.00030259 |  | 12 | -0.00079346
10 | YES | 3001.09 | 34.6524 | 0.00027319 | 0.00022888 | -0.0014775 | -0.0017064 |  | 12 | -0.0037079
10 | YES | 2500.37 | 34.623 | 0.00031571 | -0.0015602 | -0.0014775 | 8.2691e-05 |  | 12 | -0.0013695
10 | YES | 1999.65 | 34.5747 | 0.00022288 | -0.0030785 | -0.0014775 | 0.0016009 |  | 12 | -0.0018997
10 | YES | 1499.41 | 34.4919 | 0.00022583 | -0.0015526 | -0.0014775 | 7.5061e-05 |  | 12 | -0.001049
10 | YES | 1250.27 | 34.4323 | 0.00023651 | -0.0038948 | -0.0014775 | 0.0024173 |  | 12 | -0.00078583
10 | YES | 1000.01 | 34.3488 | 0.00028203 | -0.0016365 | -0.0014775 | 0.00015899 |  | 12 | -0.0006218
10 | YES | 800.236 | 34.2653 | 0.00025735 | -0.0011215 | -0.0014775 | -0.000356 |  | 13 | -0.0010872
11 | YES | 5.1104 | 32.3047 | 0.00022224 | 0 | -0.0014558 | -0.0014558 |  | 14 | -0.0021439
12 | YES | 3500.36 | 34.6663 | 0.00026718 | 0.00039673 | -0.0014341 | -0.0018308 |  | 14 | -0.00084305
12 | YES | 3001.04 | 34.6531 | 0.00024095 | -0.00079346 | -0.0014341 | -0.00064061 |  | 14 | -6.485e-05
12 | YES | 2499.79 | 34.6265 | 0.00026978 | -0.0037079 | -0.0014341 | 0.0022738 |  | 14 | -0.0010567
12 | YES | 2000.06 | 34.5801 | 0.00028982 | -0.0013695 | -0.0014341 | -6.4586e-05 |  | 14 | -0.0031586
12 | YES | 1750.73 | 34.5387 | 0.00024204 | -0.0018997 | -0.0014341 | 0.00046565 |  | 14 | -0.0013695
12 | YES | 1500.78 | 34.493 | 0.0002886 | -0.001049 | -0.0014341 | -0.00038503 |  | 14 | -0.0014267
12 | YES | 1250.34 | 34.4358 | 0.00032793 | -0.00078583 | -0.0014341 | -0.00064824 |  | 14 | -0.0010071
12 | YES | 1000.73 | 34.3599 | 0.00023286 | -0.0006218 | -0.0014341 | -0.00081227 |  | 15 | -0.00098801
13 | YES | 5.2854 | 32.288 | 0.00028753 | -0.0010872 | -0.0014123 | -0.00032515 |  | 16 | -0.00082016
14 | YES | 3498.78 | 34.6668 | 0.00026182 | -0.0021439 | -0.0013906 | 0.00075325 |  | 16 | 0.0013657
14 | YES | 3498.78 | 34.6668 | 0.00026182 | -0.00084305 | -0.0013906 | -0.00054757 |  | 16 | -6.485e-05
14 | YES | 2999.64 | 34.6522 | 0.00023523 | -6.485e-05 | -0.0013906 | -0.0013258 |  | 16 | -0.0022202
14 | YES | 2500.15 | 34.6231 | 0.00027794 | -0.0010567 | -0.0013906 | -0.00033394 |  | 16 | -0.0016365
14 | YES | 1998.29 | 34.5798 | 0.00029967 | -0.0031586 | -0.0013906 | 0.001768 |  | 16 | -0.0018883
14 | YES | 1500.45 | 34.4905 | 0.00026744 | -0.0013695 | -0.0013906 | -2.1134e-05 |  | 16 | -0.0017281
14 | YES | 1249.39 | 34.4361 | 0.00021765 | -0.0014267 | -0.0013906 | 3.6086e-05 |  | 16 | -0.001709
14 | YES | 1000.91 | 34.3579 | 0.00026336 | -0.0010071 | -0.0013906 | -0.00038353 |  | 17 | 0.0014915
15 | YES | 5.9786 | 32.2582 | 0.00052614 | -0.00098801 | -0.0013689 | -0.00038088 |  | 17 | -0.0009613
16 | YES | 3500.88 | 34.6665 | 0.00021474 | -0.00082016 | -0.0013472 | -0.000527 |  | 17 | -0.0016975
16 | YES | 3000.38 | 34.652 | 0.00025523 | 0.0013657 | -0.0013472 | -0.0027128 |  | 17 | 0.0013885
16 | YES | 2498.99 | 34.6251 | 0.00024547 | -6.485e-05 | -0.0013472 | -0.0012823 |  | 17 | -0.0024338
16 | YES | 2001.02 | 34.5803 | 0.00026163 | -0.0022202 | -0.0013472 | 0.00087299 |  | 17 | -0.00061035
16 | YES | 1749.07 | 34.5447 | 0.0002976 | -0.0016365 | -0.0013472 | 0.00028935 |  | 17 | -0.002903
16 | YES | 1502.96 | 34.4949 | 0.00029607 | -0.0018883 | -0.0013472 | 0.00054112 |  | 17 | -0.0015526
16 | YES | 1249.36 | 34.4372 | 0.00023103 | -0.0017281 | -0.0013472 | 0.0003809 |  | 19 | 0.00042343
16 | YES | 1000.24 | 34.3665 | 0.0002581 | -0.001709 | -0.0013472 | 0.00036182 |  | 19 | -0.0014992
17 | YES | 3500.22 | 34.6689 | 0.00022846 | 0.0014915 | -0.0013254 | -0.002817 |  | 19 | -0.00086594
17 | YES | 3001.23 | 34.6527 | 0.00024813 | -0.0009613 | -0.0013254 | -0.00036413 |  | 19 | -0.0038528
17 | YES | 2500.87 | 34.6239 | 0.0002918 | -0.0016975 | -0.0013254 | 0.00037211 |  | 19 | -0.0035477
17 | YES | 2000.84 | 34.5812 | 0.00027682 | 0.0013885 | -0.0013254 | -0.002714 |  | 19 | -0.0023308
17 | YES | 1750.7 | 34.5448 | 0.00029086 | -0.0024338 | -0.0013254 | 0.0011083 |  | 19 | -0.002903
17 | YES | 1500.09 | 34.4983 | 0.00028111 | -0.00061035 | -0.0013254 | -0.00071508 |  | 19 | 0.00049973
17 | YES | 1250.5 | 34.4397 | 0.00030287 | -0.002903 | -0.0013254 | 0.0015775 |  | 20 | -0.0056076
17 | YES | 999.365 | 34.3641 | 0.00036639 | -0.0015526 | -0.0013254 | 0.00022715 |  | 20 | -0.00035858
18 | NO | 4.9652 | 32.2814 | 0.0013545 | -0.002121 | -0.0013037 | 0.00081726 |  | 20 | -0.0029945
19 | YES | 4000.71 | 34.6777 | 0.00024064 | 0.00042343 | -0.001282 | -0.0017054 |  | 20 | -0.0035725
19 | YES | 3462.47 | 34.668 | 0.00024565 | -0.0014992 | -0.001282 | 0.0002172
19 | YES | 3000.08 | 34.6524 | 0.00025231 | -0.00086594 | -0.001282 | -0.00041605
19 | YES | 2498.78 | 34.6242 | 0.00024188 | -0.0038528 | -0.001282 | 0.0025709
19 | YES | 1999.85 | 34.5771 | 0.00029267 | -0.0035477 | -0.001282 | 0.0022657
19 | YES | 1499.3 | 34.4961 | 0.00028586 | -0.0023308 | -0.001282 | 0.0010488
19 | YES | 1250.26 | 34.4375 | 0.00026741 | -0.002903 | -0.001282 | 0.001621
19 | YES | 1000.96 | 34.3627 | 0.00040632 | 0.00049973 | -0.001282 | -0.0017817
20 | YES | 1500.87 | 34.4959 | 0.00028478 | -0.0056076 | -0.0012603 | 0.0043473
20 | YES | 1248.71 | 34.4301 | 0.00026286 | -0.00035858 | -0.0012603 | -0.00090167
20 | YES | 1000.69 | 34.3572 | 0.00050126 | -0.0029945 | -0.0012603 | 0.0017343
20 | YES | 4.9028 | 31.1627 | 0.00066387 | -0.0035725 | -0.0012603 | 0.0023122
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.033985 |  | X | CALC Y
m= | 1.7473e-07 | 2.329e-07 |  | Std Dev of Resid |  | 0.0010249 |  | 10.6675 | -0.0012178
b= | -0.0012197 | 0.00036764 |  | Resid Sum of Squares |  | 1.6806e-05 |  | 3359.02 | -0.00063273
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
8 | YES | 3359.02 | 34.6553 | 0.00027164 | -0.0010223 | -0.00063273 | 0.00038961 |  | 3359.02 | -0.0010223
8 | YES | 2999.35 | 34.65 | 0.0002739 | 0.00099564 | -0.00069557 | -0.0016912 |  | 2999.35 | 0.00099564
8 | YES | 2500.22 | 34.6292 | 0.00029204 | -0.0010376 | -0.00078279 | 0.00025481 |  | 2500.22 | -0.0010376
8 | YES | 2249.12 | 34.6088 | 0.00025415 | 0.00048828 | -0.00082667 | -0.0013149 |  | 2249.12 | 0.00048828
8 | YES | 2000.63 | 34.584 | 0.00026102 | -0.0016022 | -0.00087008 | 0.00073209 |  | 2000.63 | -0.0016022
8 | YES | 1749.83 | 34.5524 | 0.00026912 | -0.0026932 | -0.00091391 | 0.0017793 |  | 1749.83 | -0.0026932
8 | YES | 1499.36 | 34.5063 | 0.00038571 | -0.0014954 | -0.00095767 | 0.00053769 |  | 1499.36 | -0.0014954
8 | YES | 1249.27 | 34.4506 | 0.00033206 | -0.0012093 | -0.0010014 | 0.00020789 |  | 1249.27 | -0.0012093
8 | YES | 1000.65 | 34.3688 | 0.00020316 | -0.0015793 | -0.0010448 | 0.00053447 |  | 1000.65 | -0.0015793
8 | YES | 799.545 | 34.2813 | 0.00026137 | -0.0007515 | -0.00108 | -0.00032846 |  | 799.545 | -0.0007515
8 | YES | 600.327 | 34.1431 | 0.00028963 | -0.002842 | -0.0011148 | 0.0017272 |  | 600.327 | -0.002842
8 | YES | 400.437 | 33.9404 | 0.00025309 | 2.6703e-05 | -0.0011497 | -0.0011764 |  | 400.437 | 2.6703e-05
8 | YES | 299.392 | 33.8284 | 0.00047768 | -0.0002861 | -0.0011673 | -0.00088125 |  | 299.392 | -0.0002861
8 | YES | 249.804 | 33.8061 | 0.00041418 | -0.00071335 | -0.001176 | -0.00046266 |  | 249.804 | -0.00071335
8 | YES | 200.768 | 33.7699 | 0.00047069 | -0.0010529 | -0.0011846 | -0.00013172 |  | 200.768 | -0.0010529
8 | YES | 174.785 | 33.7305 | 0.00065778 | -0.00081253 | -0.0011891 | -0.00037659 |  | 174.785 | -0.00081253
8 | NO | 150.431 | 33.5993 | 0.00094877 | -0.0081558 | -0.0011934 | 0.0069624 |  | 75.4152 | -0.0023155
8 | NO | 125.193 | 33.1446 | 0.0015435 | -0.019547 | -0.0011978 | 0.018349 |  | 10.6675 | -0.00030899 | -0.0010117
8 | NO | 99.5611 | 32.6081 | 0.0013818 | -0.022873 | -0.0012023 | 0.021671
8 | YES | 75.4152 | 32.3448 | 0.00044928 | -0.0023155 | -0.0012065 | 0.001109
8 | NO | 49.9996 | 32.301 | 0.0032722 | 0.00089264 | -0.0012109 | -0.0021036
8 | NO | 24.4641 | 32.2205 | 0.00091053 | -8.7738e-05 | -0.0012154 | -0.0011276
8 | YES | 10.6675 | 32.2158 | 0.00018787 | -0.00030899 | -0.0012178 | -0.00090881
8 | NO | 4.8307 | 32.2161 | 0.0015046 | -0.0013008 | -0.0012188 | 8.1992e-05
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_5
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 5.0822e-08 |  | X | CALC Y
m= | -7.2674e-08 | 3.3793e-05 |  | Std Dev of Resid |  | 0.0016997 |  | 1 | -0.001267
b= | -0.0012669 | 0.00042292 |  | Resid Sum of Squares |  | 0.0002629 |  | 20 | -0.0012683
# of Values Extracted = |  |  | 104
# of Values Used in Fit = |  |  | 93
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1249.41 | 34.4595 | 0.00023282 | -0.0011063 | -0.001267 | -0.0001607 |  | 1 | -0.0011063
1 | YES | 1249.41 | 34.4595 | 0.00023282 | -0.0011063 | -0.001267 | -0.0001607 |  | 1 | -0.0011063
1 | YES | 1000.39 | 34.3692 | 0.0002916 | -0.0040703 | -0.001267 | 0.0028033 |  | 1 | -0.0040703
1 | YES | 800.385 | 34.27 | 0.00027095 | -0.0032158 | -0.001267 | 0.0019488 |  | 1 | -0.0032158
1 | YES | 4.9435 | 31.7775 | 0.00030083 | -0.0094013 | -0.001267 | 0.0081344 |  | 1 | -0.0094013
2 | NO | 5.4165 | 31.8869 | 0.0084602 | -0.0059318 | -0.001267 | 0.0046648 |  | 4 | -0.0011444
3 | NO | 5.0867 | 31.8994 | 0.0023966 | -0.01609 | -0.0012671 | 0.014823 |  | 4 | -0.0015106
4 | YES | 2000.61 | 34.5848 | 0.00022745 | -0.0011444 | -0.0012672 | -0.00012276 |  | 4 | -0.0028496
4 | YES | 1500.35 | 34.511 | 0.00025369 | -0.0015106 | -0.0012672 | 0.00024345 |  | 4 | -0.0027504
4 | YES | 1250.43 | 34.4477 | 0.00027648 | -0.0028496 | -0.0012672 | 0.0015824 |  | 5 | -0.00027084
4 | NO | 1000.36 | 34.3984 | 0.00028179 | -0.0053825 | -0.0012672 | 0.0041154 |  | 6 | -0.00014496
4 | YES | 800.791 | 34.3022 | 0.00035346 | -0.0027504 | -0.0012672 | 0.0014832 |  | 6 | -0.00010681
5 | YES | 5.1646 | 31.8274 | 0.00049417 | -0.00027084 | -0.0012672 | -0.0009964 |  | 6 | 0.00028229
6 | YES | 2998.4 | 34.6461 | 0.00027567 | -0.00014496 | -0.0012673 | -0.0011224 |  | 6 | -0.00014877
6 | YES | 2999.49 | 34.646 | 0.0002608 | -0.00010681 | -0.0012673 | -0.0011605 |  | 6 | -0.0017471
6 | YES | 2500.32 | 34.6241 | 0.00024004 | 0.00028229 | -0.0012673 | -0.0015496 |  | 6 | -0.0014305
6 | YES | 1999.72 | 34.5886 | 0.00024436 | -0.00014877 | -0.0012673 | -0.0011185 |  | 6 | -0.0007782
6 | YES | 1500.84 | 34.5095 | 0.000232 | -0.0017471 | -0.0012673 | 0.00047981 |  | 6 | 0.00080109
6 | YES | 1251.02 | 34.4531 | 0.00025644 | -0.0014305 | -0.0012673 | 0.00016319 |  | 7 | -0.00056076
6 | YES | 999.451 | 34.3677 | 0.00025206 | -0.0007782 | -0.0012673 | -0.00048912 |  | 8 | -0.00032043
6 | YES | 5.2009 | 32.1464 | 0.00020709 | 0.00080109 | -0.0012673 | -0.0020684 |  | 8 | 0.0017662
7 | YES | 4.8037 | 32.2199 | 0.00034732 | -0.00056076 | -0.0012674 | -0.00070663 |  | 8 | -0.00038528
8 | YES | 3359.02 | 34.656 | 0.00023636 | -0.00032043 | -0.0012675 | -0.00094703 |  | 8 | 0.0011177
8 | YES | 2999.35 | 34.6508 | 0.00025676 | 0.0017662 | -0.0012675 | -0.0030337 |  | 8 | -0.00099564
8 | YES | 2500.22 | 34.6298 | 0.00026352 | -0.00038528 | -0.0012675 | -0.00088218 |  | 8 | -0.0022469
8 | YES | 2249.12 | 34.6094 | 0.00025247 | 0.0011177 | -0.0012675 | -0.0023852 |  | 8 | -0.0012665
8 | YES | 2000.63 | 34.5846 | 0.00022898 | -0.00099564 | -0.0012675 | -0.00027183 |  | 8 | -0.0010643
8 | YES | 1749.83 | 34.5529 | 0.00024041 | -0.0022469 | -0.0012675 | 0.0009794 |  | 8 | -0.0013771
8 | YES | 1499.36 | 34.5065 | 0.00028308 | -0.0012665 | -0.0012675 | -9.8432e-07 |  | 8 | -0.00071716
8 | YES | 1249.27 | 34.4507 | 0.0002659 | -0.0010643 | -0.0012675 | -0.00020316 |  | 8 | -0.0031052
8 | YES | 1000.65 | 34.369 | 0.00022357 | -0.0013771 | -0.0012675 | 0.00010965 |  | 8 | -1.5259e-05
8 | YES | 799.545 | 34.2814 | 0.0002601 | -0.00071716 | -0.0012675 | -0.0005503 |  | 8 | -0.00029755
8 | YES | 600.327 | 34.1428 | 0.0002843 | -0.0031052 | -0.0012675 | 0.0018377 |  | 8 | -0.00081635
8 | YES | 400.437 | 33.9404 | 0.0002818 | -1.5259e-05 | -0.0012675 | -0.0012522 |  | 8 | -0.0012169
8 | YES | 299.392 | 33.8284 | 0.00056987 | -0.00029755 | -0.0012675 | -0.00096992 |  | 8 | -0.00057983
8 | YES | 249.804 | 33.806 | 0.00026868 | -0.00081635 | -0.0012675 | -0.00045112 |  | 8 | -0.0025291
8 | YES | 200.768 | 33.7698 | 0.00037694 | -0.0012169 | -0.0012675 | -5.0574e-05 |  | 8 | -0.00027466
8 | YES | 174.785 | 33.7307 | 0.0007299 | -0.00057983 | -0.0012675 | -0.00068763 |  | 10 | -0.00060654
8 | NO | 150.431 | 33.5991 | 0.0012477 | -0.0084229 | -0.0012675 | 0.0071554 |  | 10 | -0.00060654
8 | NO | 125.193 | 33.1429 | 0.0015026 | -0.021229 | -0.0012675 | 0.019961 |  | 10 | 0.00094223
8 | NO | 99.5611 | 32.608 | 0.0014828 | -0.023014 | -0.0012675 | 0.021747 |  | 10 | -0.0010338
8 | YES | 75.4152 | 32.3446 | 0.00037288 | -0.0025291 | -0.0012675 | 0.0012617 |  | 10 | -0.0025864
8 | NO | 49.9996 | 32.3012 | 0.0033732 | 0.0011482 | -0.0012675 | -0.0024157 |  | 10 | -0.0012093
8 | NO | 24.4641 | 32.2203 | 0.0010984 | -0.00033951 | -0.0012675 | -0.00092796 |  | 10 | -0.0036926
8 | YES | 10.6675 | 32.2158 | 0.00019545 | -0.00027466 | -0.0012675 | -0.00099281 |  | 10 | -0.0014572
8 | NO | 4.8307 | 32.2165 | 0.0028287 | -0.00087357 | -0.0012675 | -0.0003939 |  | 10 | -0.00098801
9 | NO | 5.2141 | 32.1472 | 0.0015668 | -0.00079727 | -0.0012675 | -0.00047026 |  | 11 | -0.00022888
10 | YES | 3500.17 | 34.6645 | 0.00025192 | -0.00060654 | -0.0012676 | -0.00066107 |  | 12 | 0.00086975
10 | YES | 3500.17 | 34.6645 | 0.00025192 | -0.00060654 | -0.0012676 | -0.00066107 |  | 12 | -0.00019074
10 | YES | 3001.09 | 34.6531 | 0.00023742 | 0.00094223 | -0.0012676 | -0.0022098 |  | 12 | -0.0032844
10 | YES | 2500.37 | 34.6236 | 0.00024065 | -0.0010338 | -0.0012676 | -0.00023383 |  | 12 | -0.00092316
10 | YES | 1999.65 | 34.5752 | 0.0002517 | -0.0025864 | -0.0012676 | 0.0013188 |  | 12 | -0.0016327
10 | YES | 1499.41 | 34.4923 | 0.00028475 | -0.0012093 | -0.0012676 | -5.835e-05 |  | 12 | -0.0006752
10 | YES | 1250.27 | 34.4325 | 0.00024183 | -0.0036926 | -0.0012676 | 0.002425 |  | 12 | -0.0006218
10 | YES | 1000.01 | 34.3489 | 0.00026351 | -0.0014572 | -0.0012676 | 0.0001896 |  | 12 | -0.0005455
10 | YES | 800.236 | 34.2654 | 0.00023217 | -0.00098801 | -0.0012676 | -0.0002796 |  | 13 | -0.0014267
11 | YES | 5.1104 | 32.3045 | 0.00022018 | -0.00022888 | -0.0012677 | -0.0010388 |  | 14 | -0.0017242
12 | YES | 3500.36 | 34.6668 | 0.0002296 | 0.00086975 | -0.0012678 | -0.0021375 |  | 14 | -0.00042343
12 | YES | 3001.04 | 34.6537 | 0.00026184 | -0.00019074 | -0.0012678 | -0.001077 |  | 14 | 0.00040054
12 | YES | 2499.79 | 34.6269 | 0.00023168 | -0.0032844 | -0.0012678 | 0.0020167 |  | 14 | -0.00068283
12 | YES | 2000.06 | 34.5806 | 0.00027733 | -0.00092316 | -0.0012678 | -0.0003446 |  | 14 | -0.0028839
12 | YES | 1750.73 | 34.539 | 0.00024554 | -0.0016327 | -0.0012678 | 0.00036493 |  | 14 | -0.0010605
12 | YES | 1500.78 | 34.4934 | 0.00024538 | -0.0006752 | -0.0012678 | -0.00059255 |  | 14 | -0.001461
12 | YES | 1250.34 | 34.436 | 0.00029685 | -0.0006218 | -0.0012678 | -0.00064596 |  | 14 | -0.00080872
12 | YES | 1000.73 | 34.36 | 0.00021891 | -0.0005455 | -0.0012678 | -0.00072225 |  | 15 | -0.0017052
13 | YES | 5.2854 | 32.2877 | 0.00028585 | -0.0014267 | -0.0012678 | 0.00015887 |  | 16 | -0.00043106
14 | YES | 3498.78 | 34.6672 | 0.00023424 | -0.0017242 | -0.0012679 | 0.00045634 |  | 16 | 0.0018616
14 | YES | 3498.78 | 34.6672 | 0.00023424 | -0.00042343 | -0.0012679 | -0.00084447 |  | 16 | 0.00038528
14 | YES | 2999.64 | 34.6527 | 0.00022954 | 0.00040054 | -0.0012679 | -0.0016684 |  | 16 | -0.0019531
14 | YES | 2500.15 | 34.6235 | 0.0002254 | -0.00068283 | -0.0012679 | -0.00058507 |  | 16 | -0.0013351
14 | YES | 1998.29 | 34.5801 | 0.00022996 | -0.0028839 | -0.0012679 | 0.001616 |  | 16 | -0.0017013
14 | YES | 1500.45 | 34.4908 | 0.00025156 | -0.0010605 | -0.0012679 | -0.00020741 |  | 16 | -0.0016594
14 | YES | 1249.39 | 34.436 | 0.00028428 | -0.001461 | -0.0012679 | 0.00019313 |  | 16 | -0.0015717
14 | YES | 1000.91 | 34.3581 | 0.00026429 | -0.00080872 | -0.0012679 | -0.00045918 |  | 17 | 0.0018806
15 | YES | 5.9786 | 32.2575 | 0.00073425 | -0.0017052 | -0.001268 | 0.0004372 |  | 17 | -0.00060272
16 | YES | 3500.88 | 34.6669 | 0.00022743 | -0.00043106 | -0.001268 | -0.00083698 |  | 17 | -0.0012779
16 | YES | 3000.38 | 34.6525 | 0.00024883 | 0.0018616 | -0.001268 | -0.0031296 |  | 17 | 0.0016556
16 | YES | 2498.99 | 34.6256 | 0.00023044 | 0.00038528 | -0.001268 | -0.0016533 |  | 17 | -0.0022736
16 | YES | 2001.02 | 34.5805 | 0.00027694 | -0.0019531 | -0.001268 | 0.00068507 |  | 17 | -0.00058746
16 | YES | 1749.07 | 34.545 | 0.00027885 | -0.0013351 | -0.001268 | 6.7094e-05 |  | 17 | -0.0028229
16 | YES | 1502.96 | 34.4951 | 0.00025505 | -0.0017013 | -0.001268 | 0.0004333 |  | 17 | -0.0015297
16 | YES | 1249.36 | 34.4372 | 0.00025587 | -0.0016594 | -0.001268 | 0.00039134 |  | 19 | 0.00085068
16 | YES | 1000.24 | 34.3666 | 0.00021767 | -0.0015717 | -0.001268 | 0.00030361 |  | 19 | -0.001339
17 | YES | 3500.22 | 34.6693 | 0.00029317 | 0.0018806 | -0.0012681 | -0.0031488 |  | 19 | -0.00059128
17 | YES | 3001.23 | 34.6531 | 0.00022428 | -0.00060272 | -0.0012681 | -0.0006654 |  | 19 | -0.0036888
17 | YES | 2500.87 | 34.6243 | 0.00022354 | -0.0012779 | -0.0012681 | 9.8016e-06 |  | 19 | -0.0035438
17 | YES | 2000.84 | 34.5815 | 0.00026157 | 0.0016556 | -0.0012681 | -0.0029237 |  | 19 | -0.0022964
17 | YES | 1750.7 | 34.5449 | 0.00024578 | -0.0022736 | -0.0012681 | 0.0010054 |  | 19 | -0.0028915
17 | YES | 1500.09 | 34.4983 | 0.00024923 | -0.00058746 | -0.0012681 | -0.00068066 |  | 19 | 0.00036621
17 | YES | 1250.5 | 34.4398 | 0.00032597 | -0.0028229 | -0.0012681 | 0.0015548 |  | 20 | -0.0064621
17 | YES | 999.365 | 34.3642 | 0.00031888 | -0.0015297 | -0.0012681 | 0.00026157 |  | 20 | -0.0015068
18 | NO | 4.9652 | 32.2814 | 0.0015537 | -0.0021324 | -0.0012682 | 0.00086423 |  | 20 | -0.0041771
19 | YES | 4000.71 | 34.6782 | 0.00021232 | 0.00085068 | -0.0012683 | -0.0021189 |  | 20 | -0.0048237
19 | YES | 3462.47 | 34.6682 | 0.00022521 | -0.001339 | -0.0012683 | 7.0696e-05
19 | YES | 3000.08 | 34.6527 | 0.00024903 | -0.00059128 | -0.0012683 | -0.00067699
19 | YES | 2498.78 | 34.6244 | 0.00022564 | -0.0036888 | -0.0012683 | 0.0024205
19 | YES | 1999.85 | 34.5771 | 0.00023575 | -0.0035438 | -0.0012683 | 0.0022756
19 | YES | 1499.3 | 34.4961 | 0.00024763 | -0.0022964 | -0.0012683 | 0.0010282
19 | YES | 1250.26 | 34.4375 | 0.00023933 | -0.0028915 | -0.0012683 | 0.0016233
19 | YES | 1000.96 | 34.3626 | 0.00041871 | 0.00036621 | -0.0012683 | -0.0016345
20 | YES | 1500.87 | 34.495 | 0.00024533 | -0.0064621 | -0.0012683 | 0.0051938
20 | YES | 1248.71 | 34.429 | 0.00026454 | -0.0015068 | -0.0012683 | 0.00023847
20 | YES | 1000.69 | 34.356 | 0.00059579 | -0.0041771 | -0.0012683 | 0.0029088
20 | YES | 4.9028 | 31.1615 | 0.00047108 | -0.0048237 | -0.0012683 | 0.0035553
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_6
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.005738 |  | X | CALC Y
m= | -2.9335e-08 | 6.0306e-08 |  | Std Dev of Resid |  | 0.00052681 |  | 4.8037 | -0.00048165
b= | -0.00048151 | 0.0001297 |  | Resid Sum of Squares |  | 1.1379e-05 |  | 4000.71 | -0.00059887
# of Values Extracted = |  |  | 104
# of Values Used in Fit = |  |  | 43
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.00051816 | 0.00093142 |  | 5.1646 | -0.00034714
1 | NO | 1249.41 | 34.4592 | 0.0002493 | -0.0014496 | -0.00051816 | 0.00093142 |  | 2998.4 | -0.00078964
1 | NO | 1000.39 | 34.3692 | 0.00026172 | -0.0041122 | -0.00051086 | 0.0036014 |  | 2999.49 | -0.00067902
1 | NO | 800.385 | 34.27 | 0.00025105 | -0.0032005 | -0.00050499 | 0.0026955 |  | 2500.32 | -0.00034714
1 | NO | 4.9435 | 31.7774 | 0.00029579 | -0.0095119 | -0.00048166 | 0.0090303 |  | 1999.72 | -0.0007019
2 | NO | 5.4165 | 31.8866 | 0.0079538 | -0.0062408 | -0.00048167 | 0.0057592 |  | 999.451 | -0.00087738
3 | NO | 5.0867 | 31.899 | 0.0017521 | -0.01651 | -0.00048166 | 0.016028 |  | 5.2009 | 0.00055695
4 | NO | 2000.61 | 34.5842 | 0.00027139 | -0.0017281 | -0.0005402 | 0.0011879 |  | 4.8037 | -0.00019074
4 | NO | 1500.35 | 34.5107 | 0.00034639 | -0.0017929 | -0.00052552 | 0.0012674 |  | 3359.02 | -0.0010223
4 | NO | 1250.43 | 34.4474 | 0.00025759 | -0.0030632 | -0.00051819 | 0.002545 |  | 2500.22 | -0.0010376
4 | NO | 1000.36 | 34.3983 | 0.00026999 | -0.0054588 | -0.00051086 | 0.004948 |  | 2249.12 | 0.00048828
4 | NO | 800.791 | 34.3025 | 0.00037756 | -0.0025368 | -0.000505 | 0.0020318 |  | 799.545 | -0.0007515
5 | YES | 5.1646 | 31.8274 | 0.00043663 | -0.00034714 | -0.00048166 | -0.00013453 |  | 400.437 | 2.6703e-05
6 | YES | 2998.4 | 34.6454 | 0.00027236 | -0.00078964 | -0.00056947 | 0.00022017 |  | 299.392 | -0.0002861
6 | YES | 2999.49 | 34.6454 | 0.00025291 | -0.00067902 | -0.0005695 | 0.00010952 |  | 249.804 | -0.00071335
6 | YES | 2500.32 | 34.6235 | 0.00028436 | -0.00034714 | -0.00055486 | -0.00020772 |  | 200.768 | -0.0010529
6 | YES | 1999.72 | 34.588 | 0.00025896 | -0.0007019 | -0.00054017 | 0.00016173 |  | 174.785 | -0.00081253
6 | NO | 1500.84 | 34.5091 | 0.000263 | -0.0020676 | -0.00052554 | 0.001542 |  | 24.4641 | -8.7738e-05
6 | NO | 1251.02 | 34.4529 | 0.00025925 | -0.0015984 | -0.00051821 | 0.0010802 |  | 10.6675 | -0.00030899
6 | YES | 999.451 | 34.3676 | 0.00023208 | -0.00087738 | -0.00051083 | 0.00036655 |  | 3500.17 | -0.0011749
6 | YES | 5.2009 | 32.1462 | 0.00023974 | 0.00055695 | -0.00048166 | -0.0010386 |  | 3500.17 | -0.0011749
7 | YES | 4.8037 | 32.2203 | 0.00023888 | -0.00019074 | -0.00048165 | -0.00029092 |  | 3001.09 | 0.00022888
8 | YES | 3359.02 | 34.6553 | 0.00027164 | -0.0010223 | -0.00058005 | 0.00044229 |  | 800.236 | -0.0011215
8 | NO | 2999.35 | 34.65 | 0.0002739 | 0.00099564 | -0.0005695 | -0.0015651 |  | 5.1104 | 0
8 | YES | 2500.22 | 34.6292 | 0.00029204 | -0.0010376 | -0.00055485 | 0.00048275 |  | 3500.36 | 0.00039673
8 | YES | 2249.12 | 34.6088 | 0.00025415 | 0.00048828 | -0.00054749 | -0.0010358 |  | 3001.04 | -0.00079346
8 | NO | 2000.63 | 34.584 | 0.00026102 | -0.0016022 | -0.0005402 | 0.001062 |  | 1500.78 | -0.001049
8 | NO | 1749.83 | 34.5524 | 0.00026912 | -0.0026932 | -0.00053284 | 0.0021603 |  | 1250.34 | -0.00078583
8 | NO | 1499.36 | 34.5063 | 0.00038571 | -0.0014954 | -0.00052549 | 0.00096987 |  | 1000.73 | -0.0006218
8 | NO | 1249.27 | 34.4506 | 0.00033206 | -0.0012093 | -0.00051816 | 0.0006911 |  | 5.2854 | -0.0010872
8 | NO | 1000.65 | 34.3688 | 0.00020316 | -0.0015793 | -0.00051087 | 0.0010684 |  | 3498.78 | -0.00084305
8 | YES | 799.545 | 34.2813 | 0.00026137 | -0.0007515 | -0.00050497 | 0.00024653 |  | 2999.64 | -6.485e-05
8 | NO | 600.327 | 34.1431 | 0.00028963 | -0.002842 | -0.00049912 | 0.0023428 |  | 2500.15 | -0.0010567
8 | YES | 400.437 | 33.9404 | 0.00025309 | 2.6703e-05 | -0.00049326 | -0.00051996 |  | 1000.91 | -0.0010071
8 | YES | 299.392 | 33.8284 | 0.00047768 | -0.0002861 | -0.00049029 | -0.00020419 |  | 5.9786 | -0.00098801
8 | YES | 249.804 | 33.8061 | 0.00041418 | -0.00071335 | -0.00048884 | 0.00022451 |  | 3500.88 | -0.00082016
8 | YES | 200.768 | 33.7699 | 0.00047069 | -0.0010529 | -0.0004874 | 0.00056546 |  | 2498.99 | -6.485e-05
8 | YES | 174.785 | 33.7305 | 0.00065778 | -0.00081253 | -0.00048664 | 0.00032589 |  | 3001.23 | -0.0009613
8 | NO | 150.431 | 33.5993 | 0.00094877 | -0.0081558 | -0.00048592 | 0.0076699 |  | 1500.09 | -0.00061035
8 | NO | 125.193 | 33.1446 | 0.0015435 | -0.019547 | -0.00048518 | 0.019061 |  | 4000.71 | 0.00042343
8 | NO | 99.5611 | 32.6081 | 0.0013818 | -0.022873 | -0.00048443 | 0.022388 |  | 3000.08 | -0.00086594
8 | NO | 75.4152 | 32.3448 | 0.00044928 | -0.0023155 | -0.00048372 | 0.0018318 |  | 1000.96 | 0.00049973
8 | NO | 49.9996 | 32.301 | 0.0032722 | 0.00089264 | -0.00048298 | -0.0013756 |  | 1248.71 | -0.00035858
8 | YES | 24.4641 | 32.2205 | 0.00091053 | -8.7738e-05 | -0.00048223 | -0.00039449
8 | YES | 10.6675 | 32.2158 | 0.00018787 | -0.00030899 | -0.00048182 | -0.00017283
8 | NO | 4.8307 | 32.2161 | 0.0015046 | -0.0013008 | -0.00048165 | 0.00081916
9 | NO | 5.2141 | 32.1473 | 0.0015834 | -0.0006752 | -0.00048166 | 0.00019354
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.00058419 | 0.00059074
10 | YES | 3500.17 | 34.6639 | 0.00023403 | -0.0011749 | -0.00058419 | 0.00059074
10 | YES | 3001.09 | 34.6524 | 0.00027319 | 0.00022888 | -0.00056955 | -0.00079843
10 | NO | 2500.37 | 34.623 | 0.00031571 | -0.0015602 | -0.00055486 | 0.0010054
10 | NO | 1999.65 | 34.5747 | 0.00022288 | -0.0030785 | -0.00054017 | 0.0025383
10 | NO | 1499.41 | 34.4919 | 0.00022583 | -0.0015526 | -0.0005255 | 0.0010271 |  |  |  | -0.00053104
10 | NO | 1250.27 | 34.4323 | 0.00023651 | -0.0038948 | -0.00051819 | 0.0033766
10 | NO | 1000.01 | 34.3488 | 0.00028203 | -0.0016365 | -0.00051085 | 0.0011257
10 | YES | 800.236 | 34.2653 | 0.00025735 | -0.0011215 | -0.00050499 | 0.00061653
11 | YES | 5.1104 | 32.3047 | 0.00022224 | 0 | -0.00048166 | -0.00048166
12 | YES | 3500.36 | 34.6663 | 0.00026718 | 0.00039673 | -0.00058419 | -0.00098092
12 | YES | 3001.04 | 34.6531 | 0.00024095 | -0.00079346 | -0.00056955 | 0.00022391
12 | NO | 2499.79 | 34.6265 | 0.00026978 | -0.0037079 | -0.00055484 | 0.003153
12 | NO | 2000.06 | 34.5801 | 0.00028982 | -0.0013695 | -0.00054018 | 0.0008293
12 | NO | 1750.73 | 34.5387 | 0.00024204 | -0.0018997 | -0.00053287 | 0.0013669
12 | YES | 1500.78 | 34.493 | 0.0002886 | -0.001049 | -0.00052554 | 0.0005235
12 | YES | 1250.34 | 34.4358 | 0.00032793 | -0.00078583 | -0.00051819 | 0.00026764
12 | YES | 1000.73 | 34.3599 | 0.00023286 | -0.0006218 | -0.00051087 | 0.00011093
13 | YES | 5.2854 | 32.288 | 0.00028753 | -0.0010872 | -0.00048167 | 0.00060552
14 | NO | 3498.78 | 34.6668 | 0.00026182 | -0.0021439 | -0.00058415 | 0.0015597
14 | YES | 3498.78 | 34.6668 | 0.00026182 | -0.00084305 | -0.00058415 | 0.0002589
14 | YES | 2999.64 | 34.6522 | 0.00023523 | -6.485e-05 | -0.00056951 | -0.00050466
14 | YES | 2500.15 | 34.6231 | 0.00027794 | -0.0010567 | -0.00055485 | 0.00050182
14 | NO | 1998.29 | 34.5798 | 0.00029967 | -0.0031586 | -0.00054013 | 0.0026184
14 | NO | 1500.45 | 34.4905 | 0.00026744 | -0.0013695 | -0.00052553 | 0.00084395
14 | NO | 1249.39 | 34.4361 | 0.00021765 | -0.0014267 | -0.00051816 | 0.00090854
14 | YES | 1000.91 | 34.3579 | 0.00026336 | -0.0010071 | -0.00051087 | 0.00049621
15 | YES | 5.9786 | 32.2582 | 0.00052614 | -0.00098801 | -0.00048169 | 0.00050632
16 | YES | 3500.88 | 34.6665 | 0.00021474 | -0.00082016 | -0.00058421 | 0.00023595
16 | NO | 3000.38 | 34.652 | 0.00025523 | 0.0013657 | -0.00056953 | -0.0019352
16 | YES | 2498.99 | 34.6251 | 0.00024547 | -6.485e-05 | -0.00055482 | -0.00048997
16 | NO | 2001.02 | 34.5803 | 0.00026163 | -0.0022202 | -0.00054021 | 0.0016799
16 | NO | 1749.07 | 34.5447 | 0.0002976 | -0.0016365 | -0.00053282 | 0.0011037
16 | NO | 1502.96 | 34.4949 | 0.00029607 | -0.0018883 | -0.0005256 | 0.0013627
16 | NO | 1249.36 | 34.4372 | 0.00023103 | -0.0017281 | -0.00051816 | 0.0012099
16 | NO | 1000.24 | 34.3665 | 0.0002581 | -0.001709 | -0.00051085 | 0.0011981
17 | NO | 3500.22 | 34.6689 | 0.00022846 | 0.0014915 | -0.00058419 | -0.0020757
17 | YES | 3001.23 | 34.6527 | 0.00024813 | -0.0009613 | -0.00056955 | 0.00039175
17 | NO | 2500.87 | 34.6239 | 0.0002918 | -0.0016975 | -0.00055487 | 0.0011427
17 | NO | 2000.84 | 34.5812 | 0.00027682 | 0.0013885 | -0.00054021 | -0.0019288
17 | NO | 1750.7 | 34.5448 | 0.00029086 | -0.0024338 | -0.00053287 | 0.0019009
17 | YES | 1500.09 | 34.4983 | 0.00028111 | -0.00061035 | -0.00052552 | 8.4836e-05
17 | NO | 1250.5 | 34.4397 | 0.00030287 | -0.002903 | -0.00051819 | 0.0023848
17 | NO | 999.365 | 34.3641 | 0.00036639 | -0.0015526 | -0.00051083 | 0.0010418
18 | NO | 4.9652 | 32.2814 | 0.0013545 | -0.002121 | -0.00048166 | 0.0016393
19 | YES | 4000.71 | 34.6777 | 0.00024064 | 0.00042343 | -0.00059887 | -0.0010223
19 | NO | 3462.47 | 34.668 | 0.00024565 | -0.0014992 | -0.00058308 | 0.0009161
19 | YES | 3000.08 | 34.6524 | 0.00025231 | -0.00086594 | -0.00056952 | 0.00029642
19 | NO | 2498.78 | 34.6242 | 0.00024188 | -0.0038528 | -0.00055481 | 0.003298
19 | NO | 1999.85 | 34.5771 | 0.00029267 | -0.0035477 | -0.00054018 | 0.0030075
19 | NO | 1499.3 | 34.4961 | 0.00028586 | -0.0023308 | -0.00052549 | 0.0018053
19 | NO | 1250.26 | 34.4375 | 0.00026741 | -0.002903 | -0.00051819 | 0.0023848
19 | YES | 1000.96 | 34.3627 | 0.00040632 | 0.00049973 | -0.00051087 | -0.0010106
20 | NO | 1500.87 | 34.4959 | 0.00028478 | -0.0056076 | -0.00052554 | 0.0050821
20 | YES | 1248.71 | 34.4301 | 0.00026286 | -0.00035858 | -0.00051814 | -0.00015956
20 | NO | 1000.69 | 34.3572 | 0.00050126 | -0.0029945 | -0.00051087 | 0.0024837
20 | NO | 4.9028 | 31.1627 | 0.00066387 | -0.0035725 | -0.00048166 | 0.0030908
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
YES
YES
NO
NO
NO
NO
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
NO
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
NO
YES
NO
NO
NO
YES
NO
NO
NO
YES
NO
YES
NO
NO
NO
NO
YES
NO
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE