## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2022-021

Agency: Ocean Sciences Division (OSD)

Location: Strait of Georgia / Johnstone Strait / Inlets

Project:  ATP Benthic Recovery Program

Chief Scientist: Sutherland T.

Platform: Vector

Date: 5 October 2022 – 12 October 2022

Processed by: Germaine Gatien

Date of Processing: 9 May 2023 –   15 May 2023

Number of HEX files:  27	Number of CTD casts processed: 27

Number of bottle files: 12	Number of bottle casts processed: 11 (Event 1 had no sampling)

Number of TSG files:  3		Number of TSG files processed: 3

# INSTRUMENT SUMMARY

An SBE911+ CTD #1453 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#3038), a SeaPoint fluorometer (#3949), an SBE pH sensor (#0691), a Biospherical PAR sensor (#4565), a SPAR sensor (#20518) and an altimeter (#76341).

Seasave version 7.26.7.121 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

An IOS rosette with 24 10L bottles was used.

A thermosalinograph (SeaBird 45 S/N 3411) was in use.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format. The header page had complete information about participants and equipment as well as sensor photos. There was also sampling logs for each bottle cast. Both logs included comments about problems encountered.

Waits before firing bottles were 60s or longer.

The difference between salinity sensors was 0.004psu. There were only 5 salinity samples and only 3 below 80db. Using only the 2 deepest bottles the primary salinity was high by an average of 0.0019psu and the secondary was low by 0.0019psu. The sample from 86db had a very low vertical salinity gradient, so may have the best comparison; primary salinity was high by 0.0001psu and secondary low by 0.0041psu. The primary salinity was selected for archiving and is likely within ±0.002psu.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients. For this cruise such errors should be small due to the very low vertical gradients in dissolved oxygen during most casts.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-50db except in areas of very large DO gradients

±0.04 mL/L from 50db-150db

±0.01 mL/L below 150db

There are 3 Thermosalinograph files with gaps between them. There was no flow meter, intake thermistor or loop sampling. A comparison was made of TSG data with 19 CTD casts. The results were close to those from cruise 2022-070 when the same equipment was used and more data were available for comparison. A proxy for intake temperature was derived by subtracting 0.4C° from the lab temperature and salinity was recalibrated by adding 0.055psu., based on 2022-070.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Photos of equipment were included in the log.

- Salinity, Nutrient and Dissolved Oxygen data were obtained in QF spreadsheets. CHL and NH4 were also sampled but will be analyzed by the chief scientist and are not to be included in the bottle files.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained; there were no errors in parameters.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2022-021-ctd.xmlcon.

The ROS files were converted to IOS format; the hysteresis correction was not selected since there were no very deep casts; the Tau correction was selected.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Salinity:T1:C1 was cleaned lightly for bottle #3 at 125db.

The output file was copied to *.BOT.

A preliminary header check was run. The only problem noted is that the pH data from cast #1 are bad. There was no sampling from cast #1, so there will be no bottle file.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2022-021-bot-hdr.txt.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2022-021_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2022-021oxy.csv. That file was converted into individual *.OXY files.

There was 1 sample in the DO file that had comments starting with “ALL:”, but it referred to a case where no sample number was assigned, so it won’t exist in the bottle files.

SALINITY

Salinity analysis was obtained in file QF2022-021 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 13-20 days of collection. The files were simplified and saved as 2022-021sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF_NUTS_2022-021*.xlsx. This includes a precision study. The file was simplified, saved as 2022-021nuts.csv and converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. None were found.

The minimum CTD Salinity in the spreadsheet was 27psu so no correction to silicate samples is required.

A header check was run and no problems were found.

##### Compare

Salinity

This cruise had 5 salinity samples, only 2 of which were below 100db. The deep bottles were well off the bottom of casts.

The primary salinity was found to be lower than bottles by an average of 0.0065psu (std dev 0.0034).

The secondary salinity was found to be lower than bottles by an average of 0.0048psu (std dev 0.0038).

When only the 2 deepest bottles were used the primary salinity was high by an average of 0.0019psu and the secondary was low by 0.0019psu.

Since analysis was done very soon after collection it is expected that there would be minimal errors due to evaporation/desorption. Errors due to incomplete flushing could be significant since the very steady descent rate suggests quiet waters where flushing tends to be impeded, though the 60s wait should help. The size of such errors would depend on the local vertical salinity gradient and those were high for the 2 shallow salinity samples. For cast #47 the local gradient was significantly lower than the deepest 2, so even though sampling was shallower, at 86db, the result there is likely more reliable. Primary salinity was high by 0.0001psu and secondary low by 0.0041psu.

It is clear that the sensors differ by about 0.004psu. While it is likely that the primary salinity is more accurate than the secondary, flushing errors may mean it is actually reading a little high. The secondary appears to be reading low by 0.004psu, but allowing for some flushing errors, it may be more accurate than that.

For more detail see file 2022-021-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is little variation in dissolved oxygen values from the sensor or the bottles (3.8mL/L to 4.7mL/L)

The usual fits to determine DO correction have a wider range of DO. The only data that fall outside the general fit come from the last 2 bottle casts (49 and 53) which were very shallow.

There are gradient reversals in most casts between 30 and 50db which would normally lead to different fits due to different effects of incomplete flushing. But the DO values vary so little that incomplete flushing would have almost no effect, and no such difference was found.

The DO sensor values are lower than titrated samples by about 5.5% which seems unusually large for a sensor that was recently calibrated, but the result is consistent. Any errors due to incomplete flushing or slow response time should be slight given the low variability.

There are no significant outliers.

For more detail see file 2022-021-sal-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using files 2022-021-0001.xmlcon.

The Tau function was selected but not the hysteresis function as there was no deep sampling).

A few casts were examined and all expected channels are present.

- The alignment of the DO traces look good.

- The temperature and conductivity traces are close. The only spikes noted appear to be from shed wake corruption as the CTD stopped for bottles.

- The transmissivity, fluorescence, PAR and SPAR traces look reasonable.

- The buffer bottle was left on the pH sensor during cast 1. Other casts look ok.

- The altimetry looks excellent for most casts but at least 1 cast had a spike at the bottom.

##### WILDEDIT

WILDEDIT won’t help with the sort of spikes noted above, but the routine was run in case there are any single-point spikes in other casts.

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A selection of casts were examined; the temperature profiles were complex but using +2.5s certainly improves the alignment and overall looks like a good choice for the DO sensor.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and those parameters appear to have worked well.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between the channel pairs was examined for a few casts but the differences are all very noisy as expected in shallow sampling. The same was the case on the only other cruise that has used these sensors since they were last serviced. An example of each cruise is reported below.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2022-019-0030 | 200 | -0.0006 | -0.0004 | -0.003 | shallow |
| 2022-021-0009 | 300 | +0.0007 | -0.0004 | -0.004 | steady |

The temperature and conductivity differences are small and within the range seen during 2020-019. Salinity differences are slightly higher and consistent with the bottle comparison which indicated the 2 channels differed by 0.004psu.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A header check was run.

There are no negative fluorescence values; the minimum was 0.1ug/L.

The minimum pressure was -0.44db. Examination of all plots showed the only negative pressures came from an 8s period near the end of the final cast, #67. It is not long enough to establish whether the CTD actually came out of water or if the pressure is reading slightly low. In any case, it is within the specification for the pressure sensor. No other casts had negative pressures.

A cross-reference list was checked and no problems were found.

The track plots look good and were added to the end of this report.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. We usually expect to get values  within ± 5m. There was only one cast outside that limit with a value of 5.4m. The log entry differs by 1m from that in the header and it is likely that there was some variation during the cast. The altimetry  reading varied a little during the stop at the bottom. There is no record of the water depth when the CTD was at the bottom. (In early versions of digital logs, the depth entries were repeated from the initial readings.) No change will be made to the water depth entry which was probably correct when the CTD was at the surface.

##### Shift

Fluorescence   SHIFT was run on the fluorescence channel with setting +24 records. The fluorescence was very noisy making judgement difficult, but cast #1 looked better; this setting is generally found to be the best choice.

Dissolved Oxygen  The Dissolved Oxygen voltage channel was aligned earlier. No further alignment is needed for the DO concentration channel.

Conductivity  This CTD had been used recently on cruise 2022-019 when the best settings found for alignment of conductivity were -0.9 records for the primary sensor and -0.3 records for the secondary. These profiles have such little variations that it is hard to judge, but the 2022-019 values result in much smoother T-S curves.

pH  There is too little variability and to much hysteresis to make a fine-tuned judgement about the alignment of pH with temperature, but the last time the sensor was used in 2022 a shift of +15 records was applied, so SHIFT was run with that setting. The results look ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The P, T, S and DO sensors had been used on 1 other cruise since they were last recalibrated.

- 2022-019 – No calibration sampling; no recalibration applied.

Historic ranges  No local climatology available.

Repeat Casts – The were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving because primary salinity was likely closer to bottles than the secondary salinity and the primary sensors were chosen during previous uses.

All DEL files were copied to *.EDT.

Many of the casts are in areas of active mixing with many unstable features that are likely real.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignments, but in most cases unstable features were left unedited.

Most files required editing near the top and bottom of casts.

No editing was applied to events #6, 9 and 58.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.

##### Recalibration

All salinity in bottle files are >25psu, so silicate does not need recalibration.

Pressure appears to be accurate and will not be recalibrated.

The bottle comparison had few bottles and a lot of scatter, but channel Salinity:T0:C0 appears to be within ±0.002psu. No recalibration will be applied.

Dissolved oxygen needs recalibration by multiplying by 1.055.

File 2022-021-recal1.ccf was prepared to apply a correction as follows:

CTD DO Corrected = CTD DO * 1.055

COMPARE was run and the results show that this correction was suitable.. The CTD DO is higher than bottles by an average of 0.003mL/L (stdev 0.027mL/L) excluding the same bottles as in earlier comparisons.

CALIBRATE was run on the *EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel. It worked well to reduce noise.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from roughly the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

The CTD DO was higher than the titrated samples by an average of 0.006mL/L but the standard deviation was 0.045mL/L. The recalibration was obviously effective.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-50db except in areas of very large DO gradients

±0.04 mL/L from 50db-150db

±0.01 mL/L below 150db

For more detail see file 2022-021-dox-comp3.xls.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

- For cast #1 channel pH:SBE was also removed.

- A second SBE DO channel (with umol/kg units) was derived.

- Oxygen saturation was calculated and surface values were all very low, at about 82% in Saanich Inlet and between 58% and 68% elsewhere.  These casts were well mixed which usually leads to low values, so this offers no evidence about the quality of the DO recalibration.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

- A second SBE DO channel with mass units was added for both the CTD DO and REORDER was run to get the pair of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

- Standards check was run and no problems were found..

- Bottle data were exported to a spreadsheet and no problems were found.

- A header check were run. No problems were found.

- Plots of each file were examined and no problems were found.

- A cross-reference listing was produced for the CHE files.

##### Thermosalinograph processing

There were 3 thermosalinograph files.

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

The configuration file parameters were checked and are correct.

b.) Conversion of Files

The file was converted to CNV format.

The file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. The temperature and salinity look spike-free and have good resolution.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

- The CTD files were thinned to reduce the files to a single point from the downcast within 0.5db of 2db. These were exported to a spreadsheet which was saved as 2022-021-tsg-ctd-comp.xlsx. There were 18 CTD casts with data available that overlapped with TSG data.

- The TSG ATC files were bin-averaged over 5 records (2 minutes) with standard deviations included and then opened in EXCEL. The files were reduced to the times of CTD casts. Those data were added to 2022-021-tsg-ctd-comp.xlss.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were both 0.0000º with a maximum difference of 0.0007º. So the TSG time is accurate.

d.) Comparison of Temperature and Salinity from TSG and CTD data

|  | Median | Std Dev | Median | Std Dev |
| --- | --- | --- | --- | --- |
|  | Ttsg-Tctd | Ttsg-Tctd | SALtsg-SALctd | SALtsg-SALctd |
| All Data | 0.3455 | 0.3842 | -0.7985 | 1.7777 |
| 9 lowest Std Dev in each | 0.3481 | 0.0741 | -0.0723 | 1.8970 |
| 3 Lowest std Dev in each | 0.4466 | 0.0580 | -0.0723 | 0.0314 |

The standard deviations are extremely high when all data are included.

For temperature there is not a large difference when half the casts are excluded based on standard deviations over 2 minutes. Using only the 3 casts with the lowest standard deviation leads to a higher difference, but the scatter in the plot of differences versus std. dev. suggests a slightly lower value.

For salinity the picture is different. The salinity appears to be low by 0.8psu when all data are included but the standard deviation is 1.8psu. Reducing the data set by half produces a much lower difference -0.0723 psu but the standard deviation is even higher. Using only 3 casts produces the same difference, -0.0723psu, but the standard deviations is much lower, at 0.0314psu. The plot of differences against standard deviations does suggest that lower value is reasonable. The 3 casts with the lowest salinity differences were from stations ok1, ok2 and ok3.

The variability in salinity differences could be due to real variability in salinity around 2m. While casts are generally well-mixed, there is still considerable variability in salinity during the surface soak. Temperature is generally less variable. Some of the salinity difference variability may be due to variability in the amount of bubbles in the water in the loop. This could vary in the different areas; that would lower salinity as measured by the TSG in a fairly random way. The casts between 56 and 60 look quite different from other sections; The near-surface vertical gradients were high for those casts with both temperature and salinity increasing with depth. The TSG data are likely coming from a little higher in the water column than values in the averaged 2m CTD files.

During 2022-070 temperatures were found to be high by about 0.4Cº; during 2022-019 it was ~0.38Cº. The earlier cruise was thought to be most reliable, having more data in the comparison. For salinity 2022-070 showed the salinity to be low by ~0.055psu and during 2022-019 it was low by ~0.066psu.

Using the results from 2022-070 looks like the best choice since there were more data available for that cruise.

Checks of the cruise that followed, 2022-081, were very confusing, with only 11 CTD casts in the comparison and large variability in the TSG salinity and in the comparison to CTDs. Lab temperature was higher than CTD temperature by about 0.2Cº. The TSG temperature variability was not unusually high. This difference may be due to a combination of higher flow rate, lower ambient temperature on the ship or differences in the temperature gradients near the surface so that slight differences in the depth of the CTD and TSG data sampling are significant. Some earlier cruises also found heating in the loop to be ~0.2Cº.

See 2022-021-tsg-ctd-comp.xls for details.

Calibration History

This TSG has been used on 6 previous cruises.

- For the first 4 a proxy for intake temperature was created by subtracting 0.2C from the lab temperature. The only cruise with sufficient sampling to estimate salinity calibration was from 2022-017 in June 2022 when salinity was estimated to be low by 0.02psu.

- F0r 2022-070 a proxy for intake temperature was created by subtracting 0.4C from the lab temperature and 0.055 was added to salinity

- For 2022-019 there was insufficient information so the 2022-070 recalibration was applied.

Conclusions

1. The TSG time is accurate.

2. The TSG lab temperature was higher than the CTD temperature by roughly 0.35Cº but variability was high. During 2022-070 when there were more data in the comparison it was found to be high by 0.4Cº, so that is likely the best choice for this cruise as well.

3. The TSG Salinity is low by ~0.073psu when some outliers are excluded. This is close to the results of 2022-070 when it was found to be low by 0.055psu. Using the 2022-070 results looks like the best choice.

f.) Editing

No editing was needed.

REMOVE was used to remove the following channels: Scan Number and Flag channels.

g.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2022-021-tsg-recal1.ccf to subtract 0.40Cº from Temperature:Primary and to add 0.055psu to Salinity.

h.) Preparing Final Files

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars (notes from logs of concern in processing plus few remarks on processing)

TSG:  The TSG was started a few days into the cruise.  It lost data acquisition on one occasion and was restarted.  It was shut down and restarted for tie up and crew change in Nanaimo (shutdown Nov 1, restarted Nov 3).

CTD:

1.Test Cast – all bottles fired at surface. pH sensor data bad – buffer bottle left on.

2. pH sensor data good.

9. Bottle 5 not fired, no sample 62.

47. Ship discharging during cast.

51. Shut down 2m early.

CRUISE SUMMARY – CTD & TSG

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1453 | No | Yes |  |  |
| Calibration Information - 1453 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 2023 | 15Jul2022 | Factory |  |  |  |
| Conductivity |  | 3500 | 19Jul2022 | Factory |  |  |  |
| Secondary Temp. |  | 6565 | 15Jul2022 | Factory |  |  |  |
| Secondary Cond. |  | 5043 | 15Jul2022 | Factory |  |  |  |
| SBE 43 DO sensor |  | 3038 | 26Febl2022 | Factory |  |  |  |
| Altimeter |  | 76341 | 2Oct2021 |  |  |  |  |
| Pressure Sensor |  | 1453 | 21Jul2022 | Factory |  |  |  |
| Transmissometer |  | 983DR | 28Apr2021 | IOS |  |  |  |
| PAR |  | 4565 | 24Feb2021 | Factory |  |  |  |
| SPAR |  | 20518 |  |  |  |  |  |
| pH:SBE |  | 0692 | 14Jan2020 | Factory |  |  |  |
| SeaPoint Fluor. |  | 3949 | ? | Factory |  |  |  |

| Calibration Information – TSG 3363 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 1Dec2019 | Factory |  |  |
| Conductivity | 3411 | 1Dec2019 | Factory |  |  |