## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2022/10/19 11:16:14 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2022-022-0006.mrg | 2022-022-0006.sam | 241 | 5.092 | 5.0917 | 0.05549 | -0.00033188 | 29.4368 | 29.5001 | 0.031019 | 0.063252 | 29.4368 | 29.4956 | 0.024507 | 0.058771
2022-022-0010.mrg | 2022-022-0010.sam | 283 | 5.092 | 4.9106 | 0.20995 | -0.18141 | 31.2727 | 31.2704 | 0.0020751 | -0.0022774 | 31.2727 | 31.2697 | 0.0019146 | -0.0029583
2022-022-0033.mrg | 2022-022-0033.sam | 369 | 5.135 | 4.6322 | 0.28293 | -0.50284 | 31.9646 | 31.9613 | 0.00049542 | -0.0032902 | 31.9646 | 31.9622 | 0.00052184 | -0.0023613
2022-022-0044.mrg | 2022-022-0044.sam | 241 | 1001.34 | 1001.34 | 0.24233 | -0.00018311 | 34.3817 | 34.3824 | 0.00024107 | 0.00074768 | 34.3817 | 34.3815 | 0.00024839 | -0.00019074
2022-022-0044.mrg | 2022-022-0044.sam | 440 | 750.525 | 750.539 | 0.45107 | 0.013611 | 34.3276 | 34.252 | 0.0012985 | -0.075558 | 34.3276 | 34.2509 | 0.0011897 | -0.076721
2022-022-0044.mrg | 2022-022-0044.sam | 241 | 500.196 | 500.196 | 0.45222 | 0.00045776 | 34.0774 | 34.0801 | 0.0002803 | 0.0027084 | 34.0774 | 34.079 | 0.00033108 | 0.0015564
2022-022-0053.mrg | 2022-022-0053.sam | 241 | 1250.1 | 1250.1 | 0.32322 | 0.00012207 | 34.4613 | 34.4658 | 0.00038127 | 0.0044708 | 34.4613 | 34.465 | 0.00035976 | 0.0037193
2022-022-0053.mrg | 2022-022-0053.sam | 241 | 1000.22 | 1000.22 | 0.31019 | -0.00024414 | 34.3824 | 34.3844 | 0.00027006 | 0.0020256 | 34.3824 | 34.3833 | 0.00023675 | 0.00090408
2022-022-0053.mrg | 2022-022-0053.sam | 241 | 750.029 | 750.029 | 0.10594 | 0.00018311 | 34.2675 | 34.2687 | 0.00025506 | 0.0011558 | 34.2675 | 34.2675 | 0.00023431 | 1.1444e-05
2022-022-0053.mrg | 2022-022-0053.sam | 241 | 500.468 | 500.468 | 0.30505 | 0.00012207 | 34.0657 | 34.0684 | 0.00026246 | 0.0026856 | 34.0657 | 34.0672 | 0.0002543 | 0.0015373
2022-022-0053.mrg | 2022-022-0053.sam | 241 | 4.757 | 4.7572 | 0.18738 | 0.00021553 | 32.0131 | 32.0134 | 0.00032488 | 0.00033951 | 32.0131 | 32.0129 | 0.00034754 | -0.00022888
2022-022-0061.mrg | 2022-022-0061.sam | 286 | 4.585 | 4.8455 | 0.14228 | 0.26047 | 31.5459 | 31.5451 | 0.00025207 | -0.00078201 | 31.5459 | 31.5444 | 0.00028812 | -0.0014572
2022-022-0075.mrg | 2022-022-0075.sam | 376 | 5.128 | 5.1087 | 0.046493 | -0.019303 | 31.1503 | 31.1432 | 0.0043529 | -0.0071297 | 31.1503 | 31.1425 | 0.0042406 | -0.0078068
2022-022-0088.mrg | 2022-022-0088.sam | 241 | 5.866 | 5.8663 | 0.053175 | 0.00029039 | 32.0109 | 32.0096 | 0.0003114 | -0.001339 | 32.0109 | 32.0088 | 0.00041243 | -0.0021095
2022-022-0090.mrg | 2022-022-0090.sam | 228 | 750.291 | 750.385 | 0.5715 | 0.093689 | 34.249 | 34.2487 | 0.00028608 | -0.00029373 | 34.249 | 34.2475 | 0.00025185 | -0.0015183
2022-022-0090.mrg | 2022-022-0090.sam | 241 | 500.727 | 500.727 | 0.42067 | -0.00033569 | 34.0992 | 34.1014 | 0.00026083 | 0.0022392 | 34.0992 | 34.1002 | 0.00025942 | 0.0010414
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 1499.68 | 1499.68 | 0.5101 | 0.00024414 | 34.5022 | 34.5047 | 0.00030098 | 0.0025291 | 34.5022 | 34.5037 | 0.00028366 | 0.0015411
2022-022-0095.mrg | 2022-022-0095.sam | 145 | 1249.62 | 1249.65 | 0.61257 | 0.028198 | 34.4677 | 34.4684 | 0.00027987 | 0.00070953 | 34.4677 | 34.4674 | 0.00023759 | -0.00034714
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 999.454 | 999.454 | 0.47236 | -6.1035e-05 | 34.3728 | 34.3729 | 0.00025105 | 9.9182e-05 | 34.3728 | 34.3718 | 0.00024132 | -0.00103
2022-022-0095.mrg | 2022-022-0095.sam | 208 | 749.915 | 749.742 | 0.6297 | -0.17334 | 34.2752 | 34.2775 | 0.00025949 | 0.0023193 | 34.2752 | 34.2762 | 0.00026583 | 0.00099945
2022-022-0095.mrg | 2022-022-0095.sam | 226 | 499.457 | 499.387 | 0.50834 | -0.069702 | 34.0942 | 34.0945 | 0.00045036 | 0.00032425 | 34.0942 | 34.0932 | 0.00041702 | -0.0010223
2022-022-0095.mrg | 2022-022-0095.sam | 179 | 400.778 | 400.824 | 0.60852 | 0.046356 | 33.9811 | 33.9848 | 0.00059672 | 0.0037155 | 33.9811 | 33.9833 | 0.00066962 | 0.0021591
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 299.619 | 299.619 | 0.38464 | -9.1553e-05 | 33.9111 | 33.9131 | 0.00052061 | 0.0019951 | 33.9111 | 33.9118 | 0.00038091 | 0.00067902
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 250.525 | 250.525 | 0.41374 | 0.00033569 | 33.8995 | 33.9019 | 0.00057706 | 0.0023994 | 33.8995 | 33.9002 | 0.00048614 | 0.00074005
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 200.65 | 200.65 | 0.46563 | 0.00028992 | 33.8525 | 33.8554 | 0.00087212 | 0.0028496 | 33.8525 | 33.8542 | 0.00079651 | 0.0016861
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 175.127 | 175.127 | 0.23375 | 0.00048828 | 33.7913 | 33.795 | 0.00045558 | 0.0036507 | 33.7913 | 33.7937 | 0.00065585 | 0.0024414
2022-022-0095.mrg | 2022-022-0095.sam | 237 | 150.293 | 150.311 | 0.59 | 0.017715 | 33.6966 | 33.7011 | 0.00094209 | 0.0044517 | 33.6966 | 33.6997 | 0.00098713 | 0.003128
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 125.603 | 125.603 | 0.56117 | 9.9182e-05 | 33.3481 | 33.361 | 0.0042175 | 0.012859 | 33.3481 | 33.3609 | 0.0040029 | 0.012814
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 100.246 | 100.246 | 0.38292 | 0.00014496 | 32.667 | 32.6645 | 0.0031128 | -0.0025063 | 32.667 | 32.6593 | 0.0029432 | -0.0076599
2022-022-0095.mrg | 2022-022-0095.sam | 204 | 73.93 | 73.8814 | 0.61297 | -0.048599 | 32.2668 | 32.2745 | 0.0022824 | 0.0076904 | 32.2668 | 32.2736 | 0.0018067 | 0.0067787
2022-022-0095.mrg | 2022-022-0095.sam | 445 | 31.044 | 31.0363 | 0.30283 | -0.0077286 | 32.1909 | 32.193 | 0.0026426 | 0.0020905 | 32.1909 | 32.1922 | 0.001607 | 0.0013428
2022-022-0095.mrg | 2022-022-0095.sam | 340 | 20.423 | 20.3233 | 0.41272 | -0.099653 | 32.0898 | 32.0894 | 0.00065749 | -0.00041962 | 32.0898 | 32.0881 | 0.00091529 | -0.0017166
2022-022-0095.mrg | 2022-022-0095.sam | 241 | 10.076 | 10.0765 | 0.1319 | 0.00048542 | 32.089 | 32.0899 | 0.00024317 | 0.00088119 | 32.089 | 32.0889 | 0.00026002 | -5.722e-05
2022-022-0095.mrg | 2022-022-0095.sam | 569 | 4.367 | 4.3306 | 0.38309 | -0.036411 | 32.087 | 32.0899 | 0.00021723 | 0.0029373 | 32.087 | 32.0891 | 0.0002128 | 0.0021477
2022-022-0109.mrg | 2022-022-0109.sam | 224 | 750.127 | 750.049 | 0.55096 | -0.078064 | 34.2657 | 34.2692 | 0.00034323 | 0.0035286 | 34.2657 | 34.2678 | 0.0002984 | 0.0020638
2022-022-0109.mrg | 2022-022-0109.sam | 389 | 4.716 | 4.7741 | 0.14429 | 0.0581 | 31.8741 | 31.8727 | 0.004914 | -0.0014191 | 31.8741 | 31.8719 | 0.0054338 | -0.002203
2022-022-0128.mrg | 2022-022-0128.sam | 140 | 2002.05 | 2001.89 | 0.39395 | -0.16785 | 34.5914 | 34.5919 | 0.00023599 | 0.00045013 | 34.5914 | 34.591 | 0.0002338 | -0.00041199
2022-022-0128.mrg | 2022-022-0128.sam | 213 | 1500.51 | 1500.38 | 0.6946 | -0.13037 | 34.5078 | 34.5078 | 0.00029674 | 3.0518e-05 | 34.5078 | 34.5068 | 0.00025865 | -0.0010414
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 1249.83 | 1249.83 | 0.44017 | -0.00012207 | 34.4546 | 34.4561 | 0.00031667 | 0.0015335 | 34.4546 | 34.4549 | 0.00027924 | 0.0002861
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 1000.36 | 1000.36 | 0.50214 | 0.00012207 | 34.3694 | 34.372 | 0.0003751 | 0.0026169 | 34.3694 | 34.3707 | 0.00035115 | 0.0013313
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 749.532 | 749.532 | 0.25033 | 0.00048828 | 34.2559 | 34.2565 | 0.00025029 | 0.0006218 | 34.2559 | 34.255 | 0.00025497 | -0.00086212
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 500.829 | 500.829 | 0.38202 | 0.00015259 | 34.1004 | 34.0989 | 0.0012261 | -0.0014496 | 34.1004 | 34.0975 | 0.0012083 | -0.0029373
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 401.101 | 401.101 | 0.34821 | 0.00036621 | 33.9572 | 33.9599 | 0.0002548 | 0.0027504 | 33.9572 | 33.9585 | 0.00023118 | 0.001339
2022-022-0128.mrg | 2022-022-0128.sam | 240 | 300.568 | 300.564 | 0.33876 | -0.0042114 | 33.9699 | 33.9707 | 0.00040588 | 0.00082398 | 33.9699 | 33.9694 | 0.00036272 | -0.0005455
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 250.566 | 250.566 | 0.38405 | 0.00045776 | 33.9102 | 33.9122 | 0.00023214 | 0.0020142 | 33.9102 | 33.9108 | 0.0002499 | 0.00063705
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 200.411 | 200.411 | 0.31368 | -1.5259e-05 | 33.8679 | 33.8689 | 0.00083429 | 0.001049 | 33.8679 | 33.8677 | 0.00068013 | -0.00022125
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 173.938 | 173.938 | 0.44916 | -0.00018311 | 33.8279 | 33.8303 | 0.00083463 | 0.0023918 | 33.8279 | 33.8289 | 0.00071323 | 0.00099182
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 151.035 | 151.035 | 0.27017 | -0.00048828 | 33.7458 | 33.7495 | 0.0012737 | 0.003727 | 33.7458 | 33.7485 | 0.0015023 | 0.002655
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 126.065 | 126.065 | 0.19858 | -0.00046539 | 33.573 | 33.5738 | 0.00085326 | 0.0008049 | 33.573 | 33.5728 | 0.00071287 | -0.00020218
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 99.773 | 99.773 | 0.35525 | -3.0518e-05 | 33.3783 | 33.3858 | 0.0029248 | 0.007534 | 33.3783 | 33.3873 | 0.0027723 | 0.008976
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 74.985 | 74.9852 | 0.28992 | 0.00022888 | 32.709 | 32.7255 | 0.0043044 | 0.016499 | 32.709 | 32.7261 | 0.0053973 | 0.01712
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 50.779 | 50.7786 | 0.24205 | -0.00037003 | 32.3707 | 32.3668 | 0.0013768 | -0.0039291 | 32.3707 | 32.3655 | 0.0022926 | -0.0052109
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 30.973 | 30.9731 | 0.21321 | 9.9182e-05 | 32.2033 | 32.2143 | 0.0025061 | 0.01104 | 32.2033 | 32.2143 | 0.0021059 | 0.010963
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 20.664 | 20.6644 | 0.097329 | 0.00043678 | 31.9671 | 32.016 | 0.014803 | 0.048864 | 31.9671 | 32.0143 | 0.020544 | 0.047167
2022-022-0128.mrg | 2022-022-0128.sam | 241 | 10.256 | 10.2557 | 0.30028 | -0.00026894 | 31.7971 | 31.813 | 0.0032499 | 0.015911 | 31.7971 | 31.8119 | 0.0032575 | 0.01479
2022-022-0128.mrg | 2022-022-0128.sam | 209 | 6.197 | 5.8917 | 0.37837 | -0.3053 | 31.7287 | 31.7377 | 0.0036593 | 0.0090332 | 31.7287 | 31.7367 | 0.0032691 | 0.008028
2022-022-0133.mrg | 2022-022-0133.sam | 211 | 1250.92 | 1251.08 | 0.52073 | 0.15601 | 34.4413 | 34.4416 | 0.00032376 | 0.00030518 | 34.4413 | 34.4405 | 0.00030572 | -0.00076675
2022-022-0133.mrg | 2022-022-0133.sam | 241 | 998.656 | 998.656 | 0.50402 | -0.00036621 | 34.358 | 34.3587 | 0.00033116 | 0.00069427 | 34.358 | 34.3575 | 0.00037293 | -0.00045776
2022-022-0135.mrg | 2022-022-0135.sam | 241 | 149.739 | 149.739 | 0.37289 | -4.5776e-05 | 33.9103 | 33.9129 | 0.0003768 | 0.0026398 | 33.9103 | 33.9114 | 0.00039068 | 0.0011368
2022-022-0150.mrg | 2022-022-0150.sam | 128 | 1999.94 | 1999.77 | 0.55876 | -0.17248 | 34.5792 | 34.5806 | 0.00030444 | 0.0013962 | 34.5792 | 34.5795 | 0.00028028 | 0.00029373
2022-022-0150.mrg | 2022-022-0150.sam | 237 | 1500.42 | 1500.44 | 0.44854 | 0.017212 | 34.4956 | 34.4999 | 0.00025156 | 0.0042686 | 34.4956 | 34.4987 | 0.00022971 | 0.0030594
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 1251.4 | 1251.4 | 0.35279 | 0.00012207 | 34.4513 | 34.4524 | 0.00023288 | 0.0011063 | 34.4513 | 34.4511 | 0.00024371 | -0.00016022
2022-022-0150.mrg | 2022-022-0150.sam | 238 | 1001.36 | 1001.35 | 0.6178 | -0.012634 | 34.3742 | 34.3759 | 0.00047311 | 0.0016708 | 34.3742 | 34.3745 | 0.0004914 | 0.0003128
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 749.295 | 749.295 | 0.47213 | 6.1035e-05 | 34.2766 | 34.2794 | 0.0004486 | 0.002758 | 34.2766 | 34.278 | 0.00045187 | 0.0013924
2022-022-0150.mrg | 2022-022-0150.sam | 240 | 500.81 | 500.806 | 0.53073 | -0.0038757 | 34.1071 | 34.1082 | 0.00040348 | 0.0011215 | 34.1071 | 34.1068 | 0.00036409 | -0.00029373
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 400.241 | 400.241 | 0.19731 | -0.00039673 | 34.0145 | 34.0165 | 0.00047953 | 0.0019951 | 34.0145 | 34.015 | 0.00038978 | 0.00052261
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 301.352 | 301.352 | 0.22414 | 6.1035e-05 | 33.9636 | 33.966 | 0.00021884 | 0.0023537 | 33.9636 | 33.9645 | 0.00021672 | 0.00090027
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 250.198 | 250.198 | 0.19396 | 0.00012207 | 33.9134 | 33.9168 | 0.00064921 | 0.0034027 | 33.9134 | 33.9154 | 0.00055711 | 0.0019913
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 199.789 | 199.789 | 0.44031 | -0.00027466 | 33.8586 | 33.8556 | 0.00055152 | -0.0029678 | 33.8586 | 33.8542 | 0.0006487 | -0.0044479
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 175.407 | 175.407 | 0.060676 | -9.1553e-05 | 33.8118 | 33.8122 | 0.00028533 | 0.00044632 | 33.8118 | 33.8108 | 0.00029941 | -0.00099564
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 150.409 | 150.409 | 0.443 | 0 | 33.7836 | 33.7867 | 0.00027021 | 0.0030594 | 33.7836 | 33.7854 | 0.00027093 | 0.0017776
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 125.329 | 125.329 | 0.51493 | -0.00034332 | 33.484 | 33.4837 | 0.00050711 | -0.00028992 | 33.484 | 33.4825 | 0.00058621 | -0.0015221
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 100.306 | 100.306 | 0.12015 | -0.00041199 | 33.0732 | 33.093 | 0.0029244 | 0.019814 | 33.0732 | 33.0891 | 0.0025846 | 0.01585
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 74.642 | 74.6416 | 0.2237 | -0.00041962 | 32.4291 | 32.4332 | 0.004124 | 0.0040703 | 32.4291 | 32.4342 | 0.0058505 | 0.005146
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 50.641 | 50.6408 | 0.49389 | -0.00024414 | 32.358 | 32.3589 | 0.00081204 | 0.00090027 | 32.358 | 32.3576 | 0.00091403 | -0.00037003
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 30.269 | 30.2692 | 0.18772 | 0.00023842 | 32.251 | 32.2694 | 0.006311 | 0.018356 | 32.251 | 32.2681 | 0.0064859 | 0.017124
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 20.311 | 20.3108 | 0.41393 | -0.00018311 | 32.1347 | 32.1467 | 0.0035988 | 0.011955 | 32.1347 | 32.1452 | 0.0045978 | 0.010506
2022-022-0150.mrg | 2022-022-0150.sam | 241 | 10.247 | 10.2471 | 0.26826 | 0.00014973 | 31.4817 | 31.5093 | 0.0096319 | 0.027588 | 31.4817 | 31.5111 | 0.0089727 | 0.029423
2022-022-0150.mrg | 2022-022-0150.sam | 326 | 5.509 | 5.2217 | 0.20138 | -0.28728 | 31.3694 | 31.3807 | 0.0068935 | 0.011328 | 31.3694 | 31.3777 | 0.0069767 | 0.0082607
2022-022-0156.mrg | 2022-022-0156.sam | 204 | 999.477 | 999.275 | 0.55359 | -0.20227 | 34.3517 | 34.3523 | 0.00032353 | 0.00064087 | 34.3517 | 34.351 | 0.00037777 | -0.00072479
2022-022-0156.mrg | 2022-022-0156.sam | 240 | 749.33 | 749.334 | 0.59862 | 0.0042725 | 34.2304 | 34.232 | 0.0005002 | 0.0016365 | 34.2304 | 34.2306 | 0.00029133 | 0.00021362
2022-022-0156.mrg | 2022-022-0156.sam | 241 | 499.328 | 499.328 | 0.19323 | 0.00045776 | 34.0814 | 34.0824 | 0.00028319 | 0.00096893 | 34.0814 | 34.081 | 0.00024391 | -0.00042725
2022-022-0156.mrg | 2022-022-0156.sam | 288 | 4.49 | 4.6359 | 0.1471 | 0.14594 | 31.5871 | 31.5875 | 0.00021633 | 0.00041962 | 31.5871 | 31.5865 | 0.00021485 | -0.00058937
2022-022-0178.mrg | 2022-022-0178.sam | 353 | 5.114 | 5.5254 | 0.23216 | 0.41141 | 30.6996 | 30.7481 | 0.068869 | 0.048506 | 30.6996 | 30.7383 | 0.064462 | 0.038729
2022-022-0204.mrg | 2022-022-0204.sam | 241 | 4.898 | 4.8985 | 0.043238 | 0.00046492 | 30.5909 | 30.5947 | 0.0061611 | 0.0037804 | 30.5909 | 30.5948 | 0.0064746 | 0.0039139
2022-022-0207.mrg | 2022-022-0207.sam | 241 | 299.996 | 299.996 | 0.051788 | -0.00048828 | 30.6812 | 30.6841 | 0.00025665 | 0.002943 | 30.6812 | 30.6829 | 0.00024205 | 0.0016518
2022-022-0207.mrg | 2022-022-0207.sam | 241 | 30.02 | 30.0203 | 0.032539 | 0.0002861 | 28.9906 | 28.9821 | 0.00082899 | -0.008503 | 28.9906 | 28.9807 | 0.00054577 | -0.0099144
2022-022-0226.mrg | 2022-022-0226.sam | 241 | 349.999 | 349.999 | 0.031117 | 0.00027466 | 31.0852 | 31.0898 | 0.00024109 | 0.0045547 | 31.0852 | 31.0885 | 0.00020913 | 0.0033455
2022-022-0226.mrg | 2022-022-0226.sam | 241 | 9.554 | 9.5543 | 0.032726 | 0.00031567 | 28.2908 | 28.2832 | 0.0060495 | -0.007576 | 28.2908 | 28.2843 | 0.0052249 | -0.006546
2022-022-0231.mrg | 2022-022-0231.sam | 241 | 310.386 | 310.386 | 0.028254 | 0.00033569 | 31.0472 | 31.07 | 0.00039056 | 0.022829 | 31.0472 | 31.0688 | 0.00049377 | 0.021595
2022-022-0231.mrg | 2022-022-0231.sam | 241 | 9.736 | 9.7356 | 0.071034 | -0.00041103 | 28.2853 | 28.2449 | 0.0011965 | -0.040424 | 28.2853 | 28.2435 | 0.00048149 | -0.041771
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 21
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2022-022-0006.mrg | 2022-022-0006.sam | 12 | 12 | 2022-022-0006.mrg & 2022-022-0006.sam
2 | 2022-022-0010.mrg | 2022-022-0010.sam | 14 | 14 | 2022-022-0010.mrg & 2022-022-0010.sam
3 | 2022-022-0033.mrg | 2022-022-0033.sam | 16 | 16 | 2022-022-0033.mrg & 2022-022-0033.sam
4 | 2022-022-0044.mrg | 2022-022-0044.sam | 18 | 20 | 2022-022-0044.mrg & 2022-022-0044.sam
5 | 2022-022-0053.mrg | 2022-022-0053.sam | 22 | 26 | 2022-022-0053.mrg & 2022-022-0053.sam
6 | 2022-022-0061.mrg | 2022-022-0061.sam | 28 | 28 | 2022-022-0061.mrg & 2022-022-0061.sam
7 | 2022-022-0075.mrg | 2022-022-0075.sam | 30 | 30 | 2022-022-0075.mrg & 2022-022-0075.sam
8 | 2022-022-0088.mrg | 2022-022-0088.sam | 32 | 32 | 2022-022-0088.mrg & 2022-022-0088.sam
9 | 2022-022-0090.mrg | 2022-022-0090.sam | 34 | 35 | 2022-022-0090.mrg & 2022-022-0090.sam
10 | 2022-022-0095.mrg | 2022-022-0095.sam | 37 | 54 | 2022-022-0095.mrg & 2022-022-0095.sam
11 | 2022-022-0109.mrg | 2022-022-0109.sam | 56 | 57 | 2022-022-0109.mrg & 2022-022-0109.sam
12 | 2022-022-0128.mrg | 2022-022-0128.sam | 59 | 78 | 2022-022-0128.mrg & 2022-022-0128.sam
13 | 2022-022-0133.mrg | 2022-022-0133.sam | 80 | 81 | 2022-022-0133.mrg & 2022-022-0133.sam
14 | 2022-022-0135.mrg | 2022-022-0135.sam | 83 | 83 | 2022-022-0135.mrg & 2022-022-0135.sam
15 | 2022-022-0150.mrg | 2022-022-0150.sam | 85 | 104 | 2022-022-0150.mrg & 2022-022-0150.sam
16 | 2022-022-0156.mrg | 2022-022-0156.sam | 106 | 109 | 2022-022-0156.mrg & 2022-022-0156.sam
17 | 2022-022-0178.mrg | 2022-022-0178.sam | 111 | 111 | 2022-022-0178.mrg & 2022-022-0178.sam
18 | 2022-022-0204.mrg | 2022-022-0204.sam | 113 | 113 | 2022-022-0204.mrg & 2022-022-0204.sam
19 | 2022-022-0207.mrg | 2022-022-0207.sam | 115 | 116 | 2022-022-0207.mrg & 2022-022-0207.sam
20 | 2022-022-0226.mrg | 2022-022-0226.sam | 118 | 119 | 2022-022-0226.mrg & 2022-022-0226.sam
21 | 2022-022-0231.mrg | 2022-022-0231.sam | 121 | 122 | 2022-022-0231.mrg & 2022-022-0231.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 21
2 | (Sal1 - Sal_Bot) | Pressure | 21
3 | (Sal0 - Sal_Bot) | File Pair # | 21
4 | (Sal1 - Sal_Bot) | File Pair # | 21

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0072765 |  | X | CALC Y
m= | 2.6442e-07 | 4.283e-07 |  | Std Dev of Resid |  | 0.0016161 |  | 4.3306 | 0.0012609
b= | 0.0012598 | 0.00034366 |  | Resid Sum of Squares |  | 0.00013582 |  | 2001.89 | 0.0017891
# of Values Extracted = |  |  | 91
# of Values Used in Fit = |  |  | 54
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0917 | 29.5001 | 0.031019 | 0.063252 | 0.0012611 | -0.06199 |  | 4.6322 | -0.0032902
2 | NO | 4.9106 | 31.2704 | 0.0020751 | -0.0022774 | 0.0012611 | 0.0035384 |  | 1001.34 | 0.00074768
3 | YES | 4.6322 | 31.9613 | 0.00049542 | -0.0032902 | 0.001261 | 0.0045512 |  | 500.196 | 0.0027084
4 | YES | 1001.34 | 34.3824 | 0.00024107 | 0.00074768 | 0.0015246 | 0.00077688 |  | 1000.22 | 0.0020256
4 | NO | 750.539 | 34.252 | 0.0012985 | -0.075558 | 0.0014582 | 0.077016 |  | 750.029 | 0.0011558
4 | YES | 500.196 | 34.0801 | 0.0002803 | 0.0027084 | 0.001392 | -0.0013164 |  | 500.468 | 0.0026856
5 | NO | 1250.1 | 34.4658 | 0.00038127 | 0.0044708 | 0.0015903 | -0.0028805 |  | 4.7572 | 0.00033951
5 | YES | 1000.22 | 34.3844 | 0.00027006 | 0.0020256 | 0.0015243 | -0.00050134 |  | 4.8455 | -0.00078201
5 | YES | 750.029 | 34.2687 | 0.00025506 | 0.0011558 | 0.0014581 | 0.00030225 |  | 5.8663 | -0.001339
5 | YES | 500.468 | 34.0684 | 0.00026246 | 0.0026856 | 0.0013921 | -0.0012934 |  | 750.385 | -0.00029373
5 | YES | 4.7572 | 32.0134 | 0.00032488 | 0.00033951 | 0.001261 | 0.00092153 |  | 500.727 | 0.0022392
6 | YES | 4.8455 | 31.5451 | 0.00025207 | -0.00078201 | 0.0012611 | 0.0020431 |  | 1499.68 | 0.0025291
7 | NO | 5.1087 | 31.1432 | 0.0043529 | -0.0071297 | 0.0012611 | 0.0083908 |  | 1249.65 | 0.00070953
8 | YES | 5.8663 | 32.0096 | 0.0003114 | -0.001339 | 0.0012613 | 0.0026003 |  | 999.454 | 9.9182e-05
9 | YES | 750.385 | 34.2487 | 0.00028608 | -0.00029373 | 0.0014582 | 0.0017519 |  | 749.742 | 0.0023193
9 | YES | 500.727 | 34.1014 | 0.00026083 | 0.0022392 | 0.0013922 | -0.00084705 |  | 499.387 | 0.00032425
10 | YES | 1499.68 | 34.5047 | 0.00030098 | 0.0025291 | 0.0016563 | -0.00087281 |  | 400.824 | 0.0037155
10 | YES | 1249.65 | 34.4684 | 0.00027987 | 0.00070953 | 0.0015902 | 0.00088068 |  | 299.619 | 0.0019951
10 | YES | 999.454 | 34.3729 | 0.00025105 | 9.9182e-05 | 0.0015241 | 0.0014249 |  | 250.525 | 0.0023994
10 | YES | 749.742 | 34.2775 | 0.00025949 | 0.0023193 | 0.001458 | -0.00086131 |  | 175.127 | 0.0036507
10 | YES | 499.387 | 34.0945 | 0.00045036 | 0.00032425 | 0.0013918 | 0.0010676 |  | 20.3233 | -0.00041962
10 | YES | 400.824 | 33.9848 | 0.00059672 | 0.0037155 | 0.0013658 | -0.0023498 |  | 10.0765 | 0.00088119
10 | YES | 299.619 | 33.9131 | 0.00052061 | 0.0019951 | 0.001339 | -0.00065608 |  | 4.3306 | 0.0029373
10 | YES | 250.525 | 33.9019 | 0.00057706 | 0.0023994 | 0.001326 | -0.0010734 |  | 750.049 | 0.0035286
10 | NO | 200.65 | 33.8554 | 0.00087212 | 0.0028496 | 0.0013128 | -0.0015367 |  | 2001.89 | 0.00045013
10 | YES | 175.127 | 33.795 | 0.00045558 | 0.0036507 | 0.0013061 | -0.0023446 |  | 1500.38 | 3.0518e-05
10 | NO | 150.311 | 33.7011 | 0.00094209 | 0.0044517 | 0.0012995 | -0.0031522 |  | 1249.83 | 0.0015335
10 | NO | 125.603 | 33.361 | 0.0042175 | 0.012859 | 0.001293 | -0.011566 |  | 1000.36 | 0.0026169
10 | NO | 100.246 | 32.6645 | 0.0031128 | -0.0025063 | 0.0012863 | 0.0037925 |  | 749.532 | 0.0006218
10 | NO | 73.8814 | 32.2745 | 0.0022824 | 0.0076904 | 0.0012793 | -0.0064111 |  | 401.101 | 0.0027504
10 | NO | 31.0363 | 32.193 | 0.0026426 | 0.0020905 | 0.001268 | -0.00082246 |  | 300.564 | 0.00082398
10 | YES | 20.3233 | 32.0894 | 0.00065749 | -0.00041962 | 0.0012652 | 0.0016848 |  | 250.566 | 0.0020142
10 | YES | 10.0765 | 32.0899 | 0.00024317 | 0.00088119 | 0.0012624 | 0.00038125 |  | 1251.08 | 0.00030518
10 | YES | 4.3306 | 32.0899 | 0.00021723 | 0.0029373 | 0.0012609 | -0.0016764 |  | 998.656 | 0.00069427
11 | YES | 750.049 | 34.2692 | 0.00034323 | 0.0035286 | 0.0014581 | -0.0020705 |  | 149.739 | 0.0026398
11 | NO | 4.7741 | 31.8727 | 0.004914 | -0.0014191 | 0.001261 | 0.0026801 |  | 1999.77 | 0.0013962
12 | YES | 2001.89 | 34.5919 | 0.00023599 | 0.00045013 | 0.0017891 | 0.001339 |  | 1500.44 | 0.0042686
12 | YES | 1500.38 | 34.5078 | 0.00029674 | 3.0518e-05 | 0.0016565 | 0.001626 |  | 1251.4 | 0.0011063
12 | YES | 1249.83 | 34.4561 | 0.00031667 | 0.0015335 | 0.0015903 | 5.6753e-05 |  | 1001.35 | 0.0016708
12 | YES | 1000.36 | 34.372 | 0.0003751 | 0.0026169 | 0.0015243 | -0.0010926 |  | 749.295 | 0.002758
12 | YES | 749.532 | 34.2565 | 0.00025029 | 0.0006218 | 0.001458 | 0.00083618 |  | 500.806 | 0.0011215
12 | NO | 500.829 | 34.0989 | 0.0012261 | -0.0014496 | 0.0013922 | 0.0028418 |  | 400.241 | 0.0019951
12 | YES | 401.101 | 33.9599 | 0.0002548 | 0.0027504 | 0.0013658 | -0.0013846 |  | 301.352 | 0.0023537
12 | YES | 300.564 | 33.9707 | 0.00040588 | 0.00082398 | 0.0013393 | 0.00051528 |  | 250.198 | 0.0034027
12 | YES | 250.566 | 33.9122 | 0.00023214 | 0.0020142 | 0.001326 | -0.00068813 |  | 199.789 | -0.0029678
12 | NO | 200.411 | 33.8689 | 0.00083429 | 0.001049 | 0.0013128 | 0.00026373 |  | 175.407 | 0.00044632
12 | NO | 173.938 | 33.8303 | 0.00083463 | 0.0023918 | 0.0013058 | -0.001086 |  | 150.409 | 0.0030594
12 | NO | 151.035 | 33.7495 | 0.0012737 | 0.003727 | 0.0012997 | -0.0024272 |  | 125.329 | -0.00028992
12 | NO | 126.065 | 33.5738 | 0.00085326 | 0.0008049 | 0.0012931 | 0.00048821 |  | 999.275 | 0.00064087
12 | NO | 99.773 | 33.3858 | 0.0029248 | 0.007534 | 0.0012862 | -0.0062479 |  | 749.334 | 0.0016365
12 | NO | 74.9852 | 32.7255 | 0.0043044 | 0.016499 | 0.0012796 | -0.015219 |  | 499.328 | 0.00096893
12 | NO | 50.7786 | 32.3668 | 0.0013768 | -0.0039291 | 0.0012732 | 0.0052023 |  | 4.6359 | 0.00041962
12 | NO | 30.9731 | 32.2143 | 0.0025061 | 0.01104 | 0.001268 | -0.0097717 |  | 299.996 | 0.002943
12 | NO | 20.6644 | 32.016 | 0.014803 | 0.048864 | 0.0012652 | -0.047599 |  | 349.999 | 0.0045547
12 | NO | 10.2557 | 31.813 | 0.0032499 | 0.015911 | 0.0012625 | -0.014649
12 | NO | 5.8917 | 31.7377 | 0.0036593 | 0.0090332 | 0.0012613 | -0.0077719
13 | YES | 1251.08 | 34.4416 | 0.00032376 | 0.00030518 | 0.0015906 | 0.0012854
13 | YES | 998.656 | 34.3587 | 0.00033116 | 0.00069427 | 0.0015238 | 0.00082957
14 | YES | 149.739 | 33.9129 | 0.0003768 | 0.0026398 | 0.0012994 | -0.0013404
15 | YES | 1999.77 | 34.5806 | 0.00030444 | 0.0013962 | 0.0017886 | 0.00039238
15 | YES | 1500.44 | 34.4999 | 0.00025156 | 0.0042686 | 0.0016565 | -0.0026121
15 | YES | 1251.4 | 34.4524 | 0.00023288 | 0.0011063 | 0.0015907 | 0.00048442 |  |  |  | 0.0014228 | 0.0016067
15 | YES | 1001.35 | 34.3759 | 0.00047311 | 0.0016708 | 0.0015246 | -0.00014628
15 | YES | 749.295 | 34.2794 | 0.0004486 | 0.002758 | 0.0014579 | -0.0013001
15 | YES | 500.806 | 34.1082 | 0.00040348 | 0.0011215 | 0.0013922 | 0.00027068
15 | YES | 400.241 | 34.0165 | 0.00047953 | 0.0019951 | 0.0013656 | -0.00062948
15 | YES | 301.352 | 33.966 | 0.00021884 | 0.0023537 | 0.0013395 | -0.0010142
15 | YES | 250.198 | 33.9168 | 0.00064921 | 0.0034027 | 0.0013259 | -0.0020768
15 | YES | 199.789 | 33.8556 | 0.00055152 | -0.0029678 | 0.0013126 | 0.0042804
15 | YES | 175.407 | 33.8122 | 0.00028533 | 0.00044632 | 0.0013062 | 0.00085984
15 | YES | 150.409 | 33.7867 | 0.00027021 | 0.0030594 | 0.0012996 | -0.0017598
15 | YES | 125.329 | 33.4837 | 0.00050711 | -0.00028992 | 0.0012929 | 0.0015828
15 | NO | 100.306 | 33.093 | 0.0029244 | 0.019814 | 0.0012863 | -0.018527
15 | NO | 74.6416 | 32.4332 | 0.004124 | 0.0040703 | 0.0012795 | -0.0027908
15 | NO | 50.6408 | 32.3589 | 0.00081204 | 0.00090027 | 0.0012732 | 0.0003729
15 | NO | 30.2692 | 32.2694 | 0.006311 | 0.018356 | 0.0012678 | -0.017089
15 | NO | 20.3108 | 32.1467 | 0.0035988 | 0.011955 | 0.0012652 | -0.01069
15 | NO | 10.2471 | 31.5093 | 0.0096319 | 0.027588 | 0.0012625 | -0.026325
15 | NO | 5.2217 | 31.3807 | 0.0068935 | 0.011328 | 0.0012612 | -0.010067
16 | YES | 999.275 | 34.3523 | 0.00032353 | 0.00064087 | 0.001524 | 0.00088314
16 | YES | 749.334 | 34.232 | 0.0005002 | 0.0016365 | 0.0014579 | -0.00017859
16 | YES | 499.328 | 34.0824 | 0.00028319 | 0.00096893 | 0.0013918 | 0.00042288
16 | YES | 4.6359 | 31.5875 | 0.00021633 | 0.00041962 | 0.001261 | 0.00084139
17 | NO | 5.5254 | 30.7481 | 0.068869 | 0.048506 | 0.0012612 | -0.047245
18 | NO | 4.8985 | 30.5947 | 0.0061611 | 0.0037804 | 0.0012611 | -0.0025193
19 | YES | 299.996 | 30.6841 | 0.00025665 | 0.002943 | 0.0013391 | -0.0016039
19 | NO | 30.0203 | 28.9821 | 0.00082899 | -0.008503 | 0.0012677 | 0.0097707
20 | YES | 349.999 | 31.0898 | 0.00024109 | 0.0045547 | 0.0013523 | -0.0032024
20 | NO | 9.5543 | 28.2832 | 0.0060495 | -0.007576 | 0.0012623 | 0.0088383
21 | NO | 310.386 | 31.07 | 0.00039056 | 0.022829 | 0.0013419 | -0.021487
21 | NO | 9.7356 | 28.2449 | 0.0011965 | -0.040424 | 0.0012624 | 0.041687
0.0025542
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0012578 |  | X | CALC Y
m= | 9.8401e-08 | 3.7389e-07 |  | Std Dev of Resid |  | 0.0014307 |  | 4.3306 | 0.00023474
b= | 0.00023431 | 0.00029253 |  | Resid Sum of Squares |  | 0.00011257 |  | 2001.89 | 0.0004313
# of Values Extracted = |  |  | 91
# of Values Used in Fit = |  |  | 57
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0917 | 29.4956 | 0.024507 | 0.058771 | 0.00023481 | -0.058536 |  | 4.6322 | -0.0023613
2 | NO | 4.9106 | 31.2697 | 0.0019146 | -0.0029583 | 0.00023479 | 0.0031931 |  | 1001.34 | -0.00019074
3 | YES | 4.6322 | 31.9622 | 0.00052184 | -0.0023613 | 0.00023477 | 0.0025961 |  | 500.196 | 0.0015564
4 | YES | 1001.34 | 34.3815 | 0.00024839 | -0.00019074 | 0.00033284 | 0.00052358 |  | 1000.22 | 0.00090408
4 | NO | 750.539 | 34.2509 | 0.0011897 | -0.076721 | 0.00030816 | 0.077029 |  | 750.029 | 1.1444e-05
4 | YES | 500.196 | 34.079 | 0.00033108 | 0.0015564 | 0.00028353 | -0.0012729 |  | 500.468 | 0.0015373
5 | NO | 1250.1 | 34.465 | 0.00035976 | 0.0037193 | 0.00035732 | -0.003362 |  | 4.7572 | -0.00022888
5 | YES | 1000.22 | 34.3833 | 0.00023675 | 0.00090408 | 0.00033273 | -0.00057135 |  | 4.8455 | -0.0014572
5 | YES | 750.029 | 34.2675 | 0.00023431 | 1.1444e-05 | 0.00030811 | 0.00029667 |  | 5.8663 | -0.0021095
5 | YES | 500.468 | 34.0672 | 0.0002543 | 0.0015373 | 0.00028356 | -0.0012538 |  | 750.385 | -0.0015183
5 | YES | 4.7572 | 32.0129 | 0.00034754 | -0.00022888 | 0.00023478 | 0.00046366 |  | 500.727 | 0.0010414
6 | YES | 4.8455 | 31.5444 | 0.00028812 | -0.0014572 | 0.00023479 | 0.001692 |  | 1499.68 | 0.0015411
7 | NO | 5.1087 | 31.1425 | 0.0042406 | -0.0078068 | 0.00023481 | 0.0080416 |  | 1249.65 | -0.00034714
8 | YES | 5.8663 | 32.0088 | 0.00041243 | -0.0021095 | 0.00023489 | 0.0023444 |  | 999.454 | -0.00103
9 | YES | 750.385 | 34.2475 | 0.00025185 | -0.0015183 | 0.00030815 | 0.0018264 |  | 749.742 | 0.00099945
9 | YES | 500.727 | 34.1002 | 0.00025942 | 0.0010414 | 0.00028358 | -0.00075783 |  | 499.387 | -0.0010223
10 | YES | 1499.68 | 34.5037 | 0.00028366 | 0.0015411 | 0.00038188 | -0.0011593 |  | 400.824 | 0.0021591
10 | YES | 1249.65 | 34.4674 | 0.00023759 | -0.00034714 | 0.00035728 | 0.00070441 |  | 299.619 | 0.00067902
10 | YES | 999.454 | 34.3718 | 0.00024132 | -0.00103 | 0.00033266 | 0.0013626 |  | 250.525 | 0.00074005
10 | YES | 749.742 | 34.2762 | 0.00026583 | 0.00099945 | 0.00030809 | -0.00069137 |  | 200.65 | 0.0016861
10 | YES | 499.387 | 34.0932 | 0.00041702 | -0.0010223 | 0.00028345 | 0.0013058 |  | 175.127 | 0.0024414
10 | YES | 400.824 | 33.9833 | 0.00066962 | 0.0021591 | 0.00027375 | -0.0018854 |  | 10.0765 | -5.722e-05
10 | YES | 299.619 | 33.9118 | 0.00038091 | 0.00067902 | 0.00026379 | -0.00041522 |  | 4.3306 | 0.0021477
10 | YES | 250.525 | 33.9002 | 0.00048614 | 0.00074005 | 0.00025896 | -0.00048109 |  | 750.049 | 0.0020638
10 | YES | 200.65 | 33.8542 | 0.00079651 | 0.0016861 | 0.00025405 | -0.001432 |  | 2001.89 | -0.00041199
10 | YES | 175.127 | 33.7937 | 0.00065585 | 0.0024414 | 0.00025154 | -0.0021899 |  | 1500.38 | -0.0010414
10 | NO | 150.311 | 33.6997 | 0.00098713 | 0.003128 | 0.0002491 | -0.0028789 |  | 1249.83 | 0.0002861
10 | NO | 125.603 | 33.3609 | 0.0040029 | 0.012814 | 0.00024667 | -0.012567 |  | 1000.36 | 0.0013313
10 | NO | 100.246 | 32.6593 | 0.0029432 | -0.0076599 | 0.00024417 | 0.0079041 |  | 749.532 | -0.00086212
10 | NO | 73.8814 | 32.2736 | 0.0018067 | 0.0067787 | 0.00024158 | -0.0065371 |  | 401.101 | 0.001339
10 | NO | 31.0363 | 32.1922 | 0.001607 | 0.0013428 | 0.00023736 | -0.0011054 |  | 300.564 | -0.0005455
10 | NO | 20.3233 | 32.0881 | 0.00091529 | -0.0017166 | 0.00023631 | 0.0019529 |  | 250.566 | 0.00063705
10 | YES | 10.0765 | 32.0889 | 0.00026002 | -5.722e-05 | 0.0002353 | 0.00029252 |  | 200.411 | -0.00022125
10 | YES | 4.3306 | 32.0891 | 0.0002128 | 0.0021477 | 0.00023474 | -0.0019129 |  | 173.938 | 0.00099182
11 | YES | 750.049 | 34.2678 | 0.0002984 | 0.0020638 | 0.00030812 | -0.0017556 |  | 126.065 | -0.00020218
11 | NO | 4.7741 | 31.8719 | 0.0054338 | -0.002203 | 0.00023478 | 0.0024378 |  | 1251.08 | -0.00076675
12 | YES | 2001.89 | 34.591 | 0.0002338 | -0.00041199 | 0.0004313 | 0.00084328 |  | 998.656 | -0.00045776
12 | YES | 1500.38 | 34.5068 | 0.00025865 | -0.0010414 | 0.00038195 | 0.0014234 |  | 149.739 | 0.0011368
12 | YES | 1249.83 | 34.4549 | 0.00027924 | 0.0002861 | 0.00035729 | 7.1192e-05 |  | 1999.77 | 0.00029373
12 | YES | 1000.36 | 34.3707 | 0.00035115 | 0.0013313 | 0.00033275 | -0.00099858 |  | 1500.44 | 0.0030594
12 | YES | 749.532 | 34.255 | 0.00025497 | -0.00086212 | 0.00030806 | 0.0011702 |  | 1251.4 | -0.00016022
12 | NO | 500.829 | 34.0975 | 0.0012083 | -0.0029373 | 0.00028359 | 0.0032209 |  | 1001.35 | 0.0003128
12 | YES | 401.101 | 33.9585 | 0.00023118 | 0.001339 | 0.00027378 | -0.0010652 |  | 749.295 | 0.0013924
12 | YES | 300.564 | 33.9694 | 0.00036272 | -0.0005455 | 0.00026389 | 0.00080939 |  | 500.806 | -0.00029373
12 | YES | 250.566 | 33.9108 | 0.0002499 | 0.00063705 | 0.00025897 | -0.00037809 |  | 400.241 | 0.00052261
12 | YES | 200.411 | 33.8677 | 0.00068013 | -0.00022125 | 0.00025403 | 0.00047528 |  | 301.352 | 0.00090027
12 | YES | 173.938 | 33.8289 | 0.00071323 | 0.00099182 | 0.00025143 | -0.0007404 |  | 250.198 | 0.0019913
12 | NO | 151.035 | 33.7485 | 0.0015023 | 0.002655 | 0.00024917 | -0.0024059 |  | 199.789 | -0.0044479
12 | YES | 126.065 | 33.5728 | 0.00071287 | -0.00020218 | 0.00024671 | 0.00044889 |  | 175.407 | -0.00099564
12 | NO | 99.773 | 33.3873 | 0.0027723 | 0.008976 | 0.00024413 | -0.0087319 |  | 150.409 | 0.0017776
12 | NO | 74.9852 | 32.7261 | 0.0053973 | 0.01712 | 0.00024169 | -0.016879 |  | 125.329 | -0.0015221
12 | NO | 50.7786 | 32.3655 | 0.0022926 | -0.0052109 | 0.00023931 | 0.0054502 |  | 999.275 | -0.00072479
12 | NO | 30.9731 | 32.2143 | 0.0021059 | 0.010963 | 0.00023736 | -0.010726 |  | 749.334 | 0.00021362
12 | NO | 20.6644 | 32.0143 | 0.020544 | 0.047167 | 0.00023634 | -0.04693 |  | 499.328 | -0.00042725
12 | NO | 10.2557 | 31.8119 | 0.0032575 | 0.01479 | 0.00023532 | -0.014554 |  | 4.6359 | -0.00058937
12 | NO | 5.8917 | 31.7367 | 0.0032691 | 0.008028 | 0.00023489 | -0.0077931 |  | 299.996 | 0.0016518
13 | YES | 1251.08 | 34.4405 | 0.00030572 | -0.00076675 | 0.00035742 | 0.0011242 |  | 349.999 | 0.0033455
13 | YES | 998.656 | 34.3575 | 0.00037293 | -0.00045776 | 0.00033258 | 0.00079034
14 | YES | 149.739 | 33.9114 | 0.00039068 | 0.0011368 | 0.00024904 | -0.00088774
15 | YES | 1999.77 | 34.5795 | 0.00028028 | 0.00029373 | 0.00043109 | 0.00013736
15 | YES | 1500.44 | 34.4987 | 0.00022971 | 0.0030594 | 0.00038195 | -0.0026774
15 | YES | 1251.4 | 34.4511 | 0.00024371 | -0.00016022 | 0.00035745 | 0.00051767 |  |  |  | 0.00029296 | 0.0014187
15 | YES | 1001.35 | 34.3745 | 0.0004914 | 0.0003128 | 0.00033284 | 2.0039e-05
15 | YES | 749.295 | 34.278 | 0.00045187 | 0.0013924 | 0.00030804 | -0.0010843
15 | YES | 500.806 | 34.1068 | 0.00036409 | -0.00029373 | 0.00028359 | 0.00057732
15 | YES | 400.241 | 34.015 | 0.00038978 | 0.00052261 | 0.00027369 | -0.00024892
15 | YES | 301.352 | 33.9645 | 0.00021672 | 0.00090027 | 0.00026396 | -0.00063631
15 | YES | 250.198 | 33.9154 | 0.00055711 | 0.0019913 | 0.00025893 | -0.0017323
15 | YES | 199.789 | 33.8542 | 0.0006487 | -0.0044479 | 0.00025397 | 0.0047019
15 | YES | 175.407 | 33.8108 | 0.00029941 | -0.00099564 | 0.00025157 | 0.0012472
15 | YES | 150.409 | 33.7854 | 0.00027093 | 0.0017776 | 0.00024911 | -0.0015285
15 | YES | 125.329 | 33.4825 | 0.00058621 | -0.0015221 | 0.00024664 | 0.0017687
15 | NO | 100.306 | 33.0891 | 0.0025846 | 0.01585 | 0.00024418 | -0.015606
15 | NO | 74.6416 | 32.4342 | 0.0058505 | 0.005146 | 0.00024165 | -0.0049044
15 | NO | 50.6408 | 32.3576 | 0.00091403 | -0.00037003 | 0.00023929 | 0.00060932
15 | NO | 30.2692 | 32.2681 | 0.0064859 | 0.017124 | 0.00023729 | -0.016887
15 | NO | 20.3108 | 32.1452 | 0.0045978 | 0.010506 | 0.00023631 | -0.010269
15 | NO | 10.2471 | 31.5111 | 0.0089727 | 0.029423 | 0.00023532 | -0.029187
15 | NO | 5.2217 | 31.3777 | 0.0069767 | 0.0082607 | 0.00023482 | -0.0080259
16 | YES | 999.275 | 34.351 | 0.00037777 | -0.00072479 | 0.00033264 | 0.0010574
16 | YES | 749.334 | 34.2306 | 0.00029133 | 0.00021362 | 0.00030805 | 9.4422e-05
16 | YES | 499.328 | 34.081 | 0.00024391 | -0.00042725 | 0.00028344 | 0.00071069
16 | YES | 4.6359 | 31.5865 | 0.00021485 | -0.00058937 | 0.00023477 | 0.00082414
17 | NO | 5.5254 | 30.7383 | 0.064462 | 0.038729 | 0.00023485 | -0.038494
18 | NO | 4.8985 | 30.5948 | 0.0064746 | 0.0039139 | 0.00023479 | -0.0036791
19 | YES | 299.996 | 30.6829 | 0.00024205 | 0.0016518 | 0.00026383 | -0.0013879
19 | NO | 30.0203 | 28.9807 | 0.00054577 | -0.0099144 | 0.00023726 | 0.010152
20 | YES | 349.999 | 31.0885 | 0.00020913 | 0.0033455 | 0.00026875 | -0.0030767
20 | NO | 9.5543 | 28.2843 | 0.0052249 | -0.006546 | 0.00023525 | 0.0067813
21 | NO | 310.386 | 31.0688 | 0.00049377 | 0.021595 | 0.00026485 | -0.02133
21 | NO | 9.7356 | 28.2435 | 0.00048149 | -0.041771 | 0.00023527 | 0.042006
0.0024936
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.063637 |  | X | CALC Y
m= | 0.00010326 | 5.4927e-05 |  | Std Dev of Resid |  | 0.0015696 |  | 3 | 0.00053174
b= | 0.00022197 | 0.00067354 |  | Resid Sum of Squares |  | 0.00012811 |  | 20 | 0.0022871
# of Values Extracted = |  |  | 91
# of Values Used in Fit = |  |  | 54
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0917 | 29.5001 | 0.031019 | 0.063252 | 0.00032523 | -0.062926 |  | 3 | -0.0032902
2 | NO | 4.9106 | 31.2704 | 0.0020751 | -0.0022774 | 0.00042849 | 0.0027059 |  | 4 | 0.00074768
3 | YES | 4.6322 | 31.9613 | 0.00049542 | -0.0032902 | 0.00053174 | 0.0038219 |  | 4 | 0.0027084
4 | YES | 1001.34 | 34.3824 | 0.00024107 | 0.00074768 | 0.000635 | -0.00011268 |  | 5 | 0.0020256
4 | NO | 750.539 | 34.252 | 0.0012985 | -0.075558 | 0.000635 | 0.076193 |  | 5 | 0.0011558
4 | YES | 500.196 | 34.0801 | 0.0002803 | 0.0027084 | 0.000635 | -0.0020734 |  | 5 | 0.0026856
5 | NO | 1250.1 | 34.4658 | 0.00038127 | 0.0044708 | 0.00073826 | -0.0037326 |  | 5 | 0.00033951
5 | YES | 1000.22 | 34.3844 | 0.00027006 | 0.0020256 | 0.00073826 | -0.0012873 |  | 6 | -0.00078201
5 | YES | 750.029 | 34.2687 | 0.00025506 | 0.0011558 | 0.00073826 | -0.00041759 |  | 8 | -0.001339
5 | YES | 500.468 | 34.0684 | 0.00026246 | 0.0026856 | 0.00073826 | -0.0019473 |  | 9 | -0.00029373
5 | YES | 4.7572 | 32.0134 | 0.00032488 | 0.00033951 | 0.00073826 | 0.00039875 |  | 9 | 0.0022392
6 | YES | 4.8455 | 31.5451 | 0.00025207 | -0.00078201 | 0.00084151 | 0.0016235 |  | 10 | 0.0025291
7 | NO | 5.1087 | 31.1432 | 0.0043529 | -0.0071297 | 0.00094477 | 0.0080744 |  | 10 | 0.00070953
8 | YES | 5.8663 | 32.0096 | 0.0003114 | -0.001339 | 0.001048 | 0.002387 |  | 10 | 9.9182e-05
9 | YES | 750.385 | 34.2487 | 0.00028608 | -0.00029373 | 0.0011513 | 0.001445 |  | 10 | 0.0023193
9 | YES | 500.727 | 34.1014 | 0.00026083 | 0.0022392 | 0.0011513 | -0.0010879 |  | 10 | 0.00032425
10 | YES | 1499.68 | 34.5047 | 0.00030098 | 0.0025291 | 0.0012545 | -0.0012746 |  | 10 | 0.0037155
10 | YES | 1249.65 | 34.4684 | 0.00027987 | 0.00070953 | 0.0012545 | 0.00054501 |  | 10 | 0.0019951
10 | YES | 999.454 | 34.3729 | 0.00025105 | 9.9182e-05 | 0.0012545 | 0.0011554 |  | 10 | 0.0023994
10 | YES | 749.742 | 34.2775 | 0.00025949 | 0.0023193 | 0.0012545 | -0.0010648 |  | 10 | 0.0036507
10 | YES | 499.387 | 34.0945 | 0.00045036 | 0.00032425 | 0.0012545 | 0.00093029 |  | 10 | -0.00041962
10 | YES | 400.824 | 33.9848 | 0.00059672 | 0.0037155 | 0.0012545 | -0.002461 |  | 10 | 0.00088119
10 | YES | 299.619 | 33.9131 | 0.00052061 | 0.0019951 | 0.0012545 | -0.00074055 |  | 10 | 0.0029373
10 | YES | 250.525 | 33.9019 | 0.00057706 | 0.0023994 | 0.0012545 | -0.0011449 |  | 11 | 0.0035286
10 | NO | 200.65 | 33.8554 | 0.00087212 | 0.0028496 | 0.0012545 | -0.001595 |  | 12 | 0.00045013
10 | YES | 175.127 | 33.795 | 0.00045558 | 0.0036507 | 0.0012545 | -0.0023961 |  | 12 | 3.0518e-05
10 | NO | 150.311 | 33.7011 | 0.00094209 | 0.0044517 | 0.0012545 | -0.0031972 |  | 12 | 0.0015335
10 | NO | 125.603 | 33.361 | 0.0042175 | 0.012859 | 0.0012545 | -0.011605 |  | 12 | 0.0026169
10 | NO | 100.246 | 32.6645 | 0.0031128 | -0.0025063 | 0.0012545 | 0.0037608 |  | 12 | 0.0006218
10 | NO | 73.8814 | 32.2745 | 0.0022824 | 0.0076904 | 0.0012545 | -0.0064359 |  | 12 | 0.0027504
10 | NO | 31.0363 | 32.193 | 0.0026426 | 0.0020905 | 0.0012545 | -0.00083591 |  | 12 | 0.00082398
10 | YES | 20.3233 | 32.0894 | 0.00065749 | -0.00041962 | 0.0012545 | 0.0016742 |  | 12 | 0.0020142
10 | YES | 10.0765 | 32.0899 | 0.00024317 | 0.00088119 | 0.0012545 | 0.00037335 |  | 13 | 0.00030518
10 | YES | 4.3306 | 32.0899 | 0.00021723 | 0.0029373 | 0.0012545 | -0.0016828 |  | 13 | 0.00069427
11 | YES | 750.049 | 34.2692 | 0.00034323 | 0.0035286 | 0.0013578 | -0.0021708 |  | 14 | 0.0026398
11 | NO | 4.7741 | 31.8727 | 0.004914 | -0.0014191 | 0.0013578 | 0.0027769 |  | 15 | 0.0013962
12 | YES | 2001.89 | 34.5919 | 0.00023599 | 0.00045013 | 0.0014611 | 0.0010109 |  | 15 | 0.0042686
12 | YES | 1500.38 | 34.5078 | 0.00029674 | 3.0518e-05 | 0.0014611 | 0.0014305 |  | 15 | 0.0011063
12 | YES | 1249.83 | 34.4561 | 0.00031667 | 0.0015335 | 0.0014611 | -7.2455e-05 |  | 15 | 0.0016708
12 | YES | 1000.36 | 34.372 | 0.0003751 | 0.0026169 | 0.0014611 | -0.0011558 |  | 15 | 0.002758
12 | YES | 749.532 | 34.2565 | 0.00025029 | 0.0006218 | 0.0014611 | 0.00083926 |  | 15 | 0.0011215
12 | NO | 500.829 | 34.0989 | 0.0012261 | -0.0014496 | 0.0014611 | 0.0029106 |  | 15 | 0.0019951
12 | YES | 401.101 | 33.9599 | 0.0002548 | 0.0027504 | 0.0014611 | -0.0012893 |  | 15 | 0.0023537
12 | YES | 300.564 | 33.9707 | 0.00040588 | 0.00082398 | 0.0014611 | 0.00063708 |  | 15 | 0.0034027
12 | YES | 250.566 | 33.9122 | 0.00023214 | 0.0020142 | 0.0014611 | -0.0005531 |  | 15 | -0.0029678
12 | NO | 200.411 | 33.8689 | 0.00083429 | 0.001049 | 0.0014611 | 0.00041202 |  | 15 | 0.00044632
12 | NO | 173.938 | 33.8303 | 0.00083463 | 0.0023918 | 0.0014611 | -0.00093076 |  | 15 | 0.0030594
12 | NO | 151.035 | 33.7495 | 0.0012737 | 0.003727 | 0.0014611 | -0.0022659 |  | 15 | -0.00028992
12 | NO | 126.065 | 33.5738 | 0.00085326 | 0.0008049 | 0.0014611 | 0.00065615 |  | 16 | 0.00064087
12 | NO | 99.773 | 33.3858 | 0.0029248 | 0.007534 | 0.0014611 | -0.006073 |  | 16 | 0.0016365
12 | NO | 74.9852 | 32.7255 | 0.0043044 | 0.016499 | 0.0014611 | -0.015038 |  | 16 | 0.00096893
12 | NO | 50.7786 | 32.3668 | 0.0013768 | -0.0039291 | 0.0014611 | 0.0053902 |  | 16 | 0.00041962
12 | NO | 30.9731 | 32.2143 | 0.0025061 | 0.01104 | 0.0014611 | -0.0095786 |  | 19 | 0.002943
12 | NO | 20.6644 | 32.016 | 0.014803 | 0.048864 | 0.0014611 | -0.047403 |  | 20 | 0.0045547
12 | NO | 10.2557 | 31.813 | 0.0032499 | 0.015911 | 0.0014611 | -0.01445
12 | NO | 5.8917 | 31.7377 | 0.0036593 | 0.0090332 | 0.0014611 | -0.0075721
13 | YES | 1251.08 | 34.4416 | 0.00032376 | 0.00030518 | 0.0015643 | 0.0012591
13 | YES | 998.656 | 34.3587 | 0.00033116 | 0.00069427 | 0.0015643 | 0.00087004
14 | YES | 149.739 | 33.9129 | 0.0003768 | 0.0026398 | 0.0016676 | -0.0009722
15 | YES | 1999.77 | 34.5806 | 0.00030444 | 0.0013962 | 0.0017708 | 0.00037465
15 | YES | 1500.44 | 34.4999 | 0.00025156 | 0.0042686 | 0.0017708 | -0.0024978
15 | YES | 1251.4 | 34.4524 | 0.00023288 | 0.0011063 | 0.0017708 | 0.00066457
15 | YES | 1001.35 | 34.3759 | 0.00047311 | 0.0016708 | 0.0017708 | 9.9986e-05
15 | YES | 749.295 | 34.2794 | 0.0004486 | 0.002758 | 0.0017708 | -0.0009872
15 | YES | 500.806 | 34.1082 | 0.00040348 | 0.0011215 | 0.0017708 | 0.00064931
15 | YES | 400.241 | 34.0165 | 0.00047953 | 0.0019951 | 0.0017708 | -0.00022426
15 | YES | 301.352 | 33.966 | 0.00021884 | 0.0023537 | 0.0017708 | -0.00058284
15 | YES | 250.198 | 33.9168 | 0.00064921 | 0.0034027 | 0.0017708 | -0.0016319
15 | YES | 199.789 | 33.8556 | 0.00055152 | -0.0029678 | 0.0017708 | 0.0047387
15 | YES | 175.407 | 33.8122 | 0.00028533 | 0.00044632 | 0.0017708 | 0.0013245
15 | YES | 150.409 | 33.7867 | 0.00027021 | 0.0030594 | 0.0017708 | -0.0012886
15 | YES | 125.329 | 33.4837 | 0.00050711 | -0.00028992 | 0.0017708 | 0.0020607
15 | NO | 100.306 | 33.093 | 0.0029244 | 0.019814 | 0.0017708 | -0.018043
15 | NO | 74.6416 | 32.4332 | 0.004124 | 0.0040703 | 0.0017708 | -0.0022995
15 | NO | 50.6408 | 32.3589 | 0.00081204 | 0.00090027 | 0.0017708 | 0.00087056
15 | NO | 30.2692 | 32.2694 | 0.006311 | 0.018356 | 0.0017708 | -0.016585
15 | NO | 20.3108 | 32.1467 | 0.0035988 | 0.011955 | 0.0017708 | -0.010184
15 | NO | 10.2471 | 31.5093 | 0.0096319 | 0.027588 | 0.0017708 | -0.025817
15 | NO | 5.2217 | 31.3807 | 0.0068935 | 0.011328 | 0.0017708 | -0.0095569
16 | YES | 999.275 | 34.3523 | 0.00032353 | 0.00064087 | 0.0018741 | 0.0012332
16 | YES | 749.334 | 34.232 | 0.0005002 | 0.0016365 | 0.0018741 | 0.00023757
16 | YES | 499.328 | 34.0824 | 0.00028319 | 0.00096893 | 0.0018741 | 0.00090515
16 | YES | 4.6359 | 31.5875 | 0.00021633 | 0.00041962 | 0.0018741 | 0.0014545
17 | NO | 5.5254 | 30.7481 | 0.068869 | 0.048506 | 0.0019773 | -0.046528
18 | NO | 4.8985 | 30.5947 | 0.0061611 | 0.0037804 | 0.0020806 | -0.0016998
19 | YES | 299.996 | 30.6841 | 0.00025665 | 0.002943 | 0.0021839 | -0.00075919
19 | NO | 30.0203 | 28.9821 | 0.00082899 | -0.008503 | 0.0021839 | 0.010687
20 | YES | 349.999 | 31.0898 | 0.00024109 | 0.0045547 | 0.0022871 | -0.0022676
20 | NO | 9.5543 | 28.2832 | 0.0060495 | -0.007576 | 0.0022871 | 0.0098631
21 | NO | 310.386 | 31.07 | 0.00039056 | 0.022829 | 0.0023904 | -0.020439
21 | NO | 9.7356 | 28.2449 | 0.0011965 | -0.040424 | 0.0023904 | 0.042815
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.016557 |  | X | CALC Y
m= | 4.7794e-05 | 4.9668e-05 |  | Std Dev of Resid |  | 0.0014197 |  | 3 | -0.00012042
b= | -0.0002638 | 0.00060838 |  | Resid Sum of Squares |  | 0.00011085 |  | 20 | 0.00069209
# of Values Extracted = |  |  | 91
# of Values Used in Fit = |  |  | 57
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.0917 | 29.4956 | 0.024507 | 0.058771 | -0.00021601 | -0.058987 |  | 3 | -0.0023613
2 | NO | 4.9106 | 31.2697 | 0.0019146 | -0.0029583 | -0.00016821 | 0.0027901 |  | 4 | -0.00019074
3 | YES | 4.6322 | 31.9622 | 0.00052184 | -0.0023613 | -0.00012042 | 0.0022409 |  | 4 | 0.0015564
4 | YES | 1001.34 | 34.3815 | 0.00024839 | -0.00019074 | -7.2622e-05 | 0.00011811 |  | 5 | 0.00090408
4 | NO | 750.539 | 34.2509 | 0.0011897 | -0.076721 | -7.2622e-05 | 0.076649 |  | 5 | 1.1444e-05
4 | YES | 500.196 | 34.079 | 0.00033108 | 0.0015564 | -7.2622e-05 | -0.001629 |  | 5 | 0.0015373
5 | NO | 1250.1 | 34.465 | 0.00035976 | 0.0037193 | -2.4828e-05 | -0.0037442 |  | 5 | -0.00022888
5 | YES | 1000.22 | 34.3833 | 0.00023675 | 0.00090408 | -2.4828e-05 | -0.00092891 |  | 6 | -0.0014572
5 | YES | 750.029 | 34.2675 | 0.00023431 | 1.1444e-05 | -2.4828e-05 | -3.6272e-05 |  | 8 | -0.0021095
5 | YES | 500.468 | 34.0672 | 0.0002543 | 0.0015373 | -2.4828e-05 | -0.0015621 |  | 9 | -0.0015183
5 | YES | 4.7572 | 32.0129 | 0.00034754 | -0.00022888 | -2.4828e-05 | 0.00020405 |  | 9 | 0.0010414
6 | YES | 4.8455 | 31.5444 | 0.00028812 | -0.0014572 | 2.2966e-05 | 0.0014802 |  | 10 | 0.0015411
7 | NO | 5.1087 | 31.1425 | 0.0042406 | -0.0078068 | 7.0761e-05 | 0.0078775 |  | 10 | -0.00034714
8 | YES | 5.8663 | 32.0088 | 0.00041243 | -0.0021095 | 0.00011855 | 0.0022281 |  | 10 | -0.00103
9 | YES | 750.385 | 34.2475 | 0.00025185 | -0.0015183 | 0.00016635 | 0.0016846 |  | 10 | 0.00099945
9 | YES | 500.727 | 34.1002 | 0.00025942 | 0.0010414 | 0.00016635 | -0.00087506 |  | 10 | -0.0010223
10 | YES | 1499.68 | 34.5037 | 0.00028366 | 0.0015411 | 0.00021414 | -0.001327 |  | 10 | 0.0021591
10 | YES | 1249.65 | 34.4674 | 0.00023759 | -0.00034714 | 0.00021414 | 0.00056128 |  | 10 | 0.00067902
10 | YES | 999.454 | 34.3718 | 0.00024132 | -0.00103 | 0.00021414 | 0.0012441 |  | 10 | 0.00074005
10 | YES | 749.742 | 34.2762 | 0.00026583 | 0.00099945 | 0.00021414 | -0.00078531 |  | 10 | 0.0016861
10 | YES | 499.387 | 34.0932 | 0.00041702 | -0.0010223 | 0.00021414 | 0.0012365 |  | 10 | 0.0024414
10 | YES | 400.824 | 33.9833 | 0.00066962 | 0.0021591 | 0.00021414 | -0.001945 |  | 10 | -5.722e-05
10 | YES | 299.619 | 33.9118 | 0.00038091 | 0.00067902 | 0.00021414 | -0.00046487 |  | 10 | 0.0021477
10 | YES | 250.525 | 33.9002 | 0.00048614 | 0.00074005 | 0.00021414 | -0.00052591 |  | 11 | 0.0020638
10 | YES | 200.65 | 33.8542 | 0.00079651 | 0.0016861 | 0.00021414 | -0.001472 |  | 12 | -0.00041199
10 | YES | 175.127 | 33.7937 | 0.00065585 | 0.0024414 | 0.00021414 | -0.0022273 |  | 12 | -0.0010414
10 | NO | 150.311 | 33.6997 | 0.00098713 | 0.003128 | 0.00021414 | -0.0029139 |  | 12 | 0.0002861
10 | NO | 125.603 | 33.3609 | 0.0040029 | 0.012814 | 0.00021414 | -0.012599 |  | 12 | 0.0013313
10 | NO | 100.246 | 32.6593 | 0.0029432 | -0.0076599 | 0.00021414 | 0.0078741 |  | 12 | -0.00086212
10 | NO | 73.8814 | 32.2736 | 0.0018067 | 0.0067787 | 0.00021414 | -0.0065646 |  | 12 | 0.001339
10 | NO | 31.0363 | 32.1922 | 0.001607 | 0.0013428 | 0.00021414 | -0.0011286 |  | 12 | -0.0005455
10 | NO | 20.3233 | 32.0881 | 0.00091529 | -0.0017166 | 0.00021414 | 0.0019308 |  | 12 | 0.00063705
10 | YES | 10.0765 | 32.0889 | 0.00026002 | -5.722e-05 | 0.00021414 | 0.00027136 |  | 12 | -0.00022125
10 | YES | 4.3306 | 32.0891 | 0.0002128 | 0.0021477 | 0.00021414 | -0.0019335 |  | 12 | 0.00099182
11 | YES | 750.049 | 34.2678 | 0.0002984 | 0.0020638 | 0.00026194 | -0.0018018 |  | 12 | -0.00020218
11 | NO | 4.7741 | 31.8719 | 0.0054338 | -0.002203 | 0.00026194 | 0.0024649 |  | 13 | -0.00076675
12 | YES | 2001.89 | 34.591 | 0.0002338 | -0.00041199 | 0.00030973 | 0.00072172 |  | 13 | -0.00045776
12 | YES | 1500.38 | 34.5068 | 0.00025865 | -0.0010414 | 0.00030973 | 0.0013511 |  | 14 | 0.0011368
12 | YES | 1249.83 | 34.4549 | 0.00027924 | 0.0002861 | 0.00030973 | 2.363e-05 |  | 15 | 0.00029373
12 | YES | 1000.36 | 34.3707 | 0.00035115 | 0.0013313 | 0.00030973 | -0.0010216 |  | 15 | 0.0030594
12 | YES | 749.532 | 34.255 | 0.00025497 | -0.00086212 | 0.00030973 | 0.0011719 |  | 15 | -0.00016022
12 | NO | 500.829 | 34.0975 | 0.0012083 | -0.0029373 | 0.00030973 | 0.0032471 |  | 15 | 0.0003128
12 | YES | 401.101 | 33.9585 | 0.00023118 | 0.001339 | 0.00030973 | -0.0010292 |  | 15 | 0.0013924
12 | YES | 300.564 | 33.9694 | 0.00036272 | -0.0005455 | 0.00030973 | 0.00085523 |  | 15 | -0.00029373
12 | YES | 250.566 | 33.9108 | 0.0002499 | 0.00063705 | 0.00030973 | -0.00032732 |  | 15 | 0.00052261
12 | YES | 200.411 | 33.8677 | 0.00068013 | -0.00022125 | 0.00030973 | 0.00053098 |  | 15 | 0.00090027
12 | YES | 173.938 | 33.8289 | 0.00071323 | 0.00099182 | 0.00030973 | -0.00068209 |  | 15 | 0.0019913
12 | NO | 151.035 | 33.7485 | 0.0015023 | 0.002655 | 0.00030973 | -0.0023453 |  | 15 | -0.0044479
12 | YES | 126.065 | 33.5728 | 0.00071287 | -0.00020218 | 0.00030973 | 0.00051191 |  | 15 | -0.00099564
12 | NO | 99.773 | 33.3873 | 0.0027723 | 0.008976 | 0.00030973 | -0.0086662 |  | 15 | 0.0017776
12 | NO | 74.9852 | 32.7261 | 0.0053973 | 0.01712 | 0.00030973 | -0.016811 |  | 15 | -0.0015221
12 | NO | 50.7786 | 32.3655 | 0.0022926 | -0.0052109 | 0.00030973 | 0.0055206 |  | 16 | -0.00072479
12 | NO | 30.9731 | 32.2143 | 0.0021059 | 0.010963 | 0.00030973 | -0.010654 |  | 16 | 0.00021362
12 | NO | 20.6644 | 32.0143 | 0.020544 | 0.047167 | 0.00030973 | -0.046857 |  | 16 | -0.00042725
12 | NO | 10.2557 | 31.8119 | 0.0032575 | 0.01479 | 0.00030973 | -0.01448 |  | 16 | -0.00058937
12 | NO | 5.8917 | 31.7367 | 0.0032691 | 0.008028 | 0.00030973 | -0.0077183 |  | 19 | 0.0016518
13 | YES | 1251.08 | 34.4405 | 0.00030572 | -0.00076675 | 0.00035753 | 0.0011243 |  | 20 | 0.0033455
13 | YES | 998.656 | 34.3575 | 0.00037293 | -0.00045776 | 0.00035753 | 0.00081529
14 | YES | 149.739 | 33.9114 | 0.00039068 | 0.0011368 | 0.00040532 | -0.00073146
15 | YES | 1999.77 | 34.5795 | 0.00028028 | 0.00029373 | 0.00045311 | 0.00015938
15 | YES | 1500.44 | 34.4987 | 0.00022971 | 0.0030594 | 0.00045311 | -0.0026063
15 | YES | 1251.4 | 34.4511 | 0.00024371 | -0.00016022 | 0.00045311 | 0.00061333
15 | YES | 1001.35 | 34.3745 | 0.0004914 | 0.0003128 | 0.00045311 | 0.00014031
15 | YES | 749.295 | 34.278 | 0.00045187 | 0.0013924 | 0.00045311 | -0.00093925
15 | YES | 500.806 | 34.1068 | 0.00036409 | -0.00029373 | 0.00045311 | 0.00074685
15 | YES | 400.241 | 34.015 | 0.00038978 | 0.00052261 | 0.00045311 | -6.9499e-05
15 | YES | 301.352 | 33.9645 | 0.00021672 | 0.00090027 | 0.00045311 | -0.00044715
15 | YES | 250.198 | 33.9154 | 0.00055711 | 0.0019913 | 0.00045311 | -0.0015382
15 | YES | 199.789 | 33.8542 | 0.0006487 | -0.0044479 | 0.00045311 | 0.0049011
15 | YES | 175.407 | 33.8108 | 0.00029941 | -0.00099564 | 0.00045311 | 0.0014488
15 | YES | 150.409 | 33.7854 | 0.00027093 | 0.0017776 | 0.00045311 | -0.0013245
15 | YES | 125.329 | 33.4825 | 0.00058621 | -0.0015221 | 0.00045311 | 0.0019752
15 | NO | 100.306 | 33.0891 | 0.0025846 | 0.01585 | 0.00045311 | -0.015397
15 | NO | 74.6416 | 32.4342 | 0.0058505 | 0.005146 | 0.00045311 | -0.0046929
15 | NO | 50.6408 | 32.3576 | 0.00091403 | -0.00037003 | 0.00045311 | 0.00082314
15 | NO | 30.2692 | 32.2681 | 0.0064859 | 0.017124 | 0.00045311 | -0.016671
15 | NO | 20.3108 | 32.1452 | 0.0045978 | 0.010506 | 0.00045311 | -0.010053
15 | NO | 10.2471 | 31.5111 | 0.0089727 | 0.029423 | 0.00045311 | -0.02897
15 | NO | 5.2217 | 31.3777 | 0.0069767 | 0.0082607 | 0.00045311 | -0.0078076
16 | YES | 999.275 | 34.351 | 0.00037777 | -0.00072479 | 0.00050091 | 0.0012257
16 | YES | 749.334 | 34.2306 | 0.00029133 | 0.00021362 | 0.00050091 | 0.00028729
16 | YES | 499.328 | 34.081 | 0.00024391 | -0.00042725 | 0.00050091 | 0.00092816
16 | YES | 4.6359 | 31.5865 | 0.00021485 | -0.00058937 | 0.00050091 | 0.0010903
17 | NO | 5.5254 | 30.7383 | 0.064462 | 0.038729 | 0.0005487 | -0.03818
18 | NO | 4.8985 | 30.5948 | 0.0064746 | 0.0039139 | 0.0005965 | -0.0033174
19 | YES | 299.996 | 30.6829 | 0.00024205 | 0.0016518 | 0.00064429 | -0.0010075
19 | NO | 30.0203 | 28.9807 | 0.00054577 | -0.0099144 | 0.00064429 | 0.010559
20 | YES | 349.999 | 31.0885 | 0.00020913 | 0.0033455 | 0.00069209 | -0.0026534
20 | NO | 9.5543 | 28.2843 | 0.0052249 | -0.006546 | 0.00069209 | 0.0072381
21 | NO | 310.386 | 31.0688 | 0.00049377 | 0.021595 | 0.00073988 | -0.020855
21 | NO | 9.7356 | 28.2435 | 0.00048149 | -0.041771 | 0.00073988 | 0.042511
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE