## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2022-035

Agency: Ecosystem Science Division (OSD)

Location: Saanich Inlet and North-East Pacific

Project:  Seamounts

Party Chief: Norgard T.

Platform: John P. Tully

Date: 7 June 2022 – 28 June 2022 (at sea 16 June – 27 June 2022)

Processed by: Germaine Gatien

Date of Processing: 12 August 2022 –  19 August 2022

Number of original HEX files: 8

Number of CTD files:  8	Number of CTD casts processed: 7 (1 file full of pressure spikes)

Number of bottle files: 7	Number of bottle casts processed: 6 (1 bottle cast was a test only)

Number of TSG files:   1	Number of TSG files processed: 0 (no date/time channels)

# INSTRUMENT SUMMARY

An SBE911+ CTD #0506 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometers (#1185DR and #1883DG), a SBE 43 DO sensor (#997), a Wet Labs ECO-AFL Fluorometer (#3982) and an altimeter (#76341).

Seasave version 7.26.7.121 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlabs WETStar fluorometer (#1656) and flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format and there was no paper log. The digital log was missing information about which CTD was in use and what sensors were mounted on it. There was a personnel list and useful comments about problems encountered.

The 2 CTD salinity channels differed by about 0.009psu, so it is fortunate that the salinity sampling was excellent even though there were few CTD casts. The primary salinity was found to be very close to bottles. An explanation was found after the cruise for the poor performance of the secondary conductivity sensor.

When the pressure sensor was used on 2 cruises in 2021 on the Franklin, there was some evidence that pressure might be reading high by about 1.5db and a lab test at IOS gave a similar result. However, there was a good deal of evidence that pressure was accurate during cruise 2021-078 based on observations when the CTD was in water and when it moved through the sea surface at the beginning and end of casts. A possible explanation may be that the high values were found when the CTD had been stored in a very warm lab, whereas it would have been outside on the Vector. There are no data from 2022-035 from close to the surface; the only negative pressure value recorded was from a 2-sided pressure spike at about 7db.

There were many spikes in pressure during CTD casts. Some affected records were removed in the routine processing steps. In some cases where pressure had been smoothed, spikes in salinity (and sometimes temperature) remained. Some further smoothing was done in editing. The worst effects were seen in the first 4 CTD casts, with event #6 so heavily corrupted that it was not suitable for archiving.

Event #23 included rosette firings, but the only sampling was for eDNA and Phyto; a CHE file was prepared that contains CTD data collected during bottle stops in case it is needed by those analyzing eDNA and Phyto samples.

Dissolved oxygen data generally look good, especially during bottle stops. While spikes are not obvious in the processed files, accuracy may have been compromised occasionally during downcasts. The offshore surface oxygen saturation looked a little high (105-107%), but near-surface SBE Dissolved oxygen values were slightly lower than the titrated surface samples. Surface oxygen saturation based on a titrated sample from cast #20 was 106%.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.40 mL/L from 0-100db except in areas of very large DO gradients

±0.30 mL/L from 100db-300db

±0.10 mL/L from 300db-1200db

±0.04 mL/L below 1200db

Thermosalinograph data were found but there were no date and time channels. There are two sections well separated in time and position: a 25-hour section between west of Banfield moving north to south of Haida Gwaii and a 1-hour section moving north in the Swiftsure Bank area. There were no CTD casts in those areas and no loop sampling, so even if position data are found there will be no way to judge the reliability of the data.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Salinity, Chlorophyll, Nutrient and Dissolved Oxygen data were obtained in QF spreadsheets.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained and checked and no errors were found

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2022-035-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format; the hysteresis correction was selected since there was deep sampling and the Tau correction was selected.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. None were found.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

All bottles were fired during cast #1 but no samples gathered, so it will not be processed further.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2022-035-bot-hdr.txt.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2022-035_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2022-035oxy.csv. That file was converted into individual *.OXY files.

There were no samples in the DO file that had comments starting with “ALL:”.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2022-035_CHL*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2022-035chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2022-035 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 21-24 days of collection. The files were simplified and saved as 2022-035sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF_NUTS_2022-035*.xlsx. This includes a precision study. The file was simplified, saved as 2022-035nuts.csv and converted to individual NUT files.

The file was then converted to individual files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. The only problems noted were that there was an error on one of the rosette sheets and one salinity sample was not found; likely planned but not collected. A pad value and flag 9 was added for that sample, #87 in event #27.

The minimum CTD Salinity in the spreadsheet was >31psu so there was no need to apply a correction to silicate samples.

A header check was run and no problems were found.

##### Compare

Salinity

This cruise had very few casts, but the salinity sampling was excellent since there was a complete profile from a deep cast.

When outliers were excluded based on standard deviation in the CTD salinity >0.001 and pressure <100db, the primary salinity was found to be low by 0.0002psu and the secondary low by 0.0093psu. This explains the large difference between sensors. See section 9.

The CTD technician found a blockage in the secondary conductivity cell that explains the poor performance. From this point on, only the primary salinity will be discussed.

Samples from above 200db are lower than bottles by about 0.004psu to 0.005psu except for one cast where the CTD salinity appears to be high by ~0.015psu. The standard deviations in the CTD salinity are high for all those samples. The 3 cases where the CTD appears to be low are likely due to incomplete flushing of Niskin bottles in the presence of high vertical salinity gradients so that bottle values represent conditions a little lower in the water column. The one case where the CTD was reading high was at the bottom of a cast where flushing errors have the opposite sign and CTD data tend to be very noisy.

There were 3 samples from the bottom. The 2 that have low standard deviations in the CTD data indicate that the primary salinity was low by 0.0005psu and high by 0.0002psu; standard deviations were 0.0007psu and 0.0008psu, respectively.

Since salinity was analyzed quickly, within 3 weeks, evaporation and desorption of samples should be small but not 0 and there could be a small effect from incomplete flushing. Both those errors would lead to bottle values being high. An estimate of 0.002psu seems reasonable for these errors, so the primary salinity may be reading slightly high. The standard deviation in the CTD salinity was 0.0007psu. There are also small errors in analysis but those tend to be random in sign. When last used the primary salinity was also low by -0.0002psu, but there were very few samples. No recalibration is justified.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The initial fit looked good except for 1 severe outlier. Sample #61 supposedly from 1500m is way out of line with the CTD DO and seems much higher than expected at 1500m. The files were missing that are necessary to re-examine the titration data, but the analyst felt that any errors that could be found in that data were unlikely to explain an error this large. It is assumed that there must have been some atmospheric exposure.

An initial fit based on plots of (SBE DO – Bottle DO) versus SBE DO after removing the 1 major outlier was:

CTD DO Corrected = CTD DO * 1.0391 + 0.0121 	 R2 = 0.95

A fit was based on plots of (SBE DO – Bottle DO) versus SBE DO after removing the 1 major outlier and a few others based on residuals was:

CTD DO Corrected = CTD DO * 1.0381 + 0.0140 	 R2 = 0.97

The salinity comparison suggests that flushing of bottles was incomplete in the top 100db and many of the values rejected come from those levels. Others come from areas of large DO gradient so even small flushing errors may explain these differences which are not particularly large. The rejected values mostly come from above the trendline as expected above the OMZ and below the trendline below the OMZ, as the effect of incomplete flushing has the opposite effect above and below the minimum.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further significant outliers were found.

Sample #61 was padded and flag 5 added with a comment.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. There were no values flagged as <MLD.

All extracted CHL values were low, <1.4ug/L. The usual pattern is seen with fluorescence reading higher than chlorophyll but approaching a ratio FL/CHL of 1 as CHL increased.

Since all values were low the Fluorescence vs CHL fit has a large offset:

FL = 0.92 *CHL + 0.61

When forced through the origin it has a large slope:

FL = 1.75 *CHL

##### Conversion of Full Files from Raw Data

All files were converted using 2022-035-ctd.xmlcon.

The Tau function was selected as well as the hysteresis function (there was deep sampling).

A few casts were examined and all expected channels are present.

- There is a severe problem in both conductivity and temperature channels in casts #6 down to 80db.  Cast #1 looks odd near the surface though the secondary channels may be ok. The primary conductivity has a very large spike around 33db. Cast #8 has many spikes in secondary conductivity though the overall profile looks reasonable – temperature has a few small spikes.

- As usual there was a lot of small-scale noise in the temperature and conductivity channels during upcasts.

- The Dissolved Oxygen looks reasonable except where there are temperature spikes. There are some reversals but they correspond to temperature reversals.

- The transmissivity traces look reasonable.

- Fluorescence had a dark value of ~0.06ug/L which is good.

- The altimetry looks excellent near the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A few casts were examined; both temperature channels were noisy during upcasts so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for both DO sensors. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and those parameters appear to have worked well.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between the channel pairs was examined:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2021-078-0038 | 300 | -0.0005 | -0.00068 | -0.0063 | High, Steady |
| 2021-078-0049 | 300 | -0.0005 | -0.00068 | -0.0063 | High, Steady |
| 2021-078-0071 | 350 | -0.0005 | -0.00064 | -0.0061 | High, Steady |
| 2021-078-0076 | 350 | -0.0005 | -0.00064 | -0.0061 | High, Steady |
| “ | 650 | -0.0005 | -0.00062 | -0.0059 | “ |
| 2022-035-0010 | 350 | -0.0002 | -0.00070 | -0.0078 | High, Moderate |
|  | 500 | 0 | -0.00071 | -0.0082 | “ |
|  | 1000 | 0 | -0.00069 | -0.0082 | “ |
| 2002-035-0023 | 350 | -0.0001 | -0.00077 | -0.0088 | High, Moderate |
|  | 1000 | -0.0001 | -0.00074 | -0.0090 | “ |
|  | 1940 | -0.0002 | -0.00071 | -0.0086 | “ |
| 2022-035-0027 | 350 | -0.0003 | -0.00080 | -0.0090 | High, F Steady |
|  | 1000 | -0.0001 | -0.00080 | -0.0094 | “ |
|  | 2000 | -0.0001 | -0.00077 | -0.0090 | “ |
|  | 2500 | -0.0003 | -0.00077 | -0.0088 | “ |

The temperature differences are very small, smaller then when last used and show little variation. Salinity and conductivity differences are quite high and appear to show some temporal drift, but no significant pressure dependence.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A header check was run. The fluorescence maximum value was close to the maximum, but was found to be caused by a spike associated with a pressure spike. CLEAN was rerun using the feature that will remove fluorescence values >14.94ug/L. Spikes may remain after this but any off-scale values will be padded. There are always small spikes in fluorescence, but plots will be made after running Delete to see if there are any large spikes.

A cross-reference list was checked and looks fine. A cruise track plot looks good. The list and track plots were added to the end of this report.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. We usually expect to get values close to 0 ± 5m. For this cruise only a few casts fell within that range. A second check was made using the water depths listed in the log and while the results varied, they were still far out of the expected range. This is presumably due to the nature of the areas being sampled with steep slopes. A reasonable estimate of water depth when the CTD was at the bottom of the cast was made by adding the maximum depth sampled plus the altimetry reading from the header. Since the depth would likely have been changing rapidly over seamounts, there seems little point in updating the values in the headers.

| Max Depth Sampled | Event # | Header Water Depth | Altimeter | Check1 |  | Log Water Depth | Check2 | Est water depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 49.2 | 1 | 50 | 5.1 | 4.3 |  | 50 | 4.3 | 54 |
| 164.2 | 6 | 175 | 5.3 | -5.6 |  | 165 | 4.4 | 169 |
| 1832.3 | 8 | 1885 | 5.8 | -46.8 |  | 1851 | -12.8 | 1838 |
| 995.8 | 10 | 974 | 4.0 | 25.8 |  | 995 | 4.8 | 1000 |
| 629.3 | 15 | 620 | 7.5 | 16.7 |  | 610 | 26.7 | 637 |
| 1814.0 | 20 | 1815 | 5.3 | 4.2 |  | 1836 | -16.8 | 1819 |
| 2032.3 | 23 | 2073 | 5.4 | -35.4 |  | 2059 | -21.4 | 2038 |
| 2923.0 | 27 | 2921 | 4.7 | 6.7 |  | 2968 | -40.3 | 2928 |

The pressure sensor used for this cruise was reported to read high on 2 previous cruises on the Franklin.

The CTD on the Franklin was stored in a very warm room which could possibly have affected the pressure sensor until it had time to equilibrate. Cruise 2021-078 on the Vector used the same CTD but it was stored on deck. Acquisition began early, included the 10m soak and sometimes included out-of-water values, so there was good evidence from both the beginning and end of the casts and pressure looked very good for both. Pumps were turned on and off at about 0.1 to 0.5db and the data look appropriate for being in water.

This cruise did not include any “out of water” sampling. The only pressure recorded <0 was in a 2-sided spike around 7db.

The surface check had an average of 2.3db with a range from 1.6db to 3.2db.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. Plots showed this setting was appropriate.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on the 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. Only the primary was studied. Using -0.7 and -0.8 records gave good results, with one working better than the other over cleaning different features.

SHIFT was run on all casts using -0.75 records for the primary conductivity. Salinity was recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings that needed attention:

There were problems in a number of casts:

Event #1 did not have warnings but it is full of pressure spikes as well; they are <2db jumps so not picked up by the usual DELETE warning setting. There are too many to enable easy editing at the top of the cast and the bottom looks corrupted by shed wakes. The central section can be edited.

Event 6 – At records 1298 and 1299 there was an obvious pressure spike. The pressure and depth were edited by interpolation in the SHFC0 file. There is no evidence of a problem in the other channels.

There are other spikes in this file that did not turn up as warnings and all the T and S data look bad between 40db and 80db.

Event 8 – There are many warning but in almost every case the primary salinity and conductivity are padded. Nonetheless, the plots looks awful. A test was run using WILDEDIT in a second run that removed the pressure spikes but left noise in the salinity. Examination of the DEL file shows that the problem may not be significant, but the SHFC0 pressure was edited by interpolation where there were warnings. There were about 8 areas with problems, each with several values needing attention.

Event #10 – The near-surface data looked poor, which was due to high variability and the fact DELETE  chose records separated in time. So the SHFC0 file was put through CLIP to remove the first 440 scans and DELETE was rerun.

##### Other Comparisons

Previous experience with these sensors –

The sensors had been used on 5 other cruises since they were last recalibrated.

- Problems were noted during 2 of the cruises (on the Franklin) suggesting pressure was too high by 1.5db or more. There were only 2 casts on the 2nd of those cruises. A lab test had a pressure of 1.25db. Cruise 2021-078 was on the Vector with the CTD stored on deck. Acquisition began early, included the 10m soak and sometimes included out-of-water values, so there was good evidence from both the beginning and end of the casts and pressure looked very good for both. Pumps were turned on and off at about 0.1 to 0.5db and the data look appropriate for being in water. Storage of the CTD in a warm lab may account for Franklin observations.

- Only the 5th cruise had salinity sampling with only 5 samples that found the primary low by 0.0002 and the secondary low by 0.0065psu.

- The dissolved oxygen sensor was recalibrated in October 2021 and used for only a few casts on 1 other cruise because of problems in the data that were actually due to a cable problem. There was calibration sampling from that cruise with a correction factor slope=1.0168, offset=0.0437.

Historic ranges  The only excursions from the climatology were temperatures that were very slightly low at about 300m for cast #20 and 200m for cast #23. Since both were at seamounts unusual conditions are expected and are likely not represented in the climatology. All salinity values were within the climatology though they were both fairly low at the 2 sections with low temperatures.

Repeat Casts – The were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving because the secondary salinity compared badly with bottles and a technician later found a blockage in the conductivity sensor.  Fluorescence was also displayed since at least one cast had instrumental spikes. Most had only very small spikes, with no large values below 100db. So no editing was applied to fluorescence.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes and others that appear corrupted by pressure spikes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or pressure spikes. Where both salinity and temperature were affected by spikes the record was generally removed. All files required editing.

The edited files were copied to *.EDT.

There were specific problems encountered in editing, as follows:

- Cast #1 required heavy editing near the top and bottom, but mostly looks ok. There was a fluorescence spike, but it is at a level where it could be real.

- Cast #6 was heavily corrupted in the downcast and quite a bit in the upcast. There were no useful data between 40db and 80db in the downcast. Saanich Inlet is subject to large changes from time to time but nothing this large and unstable is at all likely and there is no sign of a similar profile in the upcast. Comparison with data from another cruise that sampled station SI 2 weeks later shows no such odd data. Given the problems with pressure spikes there is no reason to believe this odd profile is reliable. While the upcast is better it is also affected by spikes and the usual problem of the CTD passing through the rosette wake. This cast will not be processed further.

After editing T-S plots were examined for all casts; there remain some small unstable features at the surface of cast #1 that will disappear when bin-averaged. No further editing was applied.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel. It worked well to reduce noise.

##### Recalibration

There was no salinity <25psu, so silicate does not require correction.

Pressure will not be recalibrated since it was found to be accurate when last used and there is no further information available from this cruise.

File 2022-035-recal1.ccf was prepared and run on all MRG and SAM files to apply a correction as follows:

CTD DO Corrected = CTD DO * 1.0381 + 0.0140

COMPARE was rerun and the results show that the correction was appropriate. The CTD DO is higher than bottles by an average of 0.0005mL/L (std dev 0.0131 mL/L). There is a slight pressure dependence with CTD tending to read slightly high for high DO and slightly low at high pressure. This is likely due to small errors due to incomplete flushing. The sign of such errors is generally opposite above and below the oxygen minimum zone.

CALIBRATE was run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from roughly the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

The CTD DO was higher than the titrated samples by an average of ~0.021mL/L but the standard deviation was 0.12mL/L. The recalibration was obviously effective, but with a lot of noise. When a few outliers were rejected due to high standard deviation in the CTD value, the average difference was +0.007mL/L but the standard deviation was still high at 0.100mL/L. The differences are mostly very small right at the surface and in deep water where comparisons are expected to be most error-free. The largest differences came from cast #8 which did have problems with pressure spikes, so that might be a factor. Only casts #6 and #8 have notable spikes in DO and cast #6 has been rejected because of overall spiking. The obvious spikes in cast #8 disappear after DELETE is run and editing likely removed some smaller ones, but there may remain some areas with less reliable oxygen values.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.40 mL/L from 0-100db except in areas of very large DO gradients

±0.30 mL/L from 100db-300db

±0.10 mL/L from 300db-1200db

±0.04 mL/L below 1200db

For more detail see file 2022-035-dox-comp3.xls.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

- A second SBE DO channel (with umol/kg units) was added.

- Oxygen saturation was calculated and surface values were ~130% in Saanich Inlet and between 105% and 107% in the offshore. The Saanich Inlet value came from about 4m and there was a large sub-surface DO maximum. The other values are a little higher than usually seen in the offshore. But surface SBE DO compares well with titrated DO values.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

- A second SBE DO channel with mass units was added for both the CTD DO and REORDER was run to get the pair of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

- Standards check was run. There was a warning that there was an empty Salinity:Bottle column in one cast, but that was deliberate since there was a planned sample; there is a relevant flag entry and a comment in the header.

- A header check were run. No problems were found.

- One file contains no sample data, but some sampling was done that may be added later or the file might be useful to those who analyze those samples.

- Plots of each file were examined and no problems were found.

- A cross-reference listing was produced for the CHE files and added to the end of this report.

##### Thermosalinograph processing

There was 1 file named TSG_2022-035.log. It was opened in EXCEL and does contain about 26 hours of TSG data, but there are no time or date channels. No other files with position information were available at the time of processing.

There is a break in the data with a big change in position and then another hour of data.

Start Position: 		48.724N	125.295W – west of Banfield

End of 1st section: 	51.733N  	130.921W – south of Haida Gwaii (south-east of ROV event #7)

Start of 2nd section: 	48.554N	124.864W – near mouth of Juan de Fuca Strait

End of 2ndt section: 	48.631N  	124.904W - near mouth of Juan de Fuca Strait

The latter section is close to the site of ROV event #30 but seems to proceeds in the opposite direction.

If a time channel is made available, this can be revisited, though it appears there will be no way to check the reliability of the data since it does not intersect with any CTD casts and there was no loop sampling.

Particulars

1. All bottles fired as test – no sampling.

6. Heavily corrupted by pressure spikes – not prepared for archive

8. Errors at 110m; continued cast hesitantly.

23. ROS cast but no sampling of SAL, NUTS, CHL or OXY. CHE file prepared in case needed by eDNA or Phyto analysts.

CRUISE SUMMARY – CTD & TSG

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0506 |  | Yes |  | Yes |  |  |  |
| 3 | SEABIRD | 21 |  | 3353 |  | n/a |  | Yes |  |  |  |
| Calibration Information - 0506 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 2374 |  | 3Feb2021 |  | Factory |  |  |  |  |
| Conductivity |  |  | 3184 |  | 3Mar2021 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  | 4883 |  | 4Feb2021 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  | 4395 |  | 3Mar2021 |  | Factory |  |  |  |  |
| Transmissometer |  |  | 1185DR |  | 28Apr2021 |  | IOS |  |  |  |  |
| Transmissometer |  |  | 1883DG |  | 28Apr2021 |  |  |  |  |  |  |
| SBE 43 DO sensor |  |  | 997 |  | 15Oct2021 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 3982 |  |  |  | Factory |  |  |  |  |
| Pressure Sensor |  |  | 0506 |  | 29Jan2021 |  | Factory |  |  |  |  |
| Altimeter |  |  | 76341 |  | 2Oct2021 |  | Factory |  |  |  |  |

| Calibration Information – TSG 3363 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3353 | 1Sept2021 | Factory |  |  |
| Conductivity | 3353 | 1Sept2021 | Factory |  |  |

CTD FILE CROSS-REFERENCE LIST

CHE FILE CROSS-REFERENCE LIST