## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2022-042

Agency: IOS, Ocean Sciences Division, Sidney, BC; Freshwater Institute, Central & Arctic Region, Winnipeg, MB.

Location:  Beaufort Sea

Project: Canadian Beaufort Sea - Marine Ecosystem Assessment MPA Survey

Chief Scientist: Majewski A.

Platform: F/V Frosti

Date: 3 August 2022 – 30 August 2022

Processed by: Germaine Gatien

Date of Processing: 17 July 2023  -  11 October 2023

Number of original HEX files: 47

Number of CTD files:  47	 	Number of bottle casts:   25

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (Arctic #1189) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#SCST1666DR), a SBE 43 DO sensor (#1489), a SeaPoint Fluorometer (#SCF2859), a Biospherical PAR sensor (#70501), a SeaPoint turbidity meter (#11074), an Optode Dissolved Oxygen sensor (16) and an altimeter (#72144).

The data acquisition program was SeaSave V 7.26.7.107.

# SUMMARY OF QUALITY AND CONCERNS

There was a digital bridge log and rosette sheet with positions, times (BE,BO and EN), water depths and comments. There was no list of equipment.

## There was 1 error in the configuration files used at sea. It would be helpful in processing to have the original calibration XML files from the factory, either stored in a convenient place or included in the folder where the raw data are placed.

Deployment generally included a subsurface soak between 2db and 10db. In a few cases there was a return to the surface before beginning the full cast, but generally not. Pumps were not turned on until the soak level was reached, so data were removed until the pumps came on. This unfortunately means that unpumped channels from the near-surface are also removed, the most useful being transmissivity and PAR data. If required, files are available that include data before the pumps were turned on.

There was considerable variation in the comparison of CTD salinity with bottles, as is usual when most casts are quite shallow. The primary salinity was higher than the secondary by about 0.003psu, with the secondary likely more accurate. Secondary channels were archived and no recalibration was applied to salinity.

There was no dissolved oxygen sampling. The dissolved oxygen calculation using pre-cruise calibration parameters was higher than that using post-cruise by about 0.75%, on average, but there was little change in that difference between early and late in this cruise. That suggests that the bulk of the drift occurred during use after this cruise. In any case, the errors due to calibration drift would be <0.07mL/L at the top of the range and <0.05mL/L for most of the profile. Few of the casts had dissolved oxygen data from close to the surface with pumps on; those that did had surface saturations between 99% and 102%. The casts that had a soak at 8 to 10m followed by a return to the surface had values between 100% and 101%. So the dissolved oxygen values appear to be reasonable. No recalibration was applied.

Optode and Rinko dissolved oxygen sensors were used during the cruise but those channels are not included in the files to be archived. Special files including that channel were prepared for the use of the chief scientist.

There were 71 casts (events #220-296) that were run using an RBR CTD. Processing of those data are described in document “2022-042-RBR_Processing_Report.docx”.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The file names were not in standard format with an underscore rather than dash between cruise number and event number. That was corrected.

The same XMLCON file was used throughout. Parameters from the file were checked against the Arctic calibration summary spreadsheet, and the only differences noted were:.

- The offset for the oxygen sensor was different from that in the spreadsheet of calibrations. Mike Dempsey confirmed that the value in the spreadsheet is correct and that used at sea was incorrect. The value was corrected.

- There is no multiplication factor or serial number for the SPAR, but those were obtained from the 2023 cruise using the same instrument and tested at sea.

- There was an offset to the pressure of -0.42db which is presumed to be based on observations at sea.

- There are 2 unidentified channels that correspond to Optode and Rinko oxygen sensors. Those will not be included in the final files, but the voltages will be included up to the REMOVE stage and files will be available for use of the research team.

After corrections, the file was saved as 2022-042-ctd.xmlcon.

Post-cruise calibrations were available for the T and C sensors, so before converting all files, tests were done on 2 files. They were converted using pre-cruise and post-cruise calibrations for T and C. The differences were:

| File Name | Press | T1-T0 | C1-C0 | S1-S0 |
| --- | --- | --- | --- | --- |
| Pre-cruise calibrations |  |  |  |  |
| 2022-042-0002 | 345 | +0.0005 | -0.00022 | -0.0034 |
| 2022-042-0214 | 345 | +0.0001 | -0.00025 | -0.0034 |
| Post-cruise calibrations |  |  |  |  |
| 2022-042-0002 | 345 | +0.0003 | +0.00130 | +0.0163 |
| 2022-042-0214 | 345 | -0.0001 | +0.00124 | +0.0164 |

These results show little difference through the cruise and show much better results using the pre-cruise calibrations. It looks like at least 1 conductivity sensor suffered some damage after this cruise.

To confirm this the salinity from 2 bottles from the 2 casts were compared, roughly, with the bottle salinity. Only the pre-cruise calibration was checked.

| Event # | Sample # | Depth | Sal0-Sal Bot | Sal1 – Sal Bot | Sal1-Sal0 |
| --- | --- | --- | --- | --- | --- |
| 2 | 5 | 300 | +0.0018 | -0.0017 | -0.0035 |
| 2 | 4 | 345 | +0.0025 | -0.0011 | -0.0036 |
| 214 | 305 | 300 | +0.0050 | +0.0016 | -0.0034 |
| 214 | 304 | 343 | +0.0033 | -0.0003 | -0.0030 |

This comparison suggests only small changes through the cruise and the differences between primary and secondary showing very little change. The large change indicated by the post-cruise calibration must have occurred after this cruise. So using the pre-cruise calibration parameters is the obvious choice. A more thorough comparison will be done later.

The dissolved oxygen data have lower values by about 0.75% if the post-cruise calibrations are used. We do expect these sensors to drift down with time. Results from early and late in this cruise show no significant drop during the cruise, so most of the drift is presumed to have occurred after 2022-042. In any case, the errors would be <0.07mL/L at the top of the range and <0.05mL/L for most of the profile.

The rosette log sheets were obtained as well as the salinity analysis spreadsheet.

The cruise summary sheet was completed.

There was no history available for the pressure, temperature, conductivity and DO sensors since their last factory calibration and this was likely the first use for this equipment since the factory visit.

##### Conversion of Full Files from Raw Data

The deployment method varied through the cruise.

- For most casts there was a soak for about 5 minutes with the pumps on for about 3 minutes followed by the drop to the bottom of the cast. The depth of the soak varied from about 2m to 8m.

- For a few casts there was a drop to about 10m, pumps were turned on at 10m, there was a wait of about 3minutes followed by a rise to near the surface before the full cast was run.

All hex files were converted using 2022-042-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. A few features were noted:

- Temperature and conductivity pairs are very close during downcasts, a little noisier on upcasts.

- Turbidity and Transmissivity mirror each other.

- Altimetry looks good.

- PAR looks good; SPAR is extremely spiky but agrees with PAR fairly well at the surface. Many of the SPAR spikes are to 0, so this is more likely electrical in origin than due to varying cloud cover or shading of the sensor.

- Descent rates were kept high, on average, though some were very noisy, others quiet.

- Dissolved oxygen traces look normal below 15db, but complex above that. Dissolved oxygen often shows a bulge around 6db that could be due to release of bubbles, shed wakes or sharp temperature gradients. Upcasts have similar bulges but in the opposite direction which look likely to be due to temperature effects.

- A few casts were found where the pumps were still on at the end of casts; conductivity dropped suddenly to near zero when the pressure was between 0db and +0.3db. No recalibration of pressure is necessary.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2022-042-ctd.xmlcon.

The files were converted to IOS format.

CLEAN was run to create BOT files.

Temperature and salinity were plotted for all BOT files. There are surface spikes, but none required editing.

A preliminary header check turned up no problems, other than the surface spikes.

The minimum fluorescence is 0.06ug/L.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged. Output was SAMAVG.

At the time of processing the only sample analysis data available that had quality flags and analysis comments attached was salinity.

The salinity analysis spreadsheet was examined to see what comments should be included in the header file. These were used to create file 2022-042-bot-hdr.txt which will updated during processing.

SALINITY

Salinity analysis was obtained in spreadsheet QF2022-042_SAL*.xlsx. The sheet with final data for merging with CTD data was saved as 2022-042SAL.csv.. Event numbers were added based on the rosette log sheets.

The analysis was done about 5 months after collection.

The file was converted into SAL files.

The SAL files were merged with CST files with extensions MRG1.

The files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files, to produce MRG files.

The MRG files were put through CLEAN to produce MRGCLN2 files; 0s were entered into any empty flag channels.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was a lot of salinity sampling but most was in shallow water where the standard deviations in the CTD data were very high during the 10s window around firing time.

Comparisons were run using a variety of methods for excluding outliers:

- When only bottles above 50db were excluded the primary was high by 0.007psu (std dev 0.017) and the secondary high by 0.004 (0.017).

- When bottles above 50db were excluded as well as bottles fired at the bottom of the cast, the primary was high by 0.006psu (std dev 0.018) and the secondary high by 0.0034 (0.018).

- When only bottles above 50db were included as well as bottles for which the CTD standard deviation during the 10s window was >0.0008psu, the primary was high by 0.0015psu (std dev 0.018) and the secondary low by 0.0012 (0.017).

- When all bottles with standard deviation <00008psu were included the primary was high by 0.0004psu (std dev 0.021) and the secondary low by 0.0016 (0.020).

- When only bottles fired at the bottom of casts are included, excluding those with standard deviation >0.0008psu,the primary was high by 0.0139psu (0.0206) and the secondary by 0.0095 (0.0253), but when bottles from above 200m were excluded those differences dropped to 0.0072 and 0.0006. So, there is slight evidence that bottles fired close to the sea floor tend towards being higher than CTD values whereas bottles fired above bottom have a tendency to read close to or lower than CTD values. This is expected since any flushing errors at the bottom would have the opposite sign to those fired on upcasts. For these data, the difference does not look significant below 200db, so flushing was probably good near the bottom of most casts.

Plots were also made of differences versus File Pair # (proxy for time) the trend was slight for both, especially so for the secondary salinity.

The Primary salinity appears to be higher than the secondary by about 0.003psu in most of the comparisons, though the difference is smaller . It is likely that the secondary salinity is more accurate than the primary; it reads slightly low when limited to bottles below 50db and with standard deviation <0.0008psu. Given that there was a long wait before analysis there are likely small errors due to desorption &/or evaporation of samples and/or incomplete flushing, both channels likely are reading higher than appears in the comparison and the secondary is likely more accurate. However, variability is high in these comparisons, so this is weak evidence.

Neither salinity channel looks like recalibration is justified, but if a mix of channels are used, then subtracting 0.003psu from the primary would bring the two into better alignment.

There were no significant outliers requiring adjustments of quality flags..

For details see file 2022-042-sal-comp1.xls.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts. Noise in the both temperature and dissolved oxygen make the tests difficult. The best setting varied from feature to feature, but a setting of +3s looked best overall.

ALIGNCTD was run using a setting of +3s to the dissolved oxygen channel.

##### CELLTM

The usual tests for the best settings for this routine are difficult to interpret because of the noisy upcast data, so the default settings was used.

CELLTM was run using (α = 0.0245, β = 9.5) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Program DERIVE was run a second time to look at differences between channel pairs for a few deeper casts. The results were consistent over time with a very slight increase with pressure for conductivity and salinity. Temperature differences are very small.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2022-042-0002 | 300 | +0.0003 | -0.00023 | -0.0035 | High, Steady |
| “ | 500 | +0.0002 | -0.00025 | -0.0036 | “ |
| 2022-042-0067 | 300 | +0.0003 | -0.00026 | -0.0035 | High, Steady |
| “ | 600 | +0.0002 | -0.00026 | -0.0038 | “ |
| 2022-042-0189 | 300 | +0.0003 | -0.00023 | -0.0035 | F.High, Noisy |
| “ | 450 | +0.0003 | -0.00025 | -0.0036 | “ |

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

Most casts had a brief soak at the surface and were lowered to between 2 and 10db. Pumps were then turned on and there followed a soak and then the full cast was run.

For 4 casts (#22, 50, 189, & 214) there was a return to the surface before the full cast was run, but the wait at the surface was short; a wait of at least 30s is recommended to allow water to settle from upcast.

To ensure that DELETE does not include the data from before the pumps came on CLIP was run to remove all initial data with pumps OFF. DELETE will remove the upcast before the full casts for the 4 casts that did return to about 2db before running the full cast.

##### Checking Headers

An initial cross-reference list was produced and compared to the Bridge Log. There were 3 station names that differed slightly, but for cast #151 the file header is assumed to be right. For events #137 and #189 the station names were changed to match the log.

The cruise tracks were plotted and look good; they were added to the end of this report.

A header check was done. There are some spikes that are likely from the surface and will be removed when DELETE is run. There were no negative fluorescence values.

The surface check was run on the files before CLIP was applied. The average value was +0.29db with a minimum value of +0.12db. Those are consistent with the upcast results of conductivity dropping very low when pressure was 0.0db to 0.3db. Those results suggest pressure is slightly high but well within specifications. The pressure offset had been updated in the configuration file which is presumed to be due to tests at sea, though it is not know when those occurred. No recalibration is needed.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header. Where that number was > 5m checks were made to see if the log entry differed from the header entry and whether the altimetry signal at the bottom provided a good header value. The altimetry headers checked looked fine. There were 6 casts that were examined to see if the water depth entries were reasonable.

- In a few cases (44 & 139) the log notes there was shoaling and the check value is only slightly >5m so no change was made.

- For event 40 the bottom depth was changed to 375m to match the log entry.

- For event 60 a sum of the depth sampled + Altimetry was entered (71m).

- For event 94 there was no bottom depth entered in the header; it was entered in the file header based on the log entry (43m).

- For event 141 the CTD did not get within 15m of the bottom so there is no altimetry header entry.

- For event 186 a sum of the depth sampled + Altimetry was entered (300m).

- For event #206 the bottom depth was missing and was entered based on the log (93m).

These changes were made to the CLN and CLIP for 5 casts and SAM & SAMAVG files for 2 casts. MERGE and CLEAN were rerun for the bottle files.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel in all casts using the usual advance of +24 records. Plots made after this step were hard to interpret due to noisy fluorescence and many stops for bottles but there is clearly some improvement in matching the fluorescence offset to the temperature offset.

Dissolved Oxygen

Channel Oxygen:Voltage:SBE was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel; it is difficult to judge but the alignment looked reasonable.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with the settings -1.3s for the primary and -0.7s for the secondary. SHIFT was run twice using those settings.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The output files were copied to *.EDT.

##### Other Comparisons

Previous experience with these sensors –

There was no history available for the pressure, temperature, conductivity and DO sensors since their latest factory calibration.

Historic ranges – Local climatology was not available.

Repeat Casts – The only casts that are close together in time and space are too shallow to test repeatability.

Post-Cruise Calibration – There were post-cruise calibrations in late 2022 for the temperature , conductivity and dissolved oxygen sensors. The results are discussed in section 2. No significant calibration drift appears to have occurred during this cruise.

##### DETAILED EDITING

The secondary temperature and salinity channels were selected for editing, though they were only slightly less noisy than the primary.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and misalignment of T and C and records corrupted by shed wakes. For most casts the editing was mostly at the top and bottom of casts and salinity editing was very light.

All files were edited except for those from casts #36, 88 and 149 which needed no editing. The output files were copied to EDT.

T-S plots were examined and the results looked ok; there are some small unstable features in a few casts. There was no obvious instrumental cause of some of these features so they may be real. In other cases the descent rate is low and T and S values reflect small variations in descent rate; it is impossible to determine which data are better.

##### Initial Recalibration

There was no Dissolved Oxygen calibration sampling. Surface saturation of DO was derived and plotted.

CHANGE UNITS was used to derive dissolved oxygen in mass units.

DERIVED QUANTITIES was used to calculate DO surface saturation.

Plots were made of DO surface saturation and values ranged from 96% to 104% with most between 100% and 101%.

Summary of calibration review:

- Pressure checks indicate no recalibration is needed.

- The secondary salinity appears to be very close to bottles. There was no significant drift during the cruise; drift noted in post-cruise calibrations likely occurred after this cruise.

- Dissolved oxygen values drifted little between calibrations. There may be errors due to other effects, but without bottle sampling no estimate is possible for that. Surface saturation values look normal, so calibration drift is likely low.

No recalibration was applied.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Special Files (CTDspec)

REMOVE was run to remove the following channels: : Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity:T0:C0, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag from all casts. Channels uploy0, upoly1 were left in these files which were prepared for in-house investigations. (*REMspec)

REORDER was not needed.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

-------------------------------------------------------------------------------------------------------------------

Data Processing Notes:

----------------------

The files in this folder were prepared for the use of IOS researchers. They include

data from Rinko and Optode oxygen sensors being studied, but not to be added to the

IOS archive.

Transmissivity, Fluorescence, Turbidity, PAR and Surface PAR data are nominal and

unedited, except that some records were removed in editing temperature and salinity.

Salinity:T1:C1 values compared well with bottle samples. A post-cruise calibration

suggested values might have drifted higher but it is likely most drift occurred

after this cruise since the difference between the 2 salinity channels was much

larger than observed during this cruise. No recalibration was applied

There was no dissolved oxygen calibration sampling but the post-cruise calibration

indicated little drift and the surface saturation looked normal.

While the CTD fluorescence data are expressed in concentration units, they do not

always compare well to extracted chlorophyll samples, particularly for casts far

from shore. No extracted chlorophyll data were available at the time of processing.

For details on the processing see the report: 2022-042_Processing_Report.docx.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run to remove the following channels: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity:T0:C0, Oxygen:Voltage:SBE, uploy0, upoly1,Altimeter, Status:Pump, Descent_Rate and Flag from all casts

REORDER was not needed.

HEADER EDIT was used to fix formats and channel names and to add comments in file 2022-042-sbe911-hdr.txt. (Comments are the same as given in the previous section except for the 1st paragraph.)

The Standards Check routine was run and no problems were found.

A header check was run on the final files and no problems were found.

A cross-reference file was produced and looks fine.

Plots of each file were examined. The only problem noted was that the PAR values are much lower than SPAR values in most cases. Checks of raw files show much higher PAR values at the surface with a sharp drop-off to low values in the top 5db. Since the pumps were off until between 2db and 8db, the comparison is not useful. The comparison looks reasonably good in the CLN files; if needed those files could be used to produce near-surface files with PAR and Transmissivity.

##### Final Bottle Files

No recalibration was required.

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run to remove the same channels as for the profile files.

A second SBE DO channel was added to the CTD DO with mass units.

REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files; an error in the Data Description was found and corrected,

Header Edit was rerun. No further errors were found.

A header check was run on the final CHE files and no problems were found.

A cross reference list looks fine.

##### Sensor History

These sensors are not included in the OSD sensor history. They are maintained by the Arctic research group, and get calibrated at the factory each year. No other cruises using these sensors in 2022 were processed by the OSD Data Products Team.

Particulars –

22. Return to ~2db after soak, followed by full cast.

50. Return to 2db after soak, followed by full cast.

143.Pumps turned on at 6db; no return to surface before drop to bottom.

189. Return to ~5.8db after soak at 10db, then down to bottom..

212. Pumps not on until CTD at about 8db – no return to surface after soak.

214. Return to ~3.4db after soak, followed by full cast.

CRUISE SUMMARY – CTD

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 1189 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 1189 |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  | 5830 |  | 30Oct2019 |  | Factory |  | 30Dec2022 | Factory |  |
| Conductivity |  |  | 4327 |  | 26Nov2019 |  | Factory |  | 15Dec2022 | Factory |  |
| Secondary Temp. |  |  | 5831 |  | 30Oct2019 |  | Factory |  | 17Oct2022 | Factory |  |
| Secondary Cond. |  |  | 4339 |  | 26Nov2019 |  | Factory |  | 22Dec2022 | Factory |  |
| Transmissometer |  |  | 1666 |  | 10Jun2022 |  | Laurier |  |  |  |  |
| SBE 43 DO sensor |  |  | 1489 |  | 1Oct2021 |  | Factory |  | 22Dec2022 | Factory |  |
| PAR |  |  | 70501 |  | 4Apr2016 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  |  | 2859 |  | 21May2020 |  | IOS |  | 20-Feb-2023 | IOS |  |
| Surface PAR |  |  | 20279 |  | 4Apr2016 |  | Factory |  | Updated at sea 8-Sept-2023 |  |  |
| Pressure Sensor |  |  | 130015 |  | 23Dec2013 |  | Factory |  |  |  |  |
| Altimeter |  |  | 72144 |  | 12May2017 |  | Factory |  |  |  |  |
| Turbidity Meter |  |  | 11074 |  | 23Mar2007 |  | Factory |  |  |  |  |