## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.6.1 | Run at 2023/02/28 12:14:06 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | OXY_BOT | OXY_SBE |  | OXY_SBE - OXY_BOT
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | OXY_BOT | OXY_SBE | OXY_SBE | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2022-070-0001.mrg | 2022-070-0001.thn1 | 2 | 173.971 | 173.451 | 0.64984 | -0.52049 | 0.044 | 0.0485 | 0.00070711 | 0.0045
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 150.433 | 150.001 | 0 | -0.43199 | 0.04 | 0.053 | 0 | 0.013
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 125.2 | 125.01 | 0 | -0.18999 | 0.142 | 0.241 | 0 | 0.099
2022-070-0001.mrg | 2022-070-0001.thn1 | 2 | 100.365 | 100.504 | 0.71913 | 0.1385 | 0.863 | 1.0745 | 0.013435 | 0.2115
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 75.857 | 75 | 0 | -0.857 | 0.805 | 0.63 | 0 | -0.175
2022-070-0001.mrg | 2022-070-0001.thn1 | 2 | 50.354 | 50.4965 | 0.71064 | 0.1425 | 2.733 | 2.851 | 0.05374 | 0.118
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 40.425 | 41.004 | 0 | 0.579 | 2.987 | 3.079 | 0 | 0.092
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 30.908 | 30.003 | 0 | -0.905 | 3.141 | 3.281 | 0 | 0.14
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 20.815 | 19.996 | 0 | -0.819 | 3.249 | 3.394 | 0 | 0.145
2022-070-0001.mrg | 2022-070-0001.thn1 | 2 | 4.025 | 4.5095 | 0.70357 | 0.4845 | 3.779 | 4.1225 | 0.16193 | 0.3435
2022-070-0001.mrg | 2022-070-0001.thn1 | 0 | 1.314 | 0 | 0 | -1.314 | 3.918 | 0 | 0 | -3.918
2022-070-0001.mrg | 2022-070-0001.thn1 | 1 | 1.602 | 2.477 | 0 | 0.875 | 5.696 | 6.125 | 0 | 0.429
2022-070-0005.mrg | 2022-070-0005.thn1 | 2 | 251.927 | 252.453 | 0.65195 | 0.526 | 1.2 | 1.2085 | 0.0035355 | 0.0084999
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 249.571 | 249.989 | 0 | 0.418 | 1.203 | 1.213 | 0 | 0.01
2022-070-0005.mrg | 2022-070-0005.thn1 | 2 | 200.989 | 200.506 | 0.72267 | -0.483 | 1.254 | 1.2595 | 0.0035355 | 0.0055001
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 175.678 | 175.221 | 0 | -0.457 | 1.318 | 1.3 | 0 | -0.018
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 149.697 | 149.987 | 0 | 0.28999 | 1.654 | 1.55 | 0 | -0.104
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 124.794 | 125.003 | 0 | 0.209 | 2.134 | 2.066 | 0 | -0.068
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 100.997 | 101.002 | 0 | 0.0049972 | 2.096 | 2.145 | 0 | 0.049
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 75.974 | 74.994 | 0 | -0.98 | 2.319 | 2.421 | 0 | 0.102
2022-070-0005.mrg | 2022-070-0005.thn1 | 2 | 50.355 | 50.477 | 0.71701 | 0.122 | 2.826 | 2.8605 | 0.00070722 | 0.0345
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 40.433 | 41.036 | 0 | 0.603 | 2.955 | 3.277 | 0 | 0.322
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 30.019 | 30.012 | 0 | -0.007 | 3.203 | 3.267 | 0 | 0.064
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 20.337 | 19.846 | 0 | -0.491 | 3.944 | 3.797 | 0 | -0.147
2022-070-0005.mrg | 2022-070-0005.thn1 | 2 | 10.242 | 10.491 | 0.69579 | 0.249 | 4.64 | 4.1235 | 0.027577 | -0.5165
2022-070-0005.mrg | 2022-070-0005.thn1 | 2 | 5.016 | 5.4895 | 0.6965 | 0.4735 | 4.77 | 4.543 | 0.10748 | -0.227
2022-070-0005.mrg | 2022-070-0005.thn1 | 1 | 3.229 | 3.985 | 0 | 0.756 | 4.785 | 4.711 | 0 | -0.074
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 220.51 | 219.979 | 0 | -0.53099 | 1.227 | 1.242 | 0 | 0.015
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 202.961 | 202.99 | 0 | 0.029007 | 1.229 | 1.244 | 0 | 0.015
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 177.579 | 177.016 | 0 | -0.56299 | 1.265 | 1.305 | 0 | 0.04
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 151 | 150.001 | 0 | -0.99899 | 1.526 | 1.592 | 0 | 0.066
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 124.419 | 124.984 | 0 | 0.565 | 1.76 | 1.839 | 0 | 0.079
2022-070-0008.mrg | 2022-070-0008.thn1 | 2 | 100.761 | 100.496 | 0.69862 | -0.265 | 1.836 | 2.2775 | 0.044548 | 0.4415
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 77.208 | 77.016 | 0 | -0.192 | 2.099 | 2.313 | 0 | 0.214
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 52.126 | 51.988 | 0 | -0.138 | 2.393 | 2.689 | 0 | 0.296
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 40.695 | 41.027 | 0 | 0.332 | 2.779 | 3.167 | 0 | 0.388
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 29.925 | 30.003 | 0 | 0.078001 | 3.229 | 3.523 | 0 | 0.294
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 20.543 | 19.991 | 0 | -0.552 | 3.736 | 4.188 | 0 | 0.452
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 9.677 | 10.004 | 0 | 0.327 | 4.273 | 4.277 | 0 | 0.0040002
2022-070-0008.mrg | 2022-070-0008.thn1 | 2 | 5.381 | 5.4935 | 0.71347 | 0.1125 | 4.283 | 4.2675 | 0.023335 | -0.0155
2022-070-0008.mrg | 2022-070-0008.thn1 | 1 | 2.4 | 2.595 | 0 | 0.195 | 4.728 | 4.961 | 0 | 0.233
2022-070-0012.mrg | 2022-070-0012.thn1 | 2 | 185.977 | 185.396 | 0.58266 | -0.58101 | 1.166 | 1.189 | 0.0014143 | 0.023
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 175.222 | 174.996 | 0 | -0.226 | 1.174 | 1.197 | 0 | 0.023
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 150.855 | 150.001 | 0 | -0.85399 | 1.174 | 1.201 | 0 | 0.027
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 125.474 | 125.011 | 0 | -0.463 | 1.184 | 1.211 | 0 | 0.027
2022-070-0012.mrg | 2022-070-0012.thn1 | 2 | 100.519 | 100.497 | 0.70781 | -0.0215 | 1.517 | 1.521 | 0.0070711 | 0.0040001
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 76.384 | 76.991 | 0 | 0.60699 | 1.87 | 1.846 | 0 | -0.024
2022-070-0012.mrg | 2022-070-0012.thn1 | 2 | 50.882 | 50.506 | 0.69862 | -0.376 | 2.801 | 2.702 | 0.0042427 | -0.099
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 40.222 | 41.009 | 0 | 0.787 | 2.902 | 2.919 | 0 | 0.017
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 30.364 | 29.983 | 0 | -0.381 | 3.156 | 3.096 | 0 | -0.06
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 20.9 | 19.989 | 0 | -0.911 | 3.48 | 3.443 | 0 | -0.037
2022-070-0012.mrg | 2022-070-0012.thn1 | 2 | 10.32 | 10.4915 | 0.70781 | 0.1715 | 3.943 | 4.093 | 0.0014141 | 0.15
2022-070-0012.mrg | 2022-070-0012.thn1 | 2 | 5.38 | 5.5005 | 0.71488 | 0.1205 | 3.841 | 4.0935 | 0.00070706 | 0.2525
2022-070-0012.mrg | 2022-070-0012.thn1 | 1 | 1.949 | 2.693 | 0 | 0.744 | 3.949 | 4.083 | 0 | 0.134
2022-070-0015.mrg | 2022-070-0015.thn1 | 2 | 174.863 | 175.173 | 0.1683 | 0.31 | 1.201 | 1.2195 | 0.0035355 | 0.0185
2022-070-0015.mrg | 2022-070-0015.thn1 | 0 | 151.331 | 0 | 0 | -151.331 | 1.213 | 0 | 0 | -1.213
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 124.84 | 124.991 | 0 | 0.151 | 1.252 | 1.291 | 0 | 0.039
2022-070-0015.mrg | 2022-070-0015.thn1 | 2 | 100.465 | 100.495 | 0.69791 | 0.030502 | 1.755 | 1.719 | 0.041012 | -0.036
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 75.521 | 75.006 | 0 | -0.51501 | 2.528 | 2.327 | 0 | -0.201
2022-070-0015.mrg | 2022-070-0015.thn1 | 2 | 49.872 | 49.502 | 0.70428 | -0.37 | 2.796 | 2.797 | 0.0070711 | 0.00099993
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 40.439 | 40.995 | 0 | 0.556 | 2.876 | 2.902 | 0 | 0.026
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 30.9 | 30.001 | 0 | -0.899 | 2.954 | 2.989 | 0 | 0.035
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 20.659 | 19.988 | 0 | -0.671 | 2.998 | 3.048 | 0 | 0.05
2022-070-0015.mrg | 2022-070-0015.thn1 | 2 | 10.636 | 10.5055 | 0.70781 | -0.1305 | 3.118 | 3.154 | 0.0014141 | 0.036
2022-070-0015.mrg | 2022-070-0015.thn1 | 2 | 5.489 | 5.5105 | 0.7163 | 0.0215 | 3.229 | 3.2195 | 0.026163 | -0.0095
2022-070-0015.mrg | 2022-070-0015.thn1 | 1 | 2.133 | 3.129 | 0 | 0.996 | 3.275 | 3.282 | 0 | 0.007
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 113.757 | 114 | 0 | 0.243 | 1.594 | 1.594 | 0 | 0
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 101.292 | 100.998 | 0 | -0.294 | 1.843 | 1.82 | 0 | -0.023
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 75.503 | 74.991 | 0 | -0.512 | 2.037 | 2.072 | 0 | 0.035
2022-070-0018.mrg | 2022-070-0018.thn1 | 2 | 50.294 | 50.4905 | 0.72054 | 0.1965 | 2.843 | 2.8865 | 0.014849 | 0.0435
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 40.684 | 40.992 | 0 | 0.308 | 3.176 | 3.314 | 0 | 0.138
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 30.131 | 29.985 | 0 | -0.146 | 3.384 | 3.462 | 0 | 0.078
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 20.472 | 19.999 | 0 | -0.473 | 3.42 | 3.885 | 0 | 0.465
2022-070-0018.mrg | 2022-070-0018.thn1 | 2 | 10.449 | 10.5035 | 0.70923 | 0.0545 | 3.742 | 3.9585 | 0.02192 | 0.2165
2022-070-0018.mrg | 2022-070-0018.thn1 | 2 | 5.314 | 5.504 | 0.70711 | 0.19 | 3.952 | 4.034 | 0.0070712 | 0.082
2022-070-0018.mrg | 2022-070-0018.thn1 | 1 | 2.083 | 2.212 | 0 | 0.129 | 4.01 | 4.047 | 0 | 0.037
2022-070-0022.mrg | 2022-070-0022.thn1 | 2 | 123.538 | 123.143 | 0.35851 | -0.3955 | 1.893 | 1.8955 | 0.0035355 | 0.0024999
2022-070-0022.mrg | 2022-070-0022.thn1 | 2 | 100.844 | 100.495 | 0.69226 | -0.3485 | 2.259 | 2.2795 | 0.0021212 | 0.0205
2022-070-0022.mrg | 2022-070-0022.thn1 | 1 | 75.71 | 75.01 | 0 | -0.7 | 2.762 | 2.691 | 0 | -0.071
2022-070-0022.mrg | 2022-070-0022.thn1 | 2 | 50.93 | 50.5045 | 0.70498 | -0.4255 | 3.029 | 3.0885 | 0.0035356 | 0.0595
2022-070-0022.mrg | 2022-070-0022.thn1 | 1 | 40.77 | 40.999 | 0 | 0.229 | 3.133 | 3.284 | 0 | 0.151
2022-070-0022.mrg | 2022-070-0022.thn1 | 1 | 30.52 | 29.988 | 0 | -0.532 | 3.152 | 3.391 | 0 | 0.239
2022-070-0022.mrg | 2022-070-0022.thn1 | 1 | 20.576 | 19.998 | 0 | -0.578 | 3.347 | 3.692 | 0 | 0.345
2022-070-0022.mrg | 2022-070-0022.thn1 | 2 | 10.48 | 10.5015 | 0.71488 | 0.021501 | 3.734 | 3.731 | 0.0098995 | -0.003
2022-070-0022.mrg | 2022-070-0022.thn1 | 2 | 5.419 | 5.494 | 0.71701 | 0.075 | 3.759 | 3.7395 | 0.0035356 | -0.0195
2022-070-0022.mrg | 2022-070-0022.thn1 | 1 | 2.167 | 2.708 | 0 | 0.541 | 3.788 | 3.792 | 0 | 0.004
2022-070-0026.mrg | 2022-070-0026.thn1 | 2 | 145.084 | 145.194 | 0.54446 | 0.11 | 2.075 | 2.058 | 0.012728 | -0.017
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 125.159 | 124.991 | 0 | -0.168 | 2.13 | 2.187 | 0 | 0.057
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 101.933 | 100.995 | 0 | -0.938 | 2.701 | 2.678 | 0 | -0.023
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 75.768 | 75.007 | 0 | -0.76099 | 2.953 | 2.997 | 0 | 0.044
2022-070-0026.mrg | 2022-070-0026.thn1 | 2 | 50.354 | 50.4985 | 0.69226 | 0.1445 | 3.192 | 3.2705 | 0.028991 | 0.0785
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 40.399 | 41.004 | 0 | 0.605 | 3.345 | 3.369 | 0 | 0.024
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 30.437 | 29.996 | 0 | -0.441 | 3.37 | 3.404 | 0 | 0.034
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 20.46 | 19.988 | 0 | -0.472 | 3.37 | 3.471 | 0 | 0.101
2022-070-0026.mrg | 2022-070-0026.thn1 | 2 | 10.381 | 10.512 | 0.71701 | 0.131 | 3.414 | 3.465 | 0.0028284 | 0.051
2022-070-0026.mrg | 2022-070-0026.thn1 | 2 | 5.504 | 5.5 | 0.70993 | -0.0040002 | 3.445 | 3.485 | 0.0056569 | 0.04
2022-070-0026.mrg | 2022-070-0026.thn1 | 1 | 2.188 | 2.202 | 0 | 0.014 | 3.465 | 3.543 | 0 | 0.078
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 218.607 | 218.883 | 0 | 0.276 | 2.662 | 2.662 | 0 | 0
2022-070-0031.mrg | 2022-070-0031.thn1 | 2 | 200.849 | 200.508 | 0.70215 | -0.34149 | 2.667 | 2.6645 | 0.00070706 | -0.0025001
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 174.919 | 175.002 | 0 | 0.082993 | 2.671 | 2.687 | 0 | 0.016
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 150.246 | 150.011 | 0 | -0.235 | 2.692 | 2.71 | 0 | 0.018
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 125.556 | 125.005 | 0 | -0.551 | 2.748 | 2.799 | 0 | 0.051
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 101.598 | 101.009 | 0 | -0.589 | 3.005 | 3.01 | 0 | 0.0049999
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 76.217 | 77.007 | 0 | 0.79 | 3.152 | 3.255 | 0 | 0.103
2022-070-0031.mrg | 2022-070-0031.thn1 | 2 | 50.361 | 50.495 | 0.71276 | 0.134 | 3.36 | 3.36 | 0.0028284 | 2.3842e-07
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 40.31 | 41.01 | 0 | 0.7 | 3.42 | 3.376 | 0 | -0.044
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 30.159 | 29.993 | 0 | -0.166 | 3.423 | 3.516 | 0 | 0.093
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 21.346 | 22 | 0 | 0.654 | 3.477 | 3.699 | 0 | 0.222
2022-070-0031.mrg | 2022-070-0031.thn1 | 2 | 11.356 | 11.4935 | 0.70498 | 0.1375 | 3.843 | 3.7645 | 0.030406 | -0.0785
2022-070-0031.mrg | 2022-070-0031.thn1 | 2 | 5.127 | 5.4995 | 0.72478 | 0.3725 | 4.035 | 4.353 | 0.13435 | 0.318
2022-070-0031.mrg | 2022-070-0031.thn1 | 1 | 2.02 | 2.748 | 0 | 0.728 | 4.157 | 5.394 | 0 | 1.237
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 226.953 | 226.669 | 0.95459 | -0.284 | 2.689 | 2.715 | 0.0056569 | 0.026
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 200.773 | 200.49 | 0.71277 | -0.28301 | 2.683 | 2.7145 | 0.006364 | 0.0315
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 177.929 | 176.984 | 0 | -0.94501 | 2.718 | 2.747 | 0 | 0.029
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 152.28 | 152.993 | 0 | 0.713 | 2.846 | 2.806 | 0 | -0.04
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 124.759 | 124.993 | 0 | 0.23399 | 3.021 | 3.062 | 0 | 0.041
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 100.718 | 100.496 | 0.70003 | -0.222 | 3.095 | 3.1095 | 0.006364 | 0.0145
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 75.737 | 75.007 | 0 | -0.73 | 3.284 | 3.205 | 0 | -0.079
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 51.033 | 51.509 | 0.70286 | 0.476 | 3.356 | 3.39 | 0.0084852 | 0.034
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 40.858 | 40.997 | 0 | 0.139 | 3.376 | 3.438 | 0 | 0.062
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 30.841 | 29.982 | 0 | -0.859 | 3.398 | 3.506 | 0 | 0.108
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 20.802 | 19.989 | 0 | -0.813 | 3.478 | 3.584 | 0 | 0.106
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 10.629 | 10.501 | 0.70711 | -0.128 | 3.979 | 3.885 | 0.014142 | -0.094
2022-070-0036.mrg | 2022-070-0036.thn1 | 2 | 5.121 | 5.493 | 0.70145 | 0.372 | 4.076 | 3.9215 | 0.0021213 | -0.1545
2022-070-0036.mrg | 2022-070-0036.thn1 | 1 | 2.352 | 2.71 | 0 | 0.358 | 4.192 | 3.974 | 0 | -0.218
2022-070-0048.mrg | 2022-070-0048.thn1 | 2 | 171.638 | 171.333 | 0.59468 | -0.3045 | 2.781 | 2.808 | 0.0084852 | 0.027
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 150.523 | 149.986 | 0 | -0.537 | 2.904 | 2.942 | 0 | 0.038
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 124.733 | 125.003 | 0 | 0.27 | 2.912 | 2.915 | 0 | 0.003
2022-070-0048.mrg | 2022-070-0048.thn1 | 2 | 100.168 | 100.512 | 0.71347 | 0.34351 | 3.136 | 3.1775 | 0.0049497 | 0.0415
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 74.891 | 75.002 | 0 | 0.111 | 3.27 | 3.298 | 0 | 0.028
2022-070-0048.mrg | 2022-070-0048.thn1 | 2 | 50.976 | 50.4985 | 0.71064 | -0.47751 | 3.316 | 3.3765 | 0.019092 | 0.0605
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 40.496 | 41.003 | 0 | 0.507 | 3.388 | 3.446 | 0 | 0.058
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 30.708 | 30.014 | 0 | -0.694 | 3.519 | 3.526 | 0 | 0.007
2022-070-0048.mrg | 2022-070-0048.thn1 | 1 | 20.46 | 19.997 | 0 | -0.463 | 3.641 | 3.7 | 0 | 0.059
2022-070-0048.mrg | 2022-070-0048.thn1 | 2 | 10.634 | 10.4875 | 0.70216 | -0.1465 | 4.196 | 4.647 | 0.43558 | 0.451
2022-070-0048.mrg | 2022-070-0048.thn1 | 2 | 4.732 | 4.511 | 0.70711 | -0.221 | 5.495 | 5.543 | 0.10889 | 0.048
2022-070-0048.mrg | 2022-070-0048.thn1 | 0 | 1.973 | 0 | 0 | -1.973 | 6.35 | 0 | 0 | -6.35
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 317.723 | 317.761 | 0 | 0.037994 | 2.585 | 2.595 | 0 | 0.01
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 300.981 | 300.002 | 0 | -0.97897 | 2.594 | 2.616 | 0 | 0.022
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 250.943 | 250.007 | 0 | -0.93599 | 2.767 | 2.78 | 0 | 0.013
2022-070-0052.mrg | 2022-070-0052.thn1 | 2 | 200.755 | 200.494 | 0.69437 | -0.261 | 2.771 | 2.7875 | 0.0049497 | 0.0165
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 175.766 | 174.997 | 0 | -0.76901 | 2.782 | 2.812 | 0 | 0.03
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 150.836 | 150.003 | 0 | -0.83299 | 2.802 | 2.81 | 0 | 0.0079999
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 125.25 | 124.996 | 0 | -0.254 | 2.811 | 2.856 | 0 | 0.045
2022-070-0052.mrg | 2022-070-0052.thn1 | 2 | 100.946 | 100.492 | 0.70994 | -0.45399 | 2.984 | 3.051 | 0.022627 | 0.067
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 75.596 | 75.015 | 0 | -0.581 | 3.05 | 3.075 | 0 | 0.025
2022-070-0052.mrg | 2022-070-0052.thn1 | 2 | 50.55 | 50.5025 | 0.70498 | -0.047497 | 3.287 | 3.338 | 0.011314 | 0.051
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 40.824 | 40.99 | 0 | 0.166 | 3.321 | 3.363 | 0 | 0.042
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 30.487 | 29.993 | 0 | -0.494 | 3.338 | 3.383 | 0 | 0.045
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 20.559 | 19.999 | 0 | -0.56 | 3.391 | 3.494 | 0 | 0.103
2022-070-0052.mrg | 2022-070-0052.thn1 | 2 | 10.814 | 10.504 | 0.71276 | -0.31 | 3.908 | 4.305 | 0.039598 | 0.397
2022-070-0052.mrg | 2022-070-0052.thn1 | 2 | 5.015 | 5.4955 | 0.71064 | 0.4805 | 6.414 | 5.924 | 0.30264 | -0.49
2022-070-0052.mrg | 2022-070-0052.thn1 | 1 | 2.201 | 3.037 | 0 | 0.836 | 6.846 | 6.938 | 0 | 0.092
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 371.186 | 371.267 | 0 | 0.080994 | 2.543 | 2.553 | 0 | 0.01
2022-070-0057.mrg | 2022-070-0057.thn1 | 0 | 350.655 | 0 | 0 | -350.655 | 2.6 | 0 | 0 | -2.6
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 300.274 | 299.998 | 0 | -0.276 | 2.652 | 2.656 | 0 | 0.004
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 250.519 | 249.994 | 0 | -0.52499 | 2.673 | 2.684 | 0 | 0.011
2022-070-0057.mrg | 2022-070-0057.thn1 | 2 | 200.899 | 200.491 | 0.69863 | -0.40801 | 2.712 | 2.707 | 0.0042427 | -0.0049999
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 175.92 | 174.999 | 0 | -0.921 | 2.744 | 2.768 | 0 | 0.024
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 150.621 | 149.988 | 0 | -0.633 | 2.8 | 2.811 | 0 | 0.011
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 125.721 | 125.005 | 0 | -0.716 | 2.818 | 2.854 | 0 | 0.036
2022-070-0057.mrg | 2022-070-0057.thn1 | 2 | 100.642 | 100.503 | 0.71135 | -0.139 | 2.817 | 2.864 | 0.0084852 | 0.047
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 75.456 | 74.997 | 0 | -0.459 | 3.06 | 3.091 | 0 | 0.031
2022-070-0057.mrg | 2022-070-0057.thn1 | 2 | 50.44 | 50.493 | 0.70993 | 0.053001 | 3.274 | 3.2885 | 0.019092 | 0.0145
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 40.899 | 41 | 0 | 0.101 | 3.317 | 3.338 | 0 | 0.021
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 29.768 | 29.983 | 0 | 0.215 | 3.384 | 3.424 | 0 | 0.04
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 20.769 | 19.994 | 0 | -0.775 | 3.434 | 3.523 | 0 | 0.089
2022-070-0057.mrg | 2022-070-0057.thn1 | 2 | 10.62 | 10.4955 | 0.71347 | -0.1245 | 4.087 | 4.092 | 0.14991 | 0.0050001
2022-070-0057.mrg | 2022-070-0057.thn1 | 2 | 5.521 | 5.4975 | 0.70781 | -0.023499 | 5.974 | 6.1545 | 0.35143 | 0.1805
2022-070-0057.mrg | 2022-070-0057.thn1 | 1 | 2.383 | 2.704 | 0 | 0.321 | 6.503 | 6.702 | 0 | 0.199
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 340.794 | 339.954 | 0 | -0.84 | 2.626 | 2.631 | 0 | 0.0050001
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 300.218 | 300.006 | 0 | -0.21197 | 2.665 | 2.675 | 0 | 0.01
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 250.399 | 249.986 | 0 | -0.41301 | 2.707 | 2.717 | 0 | 0.01
2022-070-0064.mrg | 2022-070-0064.thn1 | 2 | 200.705 | 200.488 | 0.70569 | -0.21699 | 2.746 | 2.7675 | 0.0049497 | 0.0215
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 175.387 | 174.996 | 0 | -0.39099 | 2.783 | 2.822 | 0 | 0.039
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 150.326 | 150.006 | 0 | -0.32001 | 2.797 | 2.824 | 0 | 0.027
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 125.384 | 124.991 | 0 | -0.393 | 2.825 | 2.836 | 0 | 0.011
2022-070-0064.mrg | 2022-070-0064.thn1 | 2 | 100.407 | 100.494 | 0.7163 | 0.087501 | 2.825 | 2.8485 | 0.0021213 | 0.0235
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 75.601 | 74.992 | 0 | -0.609 | 2.925 | 2.953 | 0 | 0.028
2022-070-0064.mrg | 2022-070-0064.thn1 | 2 | 50.197 | 50.503 | 0.70004 | 0.306 | 3.225 | 3.266 | 0.022627 | 0.041
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 40.489 | 40.974 | 0 | 0.485 | 3.211 | 3.26 | 0 | 0.049
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 30.794 | 29.983 | 0 | -0.811 | 3.229 | 3.3 | 0 | 0.071
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 20.494 | 19.992 | 0 | -0.502 | 3.456 | 3.499 | 0 | 0.043
2022-070-0064.mrg | 2022-070-0064.thn1 | 2 | 10.481 | 10.494 | 0.70145 | 0.013 | 3.636 | 3.8205 | 0.078489 | 0.1845
2022-070-0064.mrg | 2022-070-0064.thn1 | 2 | 5.238 | 5.488 | 0.70004 | 0.25 | 5.318 | 4.358 | 0.25314 | -0.96
2022-070-0064.mrg | 2022-070-0064.thn1 | 1 | 2.15 | 2.717 | 0 | 0.567 | 6.262 | 6.252 | 0 | -0.01
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 275.044 | 274.923 | 0 | -0.121 | 2.675 | 2.702 | 0 | 0.027
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 250.399 | 250 | 0 | -0.399 | 2.71 | 2.734 | 0 | 0.024
2022-070-0071.mrg | 2022-070-0071.thn1 | 2 | 200.164 | 200.508 | 0.70499 | 0.3445 | 2.805 | 2.828 | 0.011314 | 0.023
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 175.472 | 175.005 | 0 | -0.46699 | 2.826 | 2.845 | 0 | 0.019
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 150.102 | 150.006 | 0 | -0.096008 | 2.853 | 2.889 | 0 | 0.036
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 125.064 | 125.01 | 0 | -0.054001 | 2.816 | 2.855 | 0 | 0.039
2022-070-0071.mrg | 2022-070-0071.thn1 | 2 | 100.247 | 100.492 | 0.71135 | 0.245 | 2.842 | 2.865 | 0.0070711 | 0.023
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 75.342 | 74.989 | 0 | -0.353 | 2.913 | 2.938 | 0 | 0.025
2022-070-0071.mrg | 2022-070-0071.thn1 | 2 | 50.204 | 50.4885 | 0.71206 | 0.2845 | 3.178 | 3.227 | 0.0042427 | 0.049
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 40.169 | 40.989 | 0 | 0.82 | 3.255 | 3.329 | 0 | 0.074
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 30.147 | 30.017 | 0 | -0.13 | 3.391 | 3.445 | 0 | 0.054
2022-070-0071.mrg | 2022-070-0071.thn1 | 1 | 20.192 | 20.011 | 0 | -0.181 | 3.661 | 3.699 | 0 | 0.038
2022-070-0071.mrg | 2022-070-0071.thn1 | 2 | 10.216 | 10.5005 | 0.6965 | 0.2845 | 4.772 | 5.176 | 0.056569 | 0.404
2022-070-0071.mrg | 2022-070-0071.thn1 | 2 | 5.133 | 5.4955 | 0.69791 | 0.3625 | 6.121 | 5.942 | 0.34365 | -0.179
2022-070-0071.mrg | 2022-070-0071.thn1 | 0 | 1.553 | 0 | 0 | -1.553 | 6.545 | 0 | 0 | -6.545
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 318.757 | 318.774 | 0 | 0.016998 | 2.581 | 2.591 | 0 | 0.01
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 300.333 | 300.012 | 0 | -0.32101 | 2.635 | 2.643 | 0 | 0.0079999
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 250.797 | 250.023 | 0 | -0.774 | 2.716 | 2.719 | 0 | 0.003
2022-070-0075.mrg | 2022-070-0075.thn1 | 2 | 200.391 | 200.496 | 0.70993 | 0.105 | 2.798 | 2.833 | 0.0014143 | 0.035
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 175.481 | 174.999 | 0 | -0.48201 | 2.847 | 2.854 | 0 | 0.0070002
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 150.304 | 150.009 | 0 | -0.295 | 2.863 | 2.877 | 0 | 0.014
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 125.313 | 125.003 | 0 | -0.31 | 2.854 | 2.892 | 0 | 0.038
2022-070-0075.mrg | 2022-070-0075.thn1 | 2 | 100.4 | 100.495 | 0.69084 | 0.095497 | 2.856 | 2.889 | 0.0028284 | 0.033
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 75.391 | 75.001 | 0 | -0.39 | 2.909 | 2.957 | 0 | 0.048
2022-070-0075.mrg | 2022-070-0075.thn1 | 2 | 50.415 | 50.496 | 0.70569 | 0.081001 | 3.128 | 3.185 | 0.011314 | 0.057
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 40.342 | 41.014 | 0 | 0.672 | 3.311 | 3.371 | 0 | 0.06
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 30.397 | 29.999 | 0 | -0.398 | 3.427 | 3.507 | 0 | 0.08
2022-070-0075.mrg | 2022-070-0075.thn1 | 1 | 20.658 | 19.987 | 0 | -0.671 | 3.538 | 3.592 | 0 | 0.054
2022-070-0075.mrg | 2022-070-0075.thn1 | 2 | 10.791 | 10.498 | 0.70852 | -0.293 | 3.765 | 4.1565 | 0.1393 | 0.3915
2022-070-0075.mrg | 2022-070-0075.thn1 | 2 | 5.191 | 5.502 | 0.72408 | 0.311 | 6.396 | 6.281 | 0.31395 | -0.115
2022-070-0075.mrg | 2022-070-0075.thn1 | 0 | 2.058 | 0 | 0 | -2.058 | 6.5 | 0 | 0 | -6.5
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 188.23 | 188.992 | 0 | 0.76201 | 2.769 | 2.788 | 0 | 0.019
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 175.471 | 174.999 | 0 | -0.472 | 2.775 | 2.802 | 0 | 0.027
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 150.512 | 149.997 | 0 | -0.515 | 2.849 | 2.891 | 0 | 0.042
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 125.584 | 125.004 | 0 | -0.58 | 2.876 | 2.897 | 0 | 0.021
2022-070-0079.mrg | 2022-070-0079.thn1 | 2 | 100.303 | 100.494 | 0.71418 | 0.191 | 2.85 | 2.8815 | 0.0063638 | 0.0315
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 75.459 | 74.992 | 0 | -0.467 | 2.874 | 2.944 | 0 | 0.07
2022-070-0079.mrg | 2022-070-0079.thn1 | 2 | 50.686 | 50.5005 | 0.7064 | -0.1855 | 3.211 | 3.286 | 0.012728 | 0.075
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 40.254 | 41.007 | 0 | 0.753 | 3.307 | 3.389 | 0 | 0.082
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 30.624 | 30.001 | 0 | -0.623 | 3.404 | 3.442 | 0 | 0.038
2022-070-0079.mrg | 2022-070-0079.thn1 | 1 | 20.348 | 19.991 | 0 | -0.357 | 3.4 | 3.669 | 0 | 0.269
2022-070-0079.mrg | 2022-070-0079.thn1 | 2 | 10.349 | 10.492 | 0.70428 | 0.143 | 4.404 | 4.7695 | 0.04879 | 0.3655
2022-070-0079.mrg | 2022-070-0079.thn1 | 2 | 5.412 | 5.5885 | 0.57346 | 0.1765 | 5.907 | 5.6055 | 0.4278 | -0.3015
2022-070-0079.mrg | 2022-070-0079.thn1 | 0 | 1.916 | 0 | 0 | -1.916 | 6.728 | 0 | 0 | -6.728
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 163.195 | 162.759 | 0 | -0.436 | 2.473 | 2.497 | 0 | 0.024
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 150.202 | 149.998 | 0 | -0.204 | 2.485 | 2.527 | 0 | 0.042
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 125.332 | 124.998 | 0 | -0.334 | 2.804 | 2.829 | 0 | 0.025
2022-070-0082.mrg | 2022-070-0082.thn1 | 2 | 100.149 | 100.49 | 0.71064 | 0.34149 | 2.81 | 2.8525 | 0.00070706 | 0.0425
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 75.19 | 74.983 | 0 | -0.207 | 2.979 | 3.019 | 0 | 0.04
2022-070-0082.mrg | 2022-070-0082.thn1 | 2 | 50.122 | 50.507 | 0.70287 | 0.385 | 3.309 | 3.3895 | 0.016264 | 0.0805
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 40.114 | 41.01 | 0 | 0.896 | 3.432 | 3.516 | 0 | 0.084
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 29.985 | 29.993 | 0 | 0.0079994 | 3.508 | 3.606 | 0 | 0.098
2022-070-0082.mrg | 2022-070-0082.thn1 | 1 | 20.247 | 19.996 | 0 | -0.251 | 4.096 | 4.116 | 0 | 0.02
2022-070-0082.mrg | 2022-070-0082.thn1 | 2 | 10.081 | 10.501 | 0.72408 | 0.42 | 5.772 | 5.843 | 0.14708 | 0.071
2022-070-0082.mrg | 2022-070-0082.thn1 | 2 | 4.93 | 4.4875 | 0.70498 | -0.4425 | 7.463 | 7.1605 | 0.16476 | -0.3025
2022-070-0082.mrg | 2022-070-0082.thn1 | 0 | 1.937 | 0 | 0 | -1.937 | 7.117 | 0 | 0 | -7.117
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 351.453 | 351.925 | 0 | 0.47198 | 2.71 | 2.715 | 0 | 0.0049999
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 300.534 | 300.001 | 0 | -0.53299 | 2.711 | 2.753 | 0 | 0.042
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 250.223 | 249.991 | 0 | -0.23201 | 2.762 | 2.787 | 0 | 0.025
2022-070-0086.mrg | 2022-070-0086.thn1 | 2 | 200.152 | 200.487 | 0.69367 | 0.33549 | 2.817 | 2.8455 | 0.0049497 | 0.0285
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 175.15 | 174.995 | 0 | -0.155 | 2.838 | 2.86 | 0 | 0.022
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 150.038 | 150.001 | 0 | -0.036987 | 2.868 | 2.89 | 0 | 0.022
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 125.037 | 124.987 | 0 | -0.050003 | 2.867 | 2.901 | 0 | 0.034
2022-070-0086.mrg | 2022-070-0086.thn1 | 2 | 100.003 | 100.506 | 0.70145 | 0.503 | 2.838 | 2.875 | 0.0084853 | 0.037
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 75.033 | 75.002 | 0 | -0.030998 | 2.868 | 2.93 | 0 | 0.062
2022-070-0086.mrg | 2022-070-0086.thn1 | 2 | 50.227 | 50.5015 | 0.70216 | 0.2745 | 3.154 | 3.2185 | 0.0091924 | 0.0645
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 40.011 | 41.001 | 0 | 0.99 | 3.376 | 3.501 | 0 | 0.125
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 30.057 | 30.003 | 0 | -0.053999 | 3.536 | 3.598 | 0 | 0.062
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 20.133 | 19.99 | 0 | -0.143 | 3.658 | 3.945 | 0 | 0.287
2022-070-0086.mrg | 2022-070-0086.thn1 | 2 | 10.28 | 10.498 | 0.70711 | 0.218 | 4.345 | 4.91 | 0.16122 | 0.565
2022-070-0086.mrg | 2022-070-0086.thn1 | 2 | 5.223 | 5.504 | 0.70711 | 0.281 | 5.966 | 6.632 | 0.11738 | 0.666
2022-070-0086.mrg | 2022-070-0086.thn1 | 1 | 2.052 | 2.655 | 0 | 0.603 | 6.752 | 6.771 | 0 | 0.019
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 312.235 | 312.258 | 0 | 0.02301 | 2.735 | 2.748 | 0 | 0.013
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 300.217 | 299.998 | 0 | -0.21902 | 2.744 | 2.767 | 0 | 0.023
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 250.35 | 250.002 | 0 | -0.34801 | 2.798 | 2.822 | 0 | 0.024
2022-070-0089.mrg | 2022-070-0089.thn1 | 2 | 200.781 | 200.509 | 0.70711 | -0.272 | 2.865 | 2.8765 | 0.00070722 | 0.0115
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 175.885 | 175.009 | 0 | -0.87599 | 2.869 | 2.896 | 0 | 0.027
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 149.895 | 150.004 | 0 | 0.10899 | 2.876 | 2.908 | 0 | 0.032
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 125.084 | 124.986 | 0 | -0.098 | 2.797 | 2.831 | 0 | 0.034
2022-070-0089.mrg | 2022-070-0089.thn1 | 2 | 100.131 | 100.499 | 0.69933 | 0.3685 | 3.013 | 3.131 | 0.06364 | 0.118
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 75.066 | 75.003 | 0 | -0.063004 | 3.474 | 3.535 | 0 | 0.061
2022-070-0089.mrg | 2022-070-0089.thn1 | 2 | 50.073 | 50.4985 | 0.71489 | 0.4255 | 3.772 | 3.869 | 0.0098995 | 0.097
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 40.015 | 40.982 | 0 | 0.967 | 3.807 | 3.85 | 0 | 0.043
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 29.956 | 29.991 | 0 | 0.035 | 3.867 | 4.062 | 0 | 0.195
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 20.049 | 19.999 | 0 | -0.049999 | 4.022 | 4.126 | 0 | 0.104
2022-070-0089.mrg | 2022-070-0089.thn1 | 2 | 10.205 | 10.503 | 0.71135 | 0.298 | 4.313 | 5.2705 | 0.39386 | 0.9575
2022-070-0089.mrg | 2022-070-0089.thn1 | 2 | 5.179 | 5.492 | 0.70428 | 0.313 | 6.382 | 6.44 | 0.15132 | 0.058
2022-070-0089.mrg | 2022-070-0089.thn1 | 1 | 1.816 | 2.709 | 0 | 0.893 | 6.692 | 6.735 | 0 | 0.043
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 151.765 | 151.799 | 0 | 0.033997 | 2.813 | 2.835 | 0 | 0.022
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 125.924 | 125.011 | 0 | -0.913 | 2.836 | 2.858 | 0 | 0.022
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 101.096 | 100.985 | 0 | -0.111 | 2.963 | 2.892 | 0 | -0.071
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 75.397 | 75.004 | 0 | -0.393 | 3.193 | 3.32 | 0 | 0.127
2022-070-0093.mrg | 2022-070-0093.thn1 | 2 | 50.582 | 50.494 | 0.70003 | -0.088001 | 3.751 | 3.788 | 0.0028284 | 0.037
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 40.649 | 40.991 | 0 | 0.342 | 3.79 | 3.805 | 0 | 0.015
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 30.749 | 29.993 | 0 | -0.756 | 3.829 | 3.911 | 0 | 0.082
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 20.021 | 20 | 0 | -0.021 | 3.983 | 4.229 | 0 | 0.246
2022-070-0093.mrg | 2022-070-0093.thn1 | 2 | 10.697 | 10.5025 | 0.69226 | -0.1945 | 4.209 | 4.6065 | 0.012021 | 0.3975
2022-070-0093.mrg | 2022-070-0093.thn1 | 2 | 5.509 | 5.496 | 0.70852 | -0.013 | 4.552 | 4.6605 | 0.017678 | 0.1085
2022-070-0093.mrg | 2022-070-0093.thn1 | 1 | 2.037 | 2.506 | 0 | 0.469 | 4.654 | 4.74 | 0 | 0.086
Create Fit Sheet
Plot
Plot

## File List
Number of Files= |  | 21
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2022-070-0001.mrg | 2022-070-0001.thn1 | 11 | 22 | 2022-070-0001.mrg & 2022-070-0001.thn1
2 | 2022-070-0005.mrg | 2022-070-0005.thn1 | 24 | 38 | 2022-070-0005.mrg & 2022-070-0005.thn1
3 | 2022-070-0008.mrg | 2022-070-0008.thn1 | 40 | 53 | 2022-070-0008.mrg & 2022-070-0008.thn1
4 | 2022-070-0012.mrg | 2022-070-0012.thn1 | 55 | 67 | 2022-070-0012.mrg & 2022-070-0012.thn1
5 | 2022-070-0015.mrg | 2022-070-0015.thn1 | 69 | 80 | 2022-070-0015.mrg & 2022-070-0015.thn1
6 | 2022-070-0018.mrg | 2022-070-0018.thn1 | 82 | 91 | 2022-070-0018.mrg & 2022-070-0018.thn1
7 | 2022-070-0022.mrg | 2022-070-0022.thn1 | 93 | 102 | 2022-070-0022.mrg & 2022-070-0022.thn1
8 | 2022-070-0026.mrg | 2022-070-0026.thn1 | 104 | 114 | 2022-070-0026.mrg & 2022-070-0026.thn1
9 | 2022-070-0031.mrg | 2022-070-0031.thn1 | 116 | 129 | 2022-070-0031.mrg & 2022-070-0031.thn1
10 | 2022-070-0036.mrg | 2022-070-0036.thn1 | 131 | 144 | 2022-070-0036.mrg & 2022-070-0036.thn1
11 | 2022-070-0048.mrg | 2022-070-0048.thn1 | 146 | 157 | 2022-070-0048.mrg & 2022-070-0048.thn1
12 | 2022-070-0052.mrg | 2022-070-0052.thn1 | 159 | 174 | 2022-070-0052.mrg & 2022-070-0052.thn1
13 | 2022-070-0057.mrg | 2022-070-0057.thn1 | 176 | 192 | 2022-070-0057.mrg & 2022-070-0057.thn1
14 | 2022-070-0064.mrg | 2022-070-0064.thn1 | 194 | 209 | 2022-070-0064.mrg & 2022-070-0064.thn1
15 | 2022-070-0071.mrg | 2022-070-0071.thn1 | 211 | 225 | 2022-070-0071.mrg & 2022-070-0071.thn1
16 | 2022-070-0075.mrg | 2022-070-0075.thn1 | 227 | 242 | 2022-070-0075.mrg & 2022-070-0075.thn1
17 | 2022-070-0079.mrg | 2022-070-0079.thn1 | 244 | 256 | 2022-070-0079.mrg & 2022-070-0079.thn1
18 | 2022-070-0082.mrg | 2022-070-0082.thn1 | 258 | 269 | 2022-070-0082.mrg & 2022-070-0082.thn1
19 | 2022-070-0086.mrg | 2022-070-0086.thn1 | 271 | 286 | 2022-070-0086.mrg & 2022-070-0086.thn1
20 | 2022-070-0089.mrg | 2022-070-0089.thn1 | 288 | 303 | 2022-070-0089.mrg & 2022-070-0089.thn1
21 | 2022-070-0093.mrg | 2022-070-0093.thn1 | 305 | 315 | 2022-070-0093.mrg & 2022-070-0093.thn1

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (OXY_SBE - OXY_BOT) | OXY_SBE | 21
2 | (OXY_SBE - OXY_BOT) | Pressure | 21
3 | (OXY_SBE - OXY_BOT) | File Pair # | 21
4 | (OXY_SBE - OXY_BOT) | OXY_SBE | 3

## Fit_1
Calculation of Linear Least Squares Fit For
(OXY_SBE - OXY_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (OXY_SBE - OXY_BOT), Prop=OXY_SBE
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.10332 |  | X | CALC Y
m= | 0.042608 | 0.0078449 |  | Std Dev of Resid |  | 0.12388 |  | 0.0485 | -0.070054
b= | -0.07212 | 0.025403 |  | Resid Sum of Squares |  | 3.9288 |  | 6.938 | 0.22349
# of Values Extracted = |  |  | 285
# of Values Used in Fit = |  |  | 258
File | Use | Pressure | OXY_SBE | OXY_SBE | Diff (OXY_SBE - OXY_BOT) |  |  |  | Fit vs. | OXY_SBE
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 173.451 | 0.0485 | 0.00070711 | 0.0045 | -0.070054 | -0.074554 |  | 0.0485 | 0.0045
1 | YES | 150.001 | 0.053 | 0 | 0.013 | -0.069862 | -0.082862 |  | 0.053 | 0.013
1 | YES | 125.01 | 0.241 | 0 | 0.099 | -0.061852 | -0.16085 |  | 0.241 | 0.099
1 | YES | 100.504 | 1.0745 | 0.013435 | 0.2115 | -0.026338 | -0.23784 |  | 1.0745 | 0.2115
1 | YES | 75 | 0.63 | 0 | -0.175 | -0.045277 | 0.12972 |  | 0.63 | -0.175
1 | YES | 50.4965 | 2.851 | 0.05374 | 0.118 | 0.049355 | -0.068645 |  | 2.851 | 0.118
1 | YES | 41.004 | 3.079 | 0 | 0.092 | 0.05907 | -0.03293 |  | 3.079 | 0.092
1 | YES | 30.003 | 3.281 | 0 | 0.14 | 0.067676 | -0.072324 |  | 3.281 | 0.14
1 | YES | 19.996 | 3.394 | 0 | 0.145 | 0.072491 | -0.072509 |  | 3.394 | 0.145
1 | NO | 4.5095 | 4.1225 | 0.16193 | 0.3435 | 0.10353 | -0.23997 |  | 6.125 | 0.429
1 | NO | 0 | 0 | 0 | -3.918 | -0.07212 | 3.8459 |  | 1.2085 | 0.0084999
1 | YES | 2.477 | 6.125 | 0 | 0.429 | 0.18885 | -0.24015 |  | 1.213 | 0.01
2 | YES | 252.453 | 1.2085 | 0.0035355 | 0.0084999 | -0.020629 | -0.029129 |  | 1.2595 | 0.0055001
2 | YES | 249.989 | 1.213 | 0 | 0.01 | -0.020437 | -0.030437 |  | 1.3 | -0.018
2 | YES | 200.506 | 1.2595 | 0.0035355 | 0.0055001 | -0.018456 | -0.023956 |  | 1.55 | -0.104
2 | YES | 175.221 | 1.3 | 0 | -0.018 | -0.01673 | 0.00127 |  | 2.066 | -0.068
2 | YES | 149.987 | 1.55 | 0 | -0.104 | -0.0060781 | 0.097922 |  | 2.145 | 0.049
2 | YES | 125.003 | 2.066 | 0 | -0.068 | 0.015908 | 0.083908 |  | 2.421 | 0.102
2 | YES | 101.002 | 2.145 | 0 | 0.049 | 0.019274 | -0.029726 |  | 2.8605 | 0.0345
2 | YES | 74.994 | 2.421 | 0 | 0.102 | 0.031033 | -0.070967 |  | 3.277 | 0.322
2 | YES | 50.477 | 2.8605 | 0.00070722 | 0.0345 | 0.04976 | 0.01526 |  | 3.267 | 0.064
2 | YES | 41.036 | 3.277 | 0 | 0.322 | 0.067506 | -0.25449 |  | 3.797 | -0.147
2 | YES | 30.012 | 3.267 | 0 | 0.064 | 0.06708 | 0.0030799 |  | 4.1235 | -0.5165
2 | YES | 19.846 | 3.797 | 0 | -0.147 | 0.089662 | 0.23666 |  | 4.711 | -0.074
2 | YES | 10.491 | 4.1235 | 0.027577 | -0.5165 | 0.10357 | 0.62007 |  | 1.242 | 0.015
2 | NO | 5.4895 | 4.543 | 0.10748 | -0.227 | 0.12145 | 0.34845 |  | 1.244 | 0.015
2 | YES | 3.985 | 4.711 | 0 | -0.074 | 0.12861 | 0.20261 |  | 1.305 | 0.04
3 | YES | 219.979 | 1.242 | 0 | 0.015 | -0.019201 | -0.034201 |  | 1.592 | 0.066
3 | YES | 202.99 | 1.244 | 0 | 0.015 | -0.019116 | -0.034116 |  | 1.839 | 0.079
3 | YES | 177.016 | 1.305 | 0 | 0.04 | -0.016517 | -0.056517 |  | 2.2775 | 0.4415
3 | YES | 150.001 | 1.592 | 0 | 0.066 | -0.0042885 | -0.070289 |  | 2.313 | 0.214
3 | YES | 124.984 | 1.839 | 0 | 0.079 | 0.0062357 | -0.072764 |  | 2.689 | 0.296
3 | YES | 100.496 | 2.2775 | 0.044548 | 0.4415 | 0.024919 | -0.41658 |  | 3.167 | 0.388
3 | YES | 77.016 | 2.313 | 0 | 0.214 | 0.026432 | -0.18757 |  | 3.523 | 0.294
3 | YES | 51.988 | 2.689 | 0 | 0.296 | 0.042452 | -0.25355 |  | 4.188 | 0.452
3 | YES | 41.027 | 3.167 | 0 | 0.388 | 0.062819 | -0.32518 |  | 4.277 | 0.0040002
3 | YES | 30.003 | 3.523 | 0 | 0.294 | 0.077987 | -0.21601 |  | 4.2675 | -0.0155
3 | YES | 19.991 | 4.188 | 0 | 0.452 | 0.10632 | -0.34568 |  | 4.961 | 0.233
3 | YES | 10.004 | 4.277 | 0 | 0.0040002 | 0.11011 | 0.10611 |  | 1.189 | 0.023
3 | YES | 5.4935 | 4.2675 | 0.023335 | -0.0155 | 0.10971 | 0.12521 |  | 1.197 | 0.023
3 | YES | 2.595 | 4.961 | 0 | 0.233 | 0.13926 | -0.093742 |  | 1.201 | 0.027
4 | YES | 185.396 | 1.189 | 0.0014143 | 0.023 | -0.02146 | -0.04446 |  | 1.211 | 0.027
4 | YES | 174.996 | 1.197 | 0 | 0.023 | -0.021119 | -0.044119 |  | 1.521 | 0.0040001
4 | YES | 150.001 | 1.201 | 0 | 0.027 | -0.020948 | -0.047948 |  | 1.846 | -0.024
4 | YES | 125.011 | 1.211 | 0 | 0.027 | -0.020522 | -0.047522 |  | 2.702 | -0.099
4 | YES | 100.497 | 1.521 | 0.0070711 | 0.0040001 | -0.0073137 | -0.011314 |  | 2.919 | 0.017
4 | YES | 76.991 | 1.846 | 0 | -0.024 | 0.0065339 | 0.030534 |  | 3.096 | -0.06
4 | YES | 50.506 | 2.702 | 0.0042427 | -0.099 | 0.043006 | 0.14201 |  | 3.443 | -0.037
4 | YES | 41.009 | 2.919 | 0 | 0.017 | 0.052252 | 0.035252 |  | 4.093 | 0.15
4 | YES | 29.983 | 3.096 | 0 | -0.06 | 0.059794 | 0.11979 |  | 4.0935 | 0.2525
4 | YES | 19.989 | 3.443 | 0 | -0.037 | 0.074579 | 0.11158 |  | 4.083 | 0.134
4 | YES | 10.4915 | 4.093 | 0.0014141 | 0.15 | 0.10227 | -0.047726 |  | 1.2195 | 0.0185
4 | YES | 5.5005 | 4.0935 | 0.00070706 | 0.2525 | 0.1023 | -0.1502 |  | 1.291 | 0.039
4 | YES | 2.693 | 4.083 | 0 | 0.134 | 0.10185 | -0.032152 |  | 1.719 | -0.036
5 | YES | 175.173 | 1.2195 | 0.0035355 | 0.0185 | -0.02016 | -0.03866 |  | 2.327 | -0.201
5 | NO | 0 | 0 | 0 | -1.213 | -0.07212 | 1.1409 |  | 2.797 | 0.00099993
5 | YES | 124.991 | 1.291 | 0 | 0.039 | -0.017114 | -0.056114 |  | 2.902 | 0.026
5 | YES | 100.495 | 1.719 | 0.041012 | -0.036 | 0.0011227 | 0.037123 |  | 2.989 | 0.035
5 | YES | 75.006 | 2.327 | 0 | -0.201 | 0.027028 | 0.22803 |  | 3.048 | 0.05
5 | YES | 49.502 | 2.797 | 0.0070711 | 0.00099993 | 0.047054 | 0.046054 |  | 3.154 | 0.036
5 | YES | 40.995 | 2.902 | 0 | 0.026 | 0.051528 | 0.025528 |  | 3.2195 | -0.0095
5 | YES | 30.001 | 2.989 | 0 | 0.035 | 0.055235 | 0.020235 |  | 3.282 | 0.007
5 | YES | 19.988 | 3.048 | 0 | 0.05 | 0.057749 | 0.0077485 |  | 1.594 | 0
5 | YES | 10.5055 | 3.154 | 0.0014141 | 0.036 | 0.062265 | 0.026265 |  | 1.82 | -0.023
5 | YES | 5.5105 | 3.2195 | 0.026163 | -0.0095 | 0.065056 | 0.074556 |  | 2.072 | 0.035
5 | YES | 3.129 | 3.282 | 0 | 0.007 | 0.067719 | 0.060719 |  | 2.8865 | 0.0435
6 | YES | 114 | 1.594 | 0 | 0 | -0.0042033 | -0.0042033 |  | 3.314 | 0.138
6 | YES | 100.998 | 1.82 | 0 | -0.023 | 0.0054261 | 0.028426 |  | 3.462 | 0.078
6 | YES | 74.991 | 2.072 | 0 | 0.035 | 0.016163 | -0.018837 |  | 3.885 | 0.465
6 | YES | 50.4905 | 2.8865 | 0.014849 | 0.0435 | 0.050868 | 0.0073676 |  | 3.9585 | 0.2165
6 | YES | 40.992 | 3.314 | 0 | 0.138 | 0.069082 | -0.068918 |  | 4.034 | 0.082
6 | YES | 29.985 | 3.462 | 0 | 0.078 | 0.075388 | -0.0026114 |  | 4.047 | 0.037
6 | YES | 19.999 | 3.885 | 0 | 0.465 | 0.093412 | -0.37159 |  | 1.8955 | 0.0024999
6 | YES | 10.5035 | 3.9585 | 0.02192 | 0.2165 | 0.096543 | -0.11996 |  | 2.2795 | 0.0205
6 | YES | 5.504 | 4.034 | 0.0070712 | 0.082 | 0.09976 | 0.01776 |  | 2.691 | -0.071
6 | YES | 2.212 | 4.047 | 0 | 0.037 | 0.10031 | 0.063314 |  | 3.0885 | 0.0595
7 | YES | 123.143 | 1.8955 | 0.0035355 | 0.0024999 | 0.008643 | 0.0061431 |  | 3.284 | 0.151
7 | YES | 100.495 | 2.2795 | 0.0021212 | 0.0205 | 0.025004 | 0.0045046 |  | 3.391 | 0.239
7 | YES | 75.01 | 2.691 | 0 | -0.071 | 0.042538 | 0.11354 |  | 3.692 | 0.345
7 | YES | 50.5045 | 3.0885 | 0.0035356 | 0.0595 | 0.059474 | -2.5679e-05 |  | 3.731 | -0.003
7 | YES | 40.999 | 3.284 | 0 | 0.151 | 0.067804 | -0.083196 |  | 3.7395 | -0.0195
7 | YES | 29.988 | 3.391 | 0 | 0.239 | 0.072363 | -0.16664 |  | 3.792 | 0.004
7 | YES | 19.998 | 3.692 | 0 | 0.345 | 0.085188 | -0.25981 |  | 2.058 | -0.017
7 | YES | 10.5015 | 3.731 | 0.0098995 | -0.003 | 0.08685 | 0.08985 |  | 2.187 | 0.057
7 | YES | 5.494 | 3.7395 | 0.0035356 | -0.0195 | 0.087212 | 0.10671 |  | 2.678 | -0.023
7 | YES | 2.708 | 3.792 | 0 | 0.004 | 0.089449 | 0.085449 |  | 2.997 | 0.044
8 | YES | 145.194 | 2.058 | 0.012728 | -0.017 | 0.015567 | 0.032567 |  | 3.2705 | 0.0785
8 | YES | 124.991 | 2.187 | 0 | 0.057 | 0.021063 | -0.035937 |  | 3.369 | 0.024
8 | YES | 100.995 | 2.678 | 0 | -0.023 | 0.041984 | 0.064984 |  | 3.404 | 0.034
8 | YES | 75.007 | 2.997 | 0 | 0.044 | 0.055576 | 0.011576 |  | 3.471 | 0.101
8 | YES | 50.4985 | 3.2705 | 0.028991 | 0.0785 | 0.067229 | -0.011271 |  | 3.465 | 0.051
8 | YES | 41.004 | 3.369 | 0 | 0.024 | 0.071426 | 0.047426 |  | 3.485 | 0.04
8 | YES | 29.996 | 3.404 | 0 | 0.034 | 0.072917 | 0.038917 |  | 3.543 | 0.078
8 | YES | 19.988 | 3.471 | 0 | 0.101 | 0.075772 | -0.025228 |  | 2.662 | 0
8 | YES | 10.512 | 3.465 | 0.0028284 | 0.051 | 0.075516 | 0.024516 |  | 2.6645 | -0.0025001
8 | YES | 5.5 | 3.485 | 0.0056569 | 0.04 | 0.076368 | 0.036368 |  | 2.687 | 0.016
8 | YES | 2.202 | 3.543 | 0 | 0.078 | 0.07884 | 0.00083955 |  | 2.71 | 0.018
9 | YES | 218.883 | 2.662 | 0 | 0 | 0.041302 | 0.041302 |  | 2.799 | 0.051
9 | YES | 200.508 | 2.6645 | 0.00070706 | -0.0025001 | 0.041409 | 0.043909 |  | 3.01 | 0.0049999
9 | YES | 175.002 | 2.687 | 0 | 0.016 | 0.042367 | 0.026367 |  | 3.255 | 0.103
9 | YES | 150.011 | 2.71 | 0 | 0.018 | 0.043347 | 0.025347 |  | 3.36 | 2.3842e-07
9 | YES | 125.005 | 2.799 | 0 | 0.051 | 0.047139 | -0.0038608 |  | 3.376 | -0.044
9 | YES | 101.009 | 3.01 | 0 | 0.0049999 | 0.05613 | 0.05113 |  | 3.516 | 0.093
9 | YES | 77.007 | 3.255 | 0 | 0.103 | 0.066569 | -0.036431 |  | 3.699 | 0.222
9 | YES | 50.495 | 3.36 | 0.0028284 | 2.3842e-07 | 0.071042 | 0.071042 |  | 3.7645 | -0.0785
9 | YES | 41.01 | 3.376 | 0 | -0.044 | 0.071724 | 0.11572 |  | 5.394 | 1.237
9 | YES | 29.993 | 3.516 | 0 | 0.093 | 0.077689 | -0.015311 |  | 2.715 | 0.026
9 | YES | 22 | 3.699 | 0 | 0.222 | 0.085486 | -0.13651 |  | 2.7145 | 0.0315
9 | YES | 11.4935 | 3.7645 | 0.030406 | -0.0785 | 0.088277 | 0.16678 |  | 2.747 | 0.029
9 | NO | 5.4995 | 4.353 | 0.13435 | 0.318 | 0.11335 | -0.20465 |  | 2.806 | -0.04
9 | YES | 2.748 | 5.394 | 0 | 1.237 | 0.15771 | -1.0793 |  | 3.062 | 0.041
10 | YES | 226.669 | 2.715 | 0.0056569 | 0.026 | 0.04356 | 0.01756 |  | 3.1095 | 0.0145
10 | YES | 200.49 | 2.7145 | 0.006364 | 0.0315 | 0.043539 | 0.012039 |  | 3.205 | -0.079
10 | YES | 176.984 | 2.747 | 0 | 0.029 | 0.044924 | 0.015924 |  | 3.39 | 0.034
10 | YES | 152.993 | 2.806 | 0 | -0.04 | 0.047438 | 0.087438 |  | 3.438 | 0.062
10 | YES | 124.993 | 3.062 | 0 | 0.041 | 0.058345 | 0.017345 |  | 3.506 | 0.108
10 | YES | 100.496 | 3.1095 | 0.006364 | 0.0145 | 0.060369 | 0.045869 |  | 3.584 | 0.106
10 | YES | 75.007 | 3.205 | 0 | -0.079 | 0.064438 | 0.14344 |  | 3.885 | -0.094
10 | YES | 51.509 | 3.39 | 0.0084852 | 0.034 | 0.072321 | 0.038321 |  | 3.9215 | -0.1545
10 | YES | 40.997 | 3.438 | 0 | 0.062 | 0.074366 | 0.012366 |  | 3.974 | -0.218
10 | YES | 29.982 | 3.506 | 0 | 0.108 | 0.077263 | -0.030737 |  | 2.808 | 0.027
10 | YES | 19.989 | 3.584 | 0 | 0.106 | 0.080587 | -0.025413 |  | 2.942 | 0.038
10 | YES | 10.501 | 3.885 | 0.014142 | -0.094 | 0.093412 | 0.18741 |  | 2.915 | 0.003
10 | YES | 5.493 | 3.9215 | 0.0021213 | -0.1545 | 0.094967 | 0.24947 |  | 3.1775 | 0.0415
10 | YES | 2.71 | 3.974 | 0 | -0.218 | 0.097204 | 0.3152 |  | 3.298 | 0.028
11 | YES | 171.333 | 2.808 | 0.0084852 | 0.027 | 0.047523 | 0.020523 |  | 3.3765 | 0.0605
11 | YES | 149.986 | 2.942 | 0 | 0.038 | 0.053232 | 0.015232 |  | 3.446 | 0.058
11 | YES | 125.003 | 2.915 | 0 | 0.003 | 0.052082 | 0.049082 |  | 3.526 | 0.007
11 | YES | 100.512 | 3.1775 | 0.0049497 | 0.0415 | 0.063266 | 0.021766 |  | 3.7 | 0.059
11 | YES | 75.002 | 3.298 | 0 | 0.028 | 0.068401 | 0.040401 |  | 2.595 | 0.01
11 | YES | 50.4985 | 3.3765 | 0.019092 | 0.0605 | 0.071745 | 0.011245 |  | 2.616 | 0.022
11 | YES | 41.003 | 3.446 | 0 | 0.058 | 0.074707 | 0.016707 |  | 2.78 | 0.013
11 | YES | 30.014 | 3.526 | 0 | 0.007 | 0.078115 | 0.071115 |  | 2.7875 | 0.0165
11 | YES | 19.997 | 3.7 | 0 | 0.059 | 0.085529 | 0.026529 |  | 2.812 | 0.03
11 | NO | 10.4875 | 4.647 | 0.43558 | 0.451 | 0.12588 | -0.32512 |  | 2.81 | 0.0079999
11 | NO | 4.511 | 5.543 | 0.10889 | 0.048 | 0.16406 | 0.11606 |  | 2.856 | 0.045
11 | NO | 0 | 0 | 0 | -6.35 | -0.07212 | 6.2779 |  | 3.051 | 0.067
12 | YES | 317.761 | 2.595 | 0 | 0.01 | 0.038447 | 0.028447 |  | 3.075 | 0.025
12 | YES | 300.002 | 2.616 | 0 | 0.022 | 0.039342 | 0.017342 |  | 3.338 | 0.051
12 | YES | 250.007 | 2.78 | 0 | 0.013 | 0.04633 | 0.03333 |  | 3.363 | 0.042
12 | YES | 200.494 | 2.7875 | 0.0049497 | 0.0165 | 0.046649 | 0.030149 |  | 3.383 | 0.045
12 | YES | 174.997 | 2.812 | 0 | 0.03 | 0.047693 | 0.017693 |  | 3.494 | 0.103
12 | YES | 150.003 | 2.81 | 0 | 0.0079999 | 0.047608 | 0.039608 |  | 4.305 | 0.397
12 | YES | 124.996 | 2.856 | 0 | 0.045 | 0.049568 | 0.0045682 |  | 6.938 | 0.092
12 | YES | 100.492 | 3.051 | 0.022627 | 0.067 | 0.057877 | -0.0091237 |  | 2.553 | 0.01
12 | YES | 75.015 | 3.075 | 0 | 0.025 | 0.058899 | 0.033899 |  | 2.656 | 0.004
12 | YES | 50.5025 | 3.338 | 0.011314 | 0.051 | 0.070105 | 0.019105 |  | 2.684 | 0.011
12 | YES | 40.99 | 3.363 | 0 | 0.042 | 0.07117 | 0.02917 |  | 2.707 | -0.0049999
12 | YES | 29.993 | 3.383 | 0 | 0.045 | 0.072022 | 0.027023 |  | 2.768 | 0.024
12 | YES | 19.999 | 3.494 | 0 | 0.103 | 0.076752 | -0.026248 |  | 2.811 | 0.011
12 | YES | 10.504 | 4.305 | 0.039598 | 0.397 | 0.11131 | -0.28569 |  | 2.854 | 0.036
12 | NO | 5.4955 | 5.924 | 0.30264 | -0.49 | 0.18029 | 0.67029 |  | 2.864 | 0.047
12 | YES | 3.037 | 6.938 | 0 | 0.092 | 0.22349 | 0.13149 |  | 3.091 | 0.031
13 | YES | 371.267 | 2.553 | 0 | 0.01 | 0.036658 | 0.026658 |  | 3.2885 | 0.0145
13 | NO | 0 | 0 | 0 | -2.6 | -0.07212 | 2.5279 |  | 3.338 | 0.021
13 | YES | 299.998 | 2.656 | 0 | 0.004 | 0.041046 | 0.037046 |  | 3.424 | 0.04
13 | YES | 249.994 | 2.684 | 0 | 0.011 | 0.042239 | 0.031239 |  | 3.523 | 0.089
13 | YES | 200.491 | 2.707 | 0.0042427 | -0.0049999 | 0.043219 | 0.048219 |  | 6.702 | 0.199
13 | YES | 174.999 | 2.768 | 0 | 0.024 | 0.045818 | 0.021819 |  | 2.631 | 0.0050001
13 | YES | 149.988 | 2.811 | 0 | 0.011 | 0.047651 | 0.03665 |  | 2.675 | 0.01
13 | YES | 125.005 | 2.854 | 0 | 0.036 | 0.049483 | 0.013483 |  | 2.717 | 0.01
13 | YES | 100.503 | 2.864 | 0.0084852 | 0.047 | 0.049909 | 0.0029089 |  | 2.7675 | 0.0215
13 | YES | 74.997 | 3.091 | 0 | 0.031 | 0.059581 | 0.028581 |  | 2.822 | 0.039
13 | YES | 50.493 | 3.2885 | 0.019092 | 0.0145 | 0.067996 | 0.053496 |  | 2.824 | 0.027
13 | YES | 41 | 3.338 | 0 | 0.021 | 0.070105 | 0.049105 |  | 2.836 | 0.011
13 | YES | 29.983 | 3.424 | 0 | 0.04 | 0.073769 | 0.033769 |  | 2.8485 | 0.0235
13 | YES | 19.994 | 3.523 | 0 | 0.089 | 0.077987 | -0.011013 |  | 2.953 | 0.028
13 | NO | 10.4955 | 4.092 | 0.14991 | 0.0050001 | 0.10223 | 0.097231 |  | 3.266 | 0.041
13 | NO | 5.4975 | 6.1545 | 0.35143 | 0.1805 | 0.19011 | 0.0096104 |  | 3.26 | 0.049
13 | YES | 2.704 | 6.702 | 0 | 0.199 | 0.21344 | 0.014438 |  | 3.3 | 0.071
14 | YES | 339.954 | 2.631 | 0 | 0.0050001 | 0.039981 | 0.034981 |  | 3.499 | 0.043
14 | YES | 300.006 | 2.675 | 0 | 0.01 | 0.041856 | 0.031856 |  | 3.8205 | 0.1845
14 | YES | 249.986 | 2.717 | 0 | 0.01 | 0.043645 | 0.033645 |  | 6.252 | -0.01
14 | YES | 200.488 | 2.7675 | 0.0049497 | 0.0215 | 0.045797 | 0.024297 |  | 2.702 | 0.027
14 | YES | 174.996 | 2.822 | 0 | 0.039 | 0.048119 | 0.0091193 |  | 2.734 | 0.024
14 | YES | 150.006 | 2.824 | 0 | 0.027 | 0.048205 | 0.021205 |  | 2.828 | 0.023
14 | YES | 124.991 | 2.836 | 0 | 0.011 | 0.048716 | 0.037716 |  | 2.845 | 0.019
14 | YES | 100.494 | 2.8485 | 0.0021213 | 0.0235 | 0.049248 | 0.025749 |  | 2.889 | 0.036
14 | YES | 74.992 | 2.953 | 0 | 0.028 | 0.053701 | 0.025701 |  | 2.855 | 0.039
14 | YES | 50.503 | 3.266 | 0.022627 | 0.041 | 0.067037 | 0.026037 |  | 2.865 | 0.023
14 | YES | 40.974 | 3.26 | 0 | 0.049 | 0.066782 | 0.017782 |  | 2.938 | 0.025
14 | YES | 29.983 | 3.3 | 0 | 0.071 | 0.068486 | -0.002514 |  | 3.227 | 0.049
14 | YES | 19.992 | 3.499 | 0 | 0.043 | 0.076965 | 0.033965 |  | 3.329 | 0.074
14 | YES | 10.494 | 3.8205 | 0.078489 | 0.1845 | 0.090663 | -0.093837 |  | 3.445 | 0.054
14 | NO | 5.488 | 4.358 | 0.25314 | -0.96 | 0.11357 | 1.0736 |  | 3.699 | 0.038
14 | YES | 2.717 | 6.252 | 0 | -0.01 | 0.19426 | 0.20426 |  | 5.176 | 0.404
15 | YES | 274.923 | 2.702 | 0 | 0.027 | 0.043006 | 0.016006 |  | 2.591 | 0.01
15 | YES | 250 | 2.734 | 0 | 0.024 | 0.04437 | 0.02037 |  | 2.643 | 0.0079999
15 | YES | 200.508 | 2.828 | 0.011314 | 0.023 | 0.048375 | 0.025375 |  | 2.719 | 0.003
15 | YES | 175.005 | 2.845 | 0 | 0.019 | 0.049099 | 0.030099 |  | 2.833 | 0.035
15 | YES | 150.006 | 2.889 | 0 | 0.036 | 0.050974 | 0.014974 |  | 2.854 | 0.0070002
15 | YES | 125.01 | 2.855 | 0 | 0.039 | 0.049525 | 0.010525 |  | 2.877 | 0.014
15 | YES | 100.492 | 2.865 | 0.0070711 | 0.023 | 0.049951 | 0.026952 |  | 2.892 | 0.038
15 | YES | 74.989 | 2.938 | 0 | 0.025 | 0.053062 | 0.028062 |  | 2.889 | 0.033
15 | YES | 50.4885 | 3.227 | 0.0042427 | 0.049 | 0.065376 | 0.016376 |  | 2.957 | 0.048
15 | YES | 40.989 | 3.329 | 0 | 0.074 | 0.069722 | -0.0042784 |  | 3.185 | 0.057
15 | YES | 30.017 | 3.445 | 0 | 0.054 | 0.074664 | 0.020664 |  | 3.371 | 0.06
15 | YES | 20.011 | 3.699 | 0 | 0.038 | 0.085486 | 0.047487 |  | 3.507 | 0.08
15 | YES | 10.5005 | 5.176 | 0.056569 | 0.404 | 0.14842 | -0.25558 |  | 3.592 | 0.054
15 | NO | 5.4955 | 5.942 | 0.34365 | -0.179 | 0.18106 | 0.36006 |  | 2.788 | 0.019
15 | NO | 0 | 0 | 0 | -6.545 | -0.07212 | 6.4729 |  | 2.802 | 0.027
16 | YES | 318.774 | 2.591 | 0 | 0.01 | 0.038277 | 0.028277 |  | 2.891 | 0.042
16 | YES | 300.012 | 2.643 | 0 | 0.0079999 | 0.040492 | 0.032493 |  | 2.897 | 0.021
16 | YES | 250.023 | 2.719 | 0 | 0.003 | 0.043731 | 0.040731 |  | 2.8815 | 0.0315
16 | YES | 200.496 | 2.833 | 0.0014143 | 0.035 | 0.048588 | 0.013588 |  | 2.944 | 0.07
16 | YES | 174.999 | 2.854 | 0 | 0.0070002 | 0.049483 | 0.042483 |  | 3.286 | 0.075
16 | YES | 150.009 | 2.877 | 0 | 0.014 | 0.050463 | 0.036463 |  | 3.389 | 0.082
16 | YES | 125.003 | 2.892 | 0 | 0.038 | 0.051102 | 0.013102 |  | 3.442 | 0.038
16 | YES | 100.495 | 2.889 | 0.0028284 | 0.033 | 0.050974 | 0.017974 |  | 3.669 | 0.269
16 | YES | 75.001 | 2.957 | 0 | 0.048 | 0.053871 | 0.0058713 |  | 4.7695 | 0.3655
16 | YES | 50.496 | 3.185 | 0.011314 | 0.057 | 0.063586 | 0.0065861 |  | 2.497 | 0.024
16 | YES | 41.014 | 3.371 | 0 | 0.06 | 0.071511 | 0.011511 |  | 2.527 | 0.042
16 | YES | 29.999 | 3.507 | 0 | 0.08 | 0.077306 | -0.0026941 |  | 2.829 | 0.025
16 | YES | 19.987 | 3.592 | 0 | 0.054 | 0.080927 | 0.026928 |  | 2.8525 | 0.0425
16 | NO | 10.498 | 4.1565 | 0.1393 | 0.3915 | 0.10498 | -0.28652 |  | 3.019 | 0.04
16 | NO | 5.502 | 6.281 | 0.31395 | -0.115 | 0.1955 | 0.3105 |  | 3.3895 | 0.0805
16 | NO | 0 | 0 | 0 | -6.5 | -0.07212 | 6.4279 |  | 3.516 | 0.084
17 | YES | 188.992 | 2.788 | 0 | 0.019 | 0.046671 | 0.027671 |  | 3.606 | 0.098
17 | YES | 174.999 | 2.802 | 0 | 0.027 | 0.047267 | 0.020267 |  | 4.116 | 0.02
17 | YES | 149.997 | 2.891 | 0 | 0.042 | 0.051059 | 0.0090591 |  | 2.715 | 0.0049999
17 | YES | 125.004 | 2.897 | 0 | 0.021 | 0.051315 | 0.030315 |  | 2.753 | 0.042
17 | YES | 100.494 | 2.8815 | 0.0063638 | 0.0315 | 0.050654 | 0.019154 |  | 2.787 | 0.025
17 | YES | 74.992 | 2.944 | 0 | 0.07 | 0.053317 | -0.016682 |  | 2.8455 | 0.0285
17 | YES | 50.5005 | 3.286 | 0.012728 | 0.075 | 0.067889 | -0.0071106 |  | 2.86 | 0.022
17 | YES | 41.007 | 3.389 | 0 | 0.082 | 0.072278 | -0.009722 |  | 2.89 | 0.022
17 | YES | 30.001 | 3.442 | 0 | 0.038 | 0.074536 | 0.036536 |  | 2.901 | 0.034
17 | YES | 19.991 | 3.669 | 0 | 0.269 | 0.084208 | -0.18479 |  | 2.875 | 0.037
17 | YES | 10.492 | 4.7695 | 0.04879 | 0.3655 | 0.1311 | -0.2344 |  | 2.93 | 0.062
17 | NO | 5.5885 | 5.6055 | 0.4278 | -0.3015 | 0.16672 | 0.46822 |  | 3.2185 | 0.0645
17 | NO | 0 | 0 | 0 | -6.728 | -0.07212 | 6.6559 |  | 3.501 | 0.125
18 | YES | 162.759 | 2.497 | 0 | 0.024 | 0.034272 | 0.010272 |  | 3.598 | 0.062
18 | YES | 149.998 | 2.527 | 0 | 0.042 | 0.03555 | -0.0064502 |  | 3.945 | 0.287
18 | YES | 124.998 | 2.829 | 0 | 0.025 | 0.048418 | 0.023417 |  | 6.771 | 0.019
18 | YES | 100.49 | 2.8525 | 0.00070706 | 0.0425 | 0.049419 | 0.0069188 |  | 2.748 | 0.013
18 | YES | 74.983 | 3.019 | 0 | 0.04 | 0.056513 | 0.016513 |  | 2.767 | 0.023
18 | YES | 50.507 | 3.3895 | 0.016264 | 0.0805 | 0.072299 | -0.0082008 |  | 2.822 | 0.024
18 | YES | 41.01 | 3.516 | 0 | 0.084 | 0.077689 | -0.0063109 |  | 2.8765 | 0.0115
18 | YES | 29.993 | 3.606 | 0 | 0.098 | 0.081524 | -0.016476 |  | 2.896 | 0.027
18 | YES | 19.996 | 4.116 | 0 | 0.02 | 0.10325 | 0.083254 |  | 2.908 | 0.032
18 | NO | 10.501 | 5.843 | 0.14708 | 0.071 | 0.17684 | 0.10584 |  | 2.831 | 0.034
18 | NO | 4.4875 | 7.1605 | 0.16476 | -0.3025 | 0.23297 | 0.53547 |  | 3.131 | 0.118
18 | NO | 0 | 0 | 0 | -7.117 | -0.07212 | 7.0449 |  | 3.535 | 0.061
19 | YES | 351.925 | 2.715 | 0 | 0.0049999 | 0.04356 | 0.03856 |  | 3.869 | 0.097
19 | YES | 300.001 | 2.753 | 0 | 0.042 | 0.045179 | 0.0031792 |  | 3.85 | 0.043
19 | YES | 249.991 | 2.787 | 0 | 0.025 | 0.046628 | 0.021628 |  | 4.062 | 0.195
19 | YES | 200.487 | 2.8455 | 0.0049497 | 0.0285 | 0.049121 | 0.02062 |  | 4.126 | 0.104
19 | YES | 174.995 | 2.86 | 0 | 0.022 | 0.049738 | 0.027739 |  | 6.735 | 0.043
19 | YES | 150.001 | 2.89 | 0 | 0.022 | 0.051017 | 0.029017 |  | 2.835 | 0.022
19 | YES | 124.987 | 2.901 | 0 | 0.034 | 0.051485 | 0.017485 |  | 2.858 | 0.022
19 | YES | 100.506 | 2.875 | 0.0084853 | 0.037 | 0.050378 | 0.013378 |  | 2.892 | -0.071
19 | YES | 75.002 | 2.93 | 0 | 0.062 | 0.052721 | -0.009279 |  | 3.32 | 0.127
19 | YES | 50.5015 | 3.2185 | 0.0091924 | 0.0645 | 0.065013 | 0.00051346 |  | 3.788 | 0.037
19 | YES | 41.001 | 3.501 | 0 | 0.125 | 0.07705 | -0.04795 |  | 3.805 | 0.015
19 | YES | 30.003 | 3.598 | 0 | 0.062 | 0.081183 | 0.019183 |  | 3.911 | 0.082
19 | YES | 19.99 | 3.945 | 0 | 0.287 | 0.095968 | -0.19103 |  | 4.229 | 0.246
19 | NO | 10.498 | 4.91 | 0.16122 | 0.565 | 0.13708 | -0.42792 |  | 4.6065 | 0.3975
19 | NO | 5.504 | 6.632 | 0.11738 | 0.666 | 0.21046 | -0.45554 |  | 4.6605 | 0.1085
19 | YES | 2.655 | 6.771 | 0 | 0.019 | 0.21638 | 0.19738 |  | 4.74 | 0.086
20 | YES | 312.258 | 2.748 | 0 | 0.013 | 0.044966 | 0.031966
20 | YES | 299.998 | 2.767 | 0 | 0.023 | 0.045776 | 0.022776
20 | YES | 250.002 | 2.822 | 0 | 0.024 | 0.048119 | 0.024119
20 | YES | 200.509 | 2.8765 | 0.00070722 | 0.0115 | 0.050441 | 0.038941
20 | YES | 175.009 | 2.896 | 0 | 0.027 | 0.051272 | 0.024272
20 | YES | 150.004 | 2.908 | 0 | 0.032 | 0.051784 | 0.019783
20 | YES | 124.986 | 2.831 | 0 | 0.034 | 0.048503 | 0.014503
20 | YES | 100.499 | 3.131 | 0.06364 | 0.118 | 0.061285 | -0.056715
20 | YES | 75.003 | 3.535 | 0 | 0.061 | 0.078499 | 0.017499
20 | YES | 50.4985 | 3.869 | 0.0098995 | 0.097 | 0.09273 | -0.0042701
20 | YES | 40.982 | 3.85 | 0 | 0.043 | 0.09192 | 0.04892
20 | YES | 29.991 | 4.062 | 0 | 0.195 | 0.10095 | -0.094047
20 | YES | 19.999 | 4.126 | 0 | 0.104 | 0.10368 | -0.0003199
20 | NO | 10.503 | 5.2705 | 0.39386 | 0.9575 | 0.15244 | -0.80506
20 | NO | 5.492 | 6.44 | 0.15132 | 0.058 | 0.20227 | 0.14427
20 | YES | 2.709 | 6.735 | 0 | 0.043 | 0.21484 | 0.17184
21 | YES | 151.799 | 2.835 | 0 | 0.022 | 0.048673 | 0.026673
21 | YES | 125.011 | 2.858 | 0 | 0.022 | 0.049653 | 0.027653
21 | YES | 100.985 | 2.892 | 0 | -0.071 | 0.051102 | 0.1221
21 | YES | 75.004 | 3.32 | 0 | 0.127 | 0.069338 | -0.057662
21 | YES | 50.494 | 3.788 | 0.0028284 | 0.037 | 0.089279 | 0.052278
21 | YES | 40.991 | 3.805 | 0 | 0.015 | 0.090003 | 0.075003
21 | YES | 29.993 | 3.911 | 0 | 0.082 | 0.094519 | 0.012519
21 | YES | 20 | 4.229 | 0 | 0.246 | 0.10807 | -0.13793
21 | YES | 10.5025 | 4.6065 | 0.012021 | 0.3975 | 0.12415 | -0.27335
21 | YES | 5.496 | 4.6605 | 0.017678 | 0.1085 | 0.12645 | 0.017954
21 | YES | 2.506 | 4.74 | 0 | 0.086 | 0.12984 | 0.043841 |  |  |  | 0.059339 | 0.13057
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. OXY_SBE
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(OXY_SBE - OXY_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (OXY_SBE - OXY_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.068011 |  | X | CALC Y
m= | -0.00042455 | 9.691e-05 |  | Std Dev of Resid |  | 0.13323 |  | 2.202 | 0.1014
b= | 0.10233 | 0.012221 |  | Resid Sum of Squares |  | 4.6685 |  | 371.267 | -0.055288
# of Values Extracted = |  |  | 285
# of Values Used in Fit = |  |  | 265
File | Use | Pressure | OXY_SBE | OXY_SBE | Diff (OXY_SBE - OXY_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 173.451 | 0.0485 | 0.00070711 | 0.0045 | 0.028695 | 0.024195 |  | 173.451 | 0.0045
1 | YES | 150.001 | 0.053 | 0 | 0.013 | 0.03865 | 0.02565 |  | 150.001 | 0.013
1 | YES | 125.01 | 0.241 | 0 | 0.099 | 0.04926 | -0.04974 |  | 125.01 | 0.099
1 | YES | 100.504 | 1.0745 | 0.013435 | 0.2115 | 0.059664 | -0.15184 |  | 100.504 | 0.2115
1 | YES | 75 | 0.63 | 0 | -0.175 | 0.070492 | 0.24549 |  | 75 | -0.175
1 | YES | 50.4965 | 2.851 | 0.05374 | 0.118 | 0.080895 | -0.037105 |  | 50.4965 | 0.118
1 | YES | 41.004 | 3.079 | 0 | 0.092 | 0.084925 | -0.0070749 |  | 41.004 | 0.092
1 | YES | 30.003 | 3.281 | 0 | 0.14 | 0.089596 | -0.050404 |  | 30.003 | 0.14
1 | YES | 19.996 | 3.394 | 0 | 0.145 | 0.093844 | -0.051156 |  | 19.996 | 0.145
1 | NO | 4.5095 | 4.1225 | 0.16193 | 0.3435 | 0.10042 | -0.24308 |  | 2.477 | 0.429
1 | NO | 0 | 0 | 0 | -3.918 | 0.10233 | 4.0203 |  | 252.453 | 0.0084999
1 | YES | 2.477 | 6.125 | 0 | 0.429 | 0.10128 | -0.32772 |  | 249.989 | 0.01
2 | YES | 252.453 | 1.2085 | 0.0035355 | 0.0084999 | -0.0048458 | -0.013346 |  | 200.506 | 0.0055001
2 | YES | 249.989 | 1.213 | 0 | 0.01 | -0.0037998 | -0.0138 |  | 175.221 | -0.018
2 | YES | 200.506 | 1.2595 | 0.0035355 | 0.0055001 | 0.017208 | 0.011708 |  | 149.987 | -0.104
2 | YES | 175.221 | 1.3 | 0 | -0.018 | 0.027943 | 0.045943 |  | 125.003 | -0.068
2 | YES | 149.987 | 1.55 | 0 | -0.104 | 0.038656 | 0.14266 |  | 101.002 | 0.049
2 | YES | 125.003 | 2.066 | 0 | -0.068 | 0.049263 | 0.11726 |  | 74.994 | 0.102
2 | YES | 101.002 | 2.145 | 0 | 0.049 | 0.059453 | 0.010453 |  | 50.477 | 0.0345
2 | YES | 74.994 | 2.421 | 0 | 0.102 | 0.070495 | -0.031505 |  | 41.036 | 0.322
2 | YES | 50.477 | 2.8605 | 0.00070722 | 0.0345 | 0.080903 | 0.046403 |  | 30.012 | 0.064
2 | YES | 41.036 | 3.277 | 0 | 0.322 | 0.084912 | -0.23709 |  | 19.846 | -0.147
2 | YES | 30.012 | 3.267 | 0 | 0.064 | 0.089592 | 0.025592 |  | 10.491 | -0.5165
2 | YES | 19.846 | 3.797 | 0 | -0.147 | 0.093908 | 0.24091 |  | 5.4895 | -0.227
2 | YES | 10.491 | 4.1235 | 0.027577 | -0.5165 | 0.097879 | 0.61438 |  | 3.985 | -0.074
2 | YES | 5.4895 | 4.543 | 0.10748 | -0.227 | 0.1 | 0.327 |  | 219.979 | 0.015
2 | YES | 3.985 | 4.711 | 0 | -0.074 | 0.10064 | 0.17464 |  | 202.99 | 0.015
3 | YES | 219.979 | 1.242 | 0 | 0.015 | 0.008941 | -0.006059 |  | 177.016 | 0.04
3 | YES | 202.99 | 1.244 | 0 | 0.015 | 0.016154 | 0.0011537 |  | 150.001 | 0.066
3 | YES | 177.016 | 1.305 | 0 | 0.04 | 0.027181 | -0.012819 |  | 124.984 | 0.079
3 | YES | 150.001 | 1.592 | 0 | 0.066 | 0.03865 | -0.02735 |  | 100.496 | 0.4415
3 | YES | 124.984 | 1.839 | 0 | 0.079 | 0.049271 | -0.029729 |  | 77.016 | 0.214
3 | YES | 100.496 | 2.2775 | 0.044548 | 0.4415 | 0.059668 | -0.38183 |  | 51.988 | 0.296
3 | YES | 77.016 | 2.313 | 0 | 0.214 | 0.069636 | -0.14436 |  | 41.027 | 0.388
3 | YES | 51.988 | 2.689 | 0 | 0.296 | 0.080262 | -0.21574 |  | 30.003 | 0.294
3 | YES | 41.027 | 3.167 | 0 | 0.388 | 0.084915 | -0.30308 |  | 19.991 | 0.452
3 | YES | 30.003 | 3.523 | 0 | 0.294 | 0.089596 | -0.2044 |  | 10.004 | 0.0040002
3 | YES | 19.991 | 4.188 | 0 | 0.452 | 0.093846 | -0.35815 |  | 5.4935 | -0.0155
3 | YES | 10.004 | 4.277 | 0 | 0.0040002 | 0.098086 | 0.094086 |  | 2.595 | 0.233
3 | YES | 5.4935 | 4.2675 | 0.023335 | -0.0155 | 0.1 | 0.1155 |  | 185.396 | 0.023
3 | YES | 2.595 | 4.961 | 0 | 0.233 | 0.10123 | -0.13177 |  | 174.996 | 0.023
4 | YES | 185.396 | 1.189 | 0.0014143 | 0.023 | 0.023623 | 0.0006233 |  | 150.001 | 0.027
4 | YES | 174.996 | 1.197 | 0 | 0.023 | 0.028039 | 0.0050386 |  | 125.011 | 0.027
4 | YES | 150.001 | 1.201 | 0 | 0.027 | 0.03865 | 0.01165 |  | 100.497 | 0.0040001
4 | YES | 125.011 | 1.211 | 0 | 0.027 | 0.04926 | 0.02226 |  | 76.991 | -0.024
4 | YES | 100.497 | 1.521 | 0.0070711 | 0.0040001 | 0.059667 | 0.055667 |  | 50.506 | -0.099
4 | YES | 76.991 | 1.846 | 0 | -0.024 | 0.069647 | 0.093647 |  | 41.009 | 0.017
4 | YES | 50.506 | 2.702 | 0.0042427 | -0.099 | 0.080891 | 0.17989 |  | 29.983 | -0.06
4 | YES | 41.009 | 2.919 | 0 | 0.017 | 0.084923 | 0.067923 |  | 19.989 | -0.037
4 | YES | 29.983 | 3.096 | 0 | -0.06 | 0.089604 | 0.1496 |  | 10.4915 | 0.15
4 | YES | 19.989 | 3.443 | 0 | -0.037 | 0.093847 | 0.13085 |  | 5.5005 | 0.2525
4 | YES | 10.4915 | 4.093 | 0.0014141 | 0.15 | 0.097879 | -0.052121 |  | 2.693 | 0.134
4 | YES | 5.5005 | 4.0935 | 0.00070706 | 0.2525 | 0.099998 | -0.1525 |  | 175.173 | 0.0185
4 | YES | 2.693 | 4.083 | 0 | 0.134 | 0.10119 | -0.03281 |  | 124.991 | 0.039
5 | YES | 175.173 | 1.2195 | 0.0035355 | 0.0185 | 0.027963 | 0.0094634 |  | 100.495 | -0.036
5 | NO | 0 | 0 | 0 | -1.213 | 0.10233 | 1.3153 |  | 75.006 | -0.201
5 | YES | 124.991 | 1.291 | 0 | 0.039 | 0.049268 | 0.010268 |  | 49.502 | 0.00099993
5 | YES | 100.495 | 1.719 | 0.041012 | -0.036 | 0.059668 | 0.095668 |  | 40.995 | 0.026
5 | YES | 75.006 | 2.327 | 0 | -0.201 | 0.07049 | 0.27149 |  | 30.001 | 0.035
5 | YES | 49.502 | 2.797 | 0.0070711 | 0.00099993 | 0.081317 | 0.080317 |  | 19.988 | 0.05
5 | YES | 40.995 | 2.902 | 0 | 0.026 | 0.084929 | 0.058929 |  | 10.5055 | 0.036
5 | YES | 30.001 | 2.989 | 0 | 0.035 | 0.089596 | 0.054596 |  | 5.5105 | -0.0095
5 | YES | 19.988 | 3.048 | 0 | 0.05 | 0.093847 | 0.043847 |  | 3.129 | 0.007
5 | YES | 10.5055 | 3.154 | 0.0014141 | 0.036 | 0.097873 | 0.061873 |  | 114 | 0
5 | YES | 5.5105 | 3.2195 | 0.026163 | -0.0095 | 0.099994 | 0.10949 |  | 100.998 | -0.023
5 | YES | 3.129 | 3.282 | 0 | 0.007 | 0.10101 | 0.094005 |  | 74.991 | 0.035
6 | YES | 114 | 1.594 | 0 | 0 | 0.053935 | 0.053935 |  | 50.4905 | 0.0435
6 | YES | 100.998 | 1.82 | 0 | -0.023 | 0.059455 | 0.082455 |  | 40.992 | 0.138
6 | YES | 74.991 | 2.072 | 0 | 0.035 | 0.070496 | 0.035496 |  | 29.985 | 0.078
6 | YES | 50.4905 | 2.8865 | 0.014849 | 0.0435 | 0.080898 | 0.037398 |  | 19.999 | 0.465
6 | YES | 40.992 | 3.314 | 0 | 0.138 | 0.08493 | -0.05307 |  | 10.5035 | 0.2165
6 | YES | 29.985 | 3.462 | 0 | 0.078 | 0.089603 | 0.011603 |  | 5.504 | 0.082
6 | YES | 19.999 | 3.885 | 0 | 0.465 | 0.093843 | -0.37116 |  | 2.212 | 0.037
6 | YES | 10.5035 | 3.9585 | 0.02192 | 0.2165 | 0.097874 | -0.11863 |  | 123.143 | 0.0024999
6 | YES | 5.504 | 4.034 | 0.0070712 | 0.082 | 0.099997 | 0.017997 |  | 100.495 | 0.0205
6 | YES | 2.212 | 4.047 | 0 | 0.037 | 0.10139 | 0.064395 |  | 75.01 | -0.071
7 | YES | 123.143 | 1.8955 | 0.0035355 | 0.0024999 | 0.050053 | 0.047553 |  | 50.5045 | 0.0595
7 | YES | 100.495 | 2.2795 | 0.0021212 | 0.0205 | 0.059668 | 0.039168 |  | 40.999 | 0.151
7 | YES | 75.01 | 2.691 | 0 | -0.071 | 0.070488 | 0.14149 |  | 29.988 | 0.239
7 | YES | 50.5045 | 3.0885 | 0.0035356 | 0.0595 | 0.080892 | 0.021392 |  | 19.998 | 0.345
7 | YES | 40.999 | 3.284 | 0 | 0.151 | 0.084927 | -0.066073 |  | 10.5015 | -0.003
7 | YES | 29.988 | 3.391 | 0 | 0.239 | 0.089602 | -0.1494 |  | 5.494 | -0.0195
7 | YES | 19.998 | 3.692 | 0 | 0.345 | 0.093843 | -0.25116 |  | 2.708 | 0.004
7 | YES | 10.5015 | 3.731 | 0.0098995 | -0.003 | 0.097875 | 0.10088 |  | 145.194 | -0.017
7 | YES | 5.494 | 3.7395 | 0.0035356 | -0.0195 | 0.1 | 0.1195 |  | 124.991 | 0.057
7 | YES | 2.708 | 3.792 | 0 | 0.004 | 0.10118 | 0.097184 |  | 100.995 | -0.023
8 | YES | 145.194 | 2.058 | 0.012728 | -0.017 | 0.040691 | 0.057691 |  | 75.007 | 0.044
8 | YES | 124.991 | 2.187 | 0 | 0.057 | 0.049268 | -0.0077316 |  | 50.4985 | 0.0785
8 | YES | 100.995 | 2.678 | 0 | -0.023 | 0.059456 | 0.082456 |  | 41.004 | 0.024
8 | YES | 75.007 | 2.997 | 0 | 0.044 | 0.070489 | 0.026489 |  | 29.996 | 0.034
8 | YES | 50.4985 | 3.2705 | 0.028991 | 0.0785 | 0.080894 | 0.0023942 |  | 19.988 | 0.101
8 | YES | 41.004 | 3.369 | 0 | 0.024 | 0.084925 | 0.060925 |  | 10.512 | 0.051
8 | YES | 29.996 | 3.404 | 0 | 0.034 | 0.089599 | 0.055598 |  | 5.5 | 0.04
8 | YES | 19.988 | 3.471 | 0 | 0.101 | 0.093847 | -0.0071525 |  | 2.202 | 0.078
8 | YES | 10.512 | 3.465 | 0.0028284 | 0.051 | 0.097871 | 0.04687 |  | 218.883 | 0
8 | YES | 5.5 | 3.485 | 0.0056569 | 0.04 | 0.099998 | 0.059998 |  | 200.508 | -0.0025001
8 | YES | 2.202 | 3.543 | 0 | 0.078 | 0.1014 | 0.023398 |  | 175.002 | 0.016
9 | YES | 218.883 | 2.662 | 0 | 0 | 0.0094063 | 0.0094063 |  | 150.011 | 0.018
9 | YES | 200.508 | 2.6645 | 0.00070706 | -0.0025001 | 0.017207 | 0.019708 |  | 125.005 | 0.051
9 | YES | 175.002 | 2.687 | 0 | 0.016 | 0.028036 | 0.012036 |  | 101.009 | 0.0049999
9 | YES | 150.011 | 2.71 | 0 | 0.018 | 0.038646 | 0.020646 |  | 77.007 | 0.103
9 | YES | 125.005 | 2.799 | 0 | 0.051 | 0.049262 | -0.0017377 |  | 50.495 | 2.3842e-07
9 | YES | 101.009 | 3.01 | 0 | 0.0049999 | 0.05945 | 0.05445 |  | 41.01 | -0.044
9 | YES | 77.007 | 3.255 | 0 | 0.103 | 0.06964 | -0.03336 |  | 29.993 | 0.093
9 | YES | 50.495 | 3.36 | 0.0028284 | 2.3842e-07 | 0.080896 | 0.080895 |  | 22 | 0.222
9 | YES | 41.01 | 3.376 | 0 | -0.044 | 0.084923 | 0.12892 |  | 11.4935 | -0.0785
9 | YES | 29.993 | 3.516 | 0 | 0.093 | 0.0896 | -0.0034 |  | 5.4995 | 0.318
9 | YES | 22 | 3.699 | 0 | 0.222 | 0.092993 | -0.12901 |  | 2.748 | 1.237
9 | YES | 11.4935 | 3.7645 | 0.030406 | -0.0785 | 0.097454 | 0.17595 |  | 226.669 | 0.026
9 | YES | 5.4995 | 4.353 | 0.13435 | 0.318 | 0.099999 | -0.218 |  | 200.49 | 0.0315
9 | YES | 2.748 | 5.394 | 0 | 1.237 | 0.10117 | -1.1358 |  | 176.984 | 0.029
10 | YES | 226.669 | 2.715 | 0.0056569 | 0.026 | 0.0061008 | -0.0199 |  | 152.993 | -0.04
10 | YES | 200.49 | 2.7145 | 0.006364 | 0.0315 | 0.017215 | -0.014285 |  | 124.993 | 0.041
10 | YES | 176.984 | 2.747 | 0 | 0.029 | 0.027195 | -0.0018054 |  | 100.496 | 0.0145
10 | YES | 152.993 | 2.806 | 0 | -0.04 | 0.03738 | 0.07738 |  | 75.007 | -0.079
10 | YES | 124.993 | 3.062 | 0 | 0.041 | 0.049267 | 0.0082674 |  | 51.509 | 0.034
10 | YES | 100.496 | 3.1095 | 0.006364 | 0.0145 | 0.059668 | 0.045168 |  | 40.997 | 0.062
10 | YES | 75.007 | 3.205 | 0 | -0.079 | 0.070489 | 0.14949 |  | 29.982 | 0.108
10 | YES | 51.509 | 3.39 | 0.0084852 | 0.034 | 0.080465 | 0.046465 |  | 19.989 | 0.106
10 | YES | 40.997 | 3.438 | 0 | 0.062 | 0.084928 | 0.022928 |  | 10.501 | -0.094
10 | YES | 29.982 | 3.506 | 0 | 0.108 | 0.089605 | -0.018395 |  | 5.493 | -0.1545
10 | YES | 19.989 | 3.584 | 0 | 0.106 | 0.093847 | -0.012153 |  | 2.71 | -0.218
10 | YES | 10.501 | 3.885 | 0.014142 | -0.094 | 0.097875 | 0.19188 |  | 171.333 | 0.027
10 | YES | 5.493 | 3.9215 | 0.0021213 | -0.1545 | 0.1 | 0.2545 |  | 149.986 | 0.038
10 | YES | 2.71 | 3.974 | 0 | -0.218 | 0.10118 | 0.31918 |  | 125.003 | 0.003
11 | YES | 171.333 | 2.808 | 0.0084852 | 0.027 | 0.029594 | 0.0025936 |  | 100.512 | 0.0415
11 | YES | 149.986 | 2.942 | 0 | 0.038 | 0.038657 | 0.00065676 |  | 75.002 | 0.028
11 | YES | 125.003 | 2.915 | 0 | 0.003 | 0.049263 | 0.046263 |  | 50.4985 | 0.0605
11 | YES | 100.512 | 3.1775 | 0.0049497 | 0.0415 | 0.059661 | 0.018161 |  | 41.003 | 0.058
11 | YES | 75.002 | 3.298 | 0 | 0.028 | 0.070491 | 0.042491 |  | 30.014 | 0.007
11 | YES | 50.4985 | 3.3765 | 0.019092 | 0.0605 | 0.080894 | 0.020394 |  | 19.997 | 0.059
11 | YES | 41.003 | 3.446 | 0 | 0.058 | 0.084926 | 0.026925 |  | 4.511 | 0.048
11 | YES | 30.014 | 3.526 | 0 | 0.007 | 0.089591 | 0.082591 |  | 317.761 | 0.01
11 | YES | 19.997 | 3.7 | 0 | 0.059 | 0.093844 | 0.034844 |  | 300.002 | 0.022
11 | NO | 10.4875 | 4.647 | 0.43558 | 0.451 | 0.097881 | -0.35312 |  | 250.007 | 0.013
11 | YES | 4.511 | 5.543 | 0.10889 | 0.048 | 0.10042 | 0.052418 |  | 200.494 | 0.0165
11 | NO | 0 | 0 | 0 | -6.35 | 0.10233 | 6.4523 |  | 174.997 | 0.03
12 | YES | 317.761 | 2.595 | 0 | 0.01 | -0.032572 | -0.042572 |  | 150.003 | 0.0079999
12 | YES | 300.002 | 2.616 | 0 | 0.022 | -0.025033 | -0.047033 |  | 124.996 | 0.045
12 | YES | 250.007 | 2.78 | 0 | 0.013 | -0.0038074 | -0.016807 |  | 100.492 | 0.067
12 | YES | 200.494 | 2.7875 | 0.0049497 | 0.0165 | 0.017213 | 0.00071342 |  | 75.015 | 0.025
12 | YES | 174.997 | 2.812 | 0 | 0.03 | 0.028038 | -0.0019618 |  | 50.5025 | 0.051
12 | YES | 150.003 | 2.81 | 0 | 0.0079999 | 0.038649 | 0.03065 |  | 40.99 | 0.042
12 | YES | 124.996 | 2.856 | 0 | 0.045 | 0.049266 | 0.0042664 |  | 29.993 | 0.045
12 | YES | 100.492 | 3.051 | 0.022627 | 0.067 | 0.059669 | -0.0073308 |  | 19.999 | 0.103
12 | YES | 75.015 | 3.075 | 0 | 0.025 | 0.070486 | 0.045486 |  | 10.504 | 0.397
12 | YES | 50.5025 | 3.338 | 0.011314 | 0.051 | 0.080893 | 0.029893 |  | 3.037 | 0.092
12 | YES | 40.99 | 3.363 | 0 | 0.042 | 0.084931 | 0.042931 |  | 371.267 | 0.01
12 | YES | 29.993 | 3.383 | 0 | 0.045 | 0.0896 | 0.0446 |  | 299.998 | 0.004
12 | YES | 19.999 | 3.494 | 0 | 0.103 | 0.093843 | -0.0091572 |  | 249.994 | 0.011
12 | YES | 10.504 | 4.305 | 0.039598 | 0.397 | 0.097874 | -0.29913 |  | 200.491 | -0.0049999
12 | NO | 5.4955 | 5.924 | 0.30264 | -0.49 | 0.1 | 0.59 |  | 174.999 | 0.024
12 | YES | 3.037 | 6.938 | 0 | 0.092 | 0.10104 | 0.0090441 |  | 149.988 | 0.011
13 | YES | 371.267 | 2.553 | 0 | 0.01 | -0.055288 | -0.065288 |  | 125.005 | 0.036
13 | NO | 0 | 0 | 0 | -2.6 | 0.10233 | 2.7023 |  | 100.503 | 0.047
13 | YES | 299.998 | 2.656 | 0 | 0.004 | -0.025031 | -0.029031 |  | 74.997 | 0.031
13 | YES | 249.994 | 2.684 | 0 | 0.011 | -0.0038019 | -0.014802 |  | 50.493 | 0.0145
13 | YES | 200.491 | 2.707 | 0.0042427 | -0.0049999 | 0.017215 | 0.022215 |  | 41 | 0.021
13 | YES | 174.999 | 2.768 | 0 | 0.024 | 0.028037 | 0.0040375 |  | 29.983 | 0.04
13 | YES | 149.988 | 2.811 | 0 | 0.011 | 0.038656 | 0.027656 |  | 19.994 | 0.089
13 | YES | 125.005 | 2.854 | 0 | 0.036 | 0.049262 | 0.013262 |  | 10.4955 | 0.0050001
13 | YES | 100.503 | 2.864 | 0.0084852 | 0.047 | 0.059665 | 0.012665 |  | 2.704 | 0.199
13 | YES | 74.997 | 3.091 | 0 | 0.031 | 0.070493 | 0.039493 |  | 339.954 | 0.0050001
13 | YES | 50.493 | 3.2885 | 0.019092 | 0.0145 | 0.080897 | 0.066396 |  | 300.006 | 0.01
13 | YES | 41 | 3.338 | 0 | 0.021 | 0.084927 | 0.063927 |  | 249.986 | 0.01
13 | YES | 29.983 | 3.424 | 0 | 0.04 | 0.089604 | 0.049604 |  | 200.488 | 0.0215
13 | YES | 19.994 | 3.523 | 0 | 0.089 | 0.093845 | 0.0048449 |  | 174.996 | 0.039
13 | YES | 10.4955 | 4.092 | 0.14991 | 0.0050001 | 0.097878 | 0.092877 |  | 150.006 | 0.027
13 | NO | 5.4975 | 6.1545 | 0.35143 | 0.1805 | 0.099999 | -0.080501 |  | 124.991 | 0.011
13 | YES | 2.704 | 6.702 | 0 | 0.199 | 0.10119 | -0.097815 |  | 100.494 | 0.0235
14 | YES | 339.954 | 2.631 | 0 | 0.0050001 | -0.041995 | -0.046995 |  | 74.992 | 0.028
14 | YES | 300.006 | 2.675 | 0 | 0.01 | -0.025035 | -0.035035 |  | 50.503 | 0.041
14 | YES | 249.986 | 2.717 | 0 | 0.01 | -0.0037985 | -0.013798 |  | 40.974 | 0.049
14 | YES | 200.488 | 2.7675 | 0.0049497 | 0.0215 | 0.017216 | -0.0042839 |  | 29.983 | 0.071
14 | YES | 174.996 | 2.822 | 0 | 0.039 | 0.028039 | -0.010961 |  | 19.992 | 0.043
14 | YES | 150.006 | 2.824 | 0 | 0.027 | 0.038648 | 0.011648 |  | 10.494 | 0.1845
14 | YES | 124.991 | 2.836 | 0 | 0.011 | 0.049268 | 0.038268 |  | 2.717 | -0.01
14 | YES | 100.494 | 2.8485 | 0.0021213 | 0.0235 | 0.059669 | 0.036169 |  | 274.923 | 0.027
14 | YES | 74.992 | 2.953 | 0 | 0.028 | 0.070495 | 0.042495 |  | 250 | 0.024
14 | YES | 50.503 | 3.266 | 0.022627 | 0.041 | 0.080892 | 0.039892 |  | 200.508 | 0.023
14 | YES | 40.974 | 3.26 | 0 | 0.049 | 0.084938 | 0.035938 |  | 175.005 | 0.019
14 | YES | 29.983 | 3.3 | 0 | 0.071 | 0.089604 | 0.018604 |  | 150.006 | 0.036
14 | YES | 19.992 | 3.499 | 0 | 0.043 | 0.093846 | 0.050846 |  | 125.01 | 0.039
14 | YES | 10.494 | 3.8205 | 0.078489 | 0.1845 | 0.097878 | -0.086622 |  | 100.492 | 0.023
14 | NO | 5.488 | 4.358 | 0.25314 | -0.96 | 0.1 | 1.06 |  | 74.989 | 0.025
14 | YES | 2.717 | 6.252 | 0 | -0.01 | 0.10118 | 0.11118 |  | 50.4885 | 0.049
15 | YES | 274.923 | 2.702 | 0 | 0.027 | -0.014386 | -0.041386 |  | 40.989 | 0.074
15 | YES | 250 | 2.734 | 0 | 0.024 | -0.0038044 | -0.027804 |  | 30.017 | 0.054
15 | YES | 200.508 | 2.828 | 0.011314 | 0.023 | 0.017207 | -0.0057925 |  | 20.011 | 0.038
15 | YES | 175.005 | 2.845 | 0 | 0.019 | 0.028035 | 0.0090347 |  | 10.5005 | 0.404
15 | YES | 150.006 | 2.889 | 0 | 0.036 | 0.038648 | 0.0026482 |  | 318.774 | 0.01
15 | YES | 125.01 | 2.855 | 0 | 0.039 | 0.04926 | 0.01026 |  | 300.012 | 0.0079999
15 | YES | 100.492 | 2.865 | 0.0070711 | 0.023 | 0.059669 | 0.03667 |  | 250.023 | 0.003
15 | YES | 74.989 | 2.938 | 0 | 0.025 | 0.070497 | 0.045497 |  | 200.496 | 0.035
15 | YES | 50.4885 | 3.227 | 0.0042427 | 0.049 | 0.080898 | 0.031898 |  | 174.999 | 0.0070002
15 | YES | 40.989 | 3.329 | 0 | 0.074 | 0.084931 | 0.010932 |  | 150.009 | 0.014
15 | YES | 30.017 | 3.445 | 0 | 0.054 | 0.08959 | 0.03559 |  | 125.003 | 0.038
15 | YES | 20.011 | 3.699 | 0 | 0.038 | 0.093838 | 0.055838 |  | 100.495 | 0.033
15 | YES | 10.5005 | 5.176 | 0.056569 | 0.404 | 0.097875 | -0.30612 |  | 75.001 | 0.048
15 | NO | 5.4955 | 5.942 | 0.34365 | -0.179 | 0.1 | 0.279 |  | 50.496 | 0.057
15 | NO | 0 | 0 | 0 | -6.545 | 0.10233 | 6.6473 |  | 41.014 | 0.06
16 | YES | 318.774 | 2.591 | 0 | 0.01 | -0.033003 | -0.043003 |  | 29.999 | 0.08
16 | YES | 300.012 | 2.643 | 0 | 0.0079999 | -0.025037 | -0.033037 |  | 19.987 | 0.054
16 | YES | 250.023 | 2.719 | 0 | 0.003 | -0.0038142 | -0.0068142 |  | 10.498 | 0.3915
16 | YES | 200.496 | 2.833 | 0.0014143 | 0.035 | 0.017213 | -0.017788 |  | 188.992 | 0.019
16 | YES | 174.999 | 2.854 | 0 | 0.0070002 | 0.028037 | 0.021037 |  | 174.999 | 0.027
16 | YES | 150.009 | 2.877 | 0 | 0.014 | 0.038647 | 0.024647 |  | 149.997 | 0.042
16 | YES | 125.003 | 2.892 | 0 | 0.038 | 0.049263 | 0.011263 |  | 125.004 | 0.021
16 | YES | 100.495 | 2.889 | 0.0028284 | 0.033 | 0.059668 | 0.026668 |  | 100.494 | 0.0315
16 | YES | 75.001 | 2.957 | 0 | 0.048 | 0.070492 | 0.022492 |  | 74.992 | 0.07
16 | YES | 50.496 | 3.185 | 0.011314 | 0.057 | 0.080895 | 0.023895 |  | 50.5005 | 0.075
16 | YES | 41.014 | 3.371 | 0 | 0.06 | 0.084921 | 0.024921 |  | 41.007 | 0.082
16 | YES | 29.999 | 3.507 | 0 | 0.08 | 0.089597 | 0.0095974 |  | 30.001 | 0.038
16 | YES | 19.987 | 3.592 | 0 | 0.054 | 0.093848 | 0.039848 |  | 19.991 | 0.269
16 | YES | 10.498 | 4.1565 | 0.1393 | 0.3915 | 0.097876 | -0.29362 |  | 10.492 | 0.3655
16 | NO | 5.502 | 6.281 | 0.31395 | -0.115 | 0.099998 | 0.215 |  | 162.759 | 0.024
16 | NO | 0 | 0 | 0 | -6.5 | 0.10233 | 6.6023 |  | 149.998 | 0.042
17 | YES | 188.992 | 2.788 | 0 | 0.019 | 0.022097 | 0.0030965 |  | 124.998 | 0.025
17 | YES | 174.999 | 2.802 | 0 | 0.027 | 0.028037 | 0.0010374 |  | 100.49 | 0.0425
17 | YES | 149.997 | 2.891 | 0 | 0.042 | 0.038652 | -0.0033481 |  | 74.983 | 0.04
17 | YES | 125.004 | 2.897 | 0 | 0.021 | 0.049263 | 0.028263 |  | 50.507 | 0.0805
17 | YES | 100.494 | 2.8815 | 0.0063638 | 0.0315 | 0.059669 | 0.028168 |  | 41.01 | 0.084
17 | YES | 74.992 | 2.944 | 0 | 0.07 | 0.070495 | 0.00049557 |  | 29.993 | 0.098
17 | YES | 50.5005 | 3.286 | 0.012728 | 0.075 | 0.080893 | 0.0058934 |  | 19.996 | 0.02
17 | YES | 41.007 | 3.389 | 0 | 0.082 | 0.084924 | 0.0029238 |  | 10.501 | 0.071
17 | YES | 30.001 | 3.442 | 0 | 0.038 | 0.089596 | 0.051597 |  | 351.925 | 0.0049999
17 | YES | 19.991 | 3.669 | 0 | 0.269 | 0.093846 | -0.17515 |  | 300.001 | 0.042
17 | YES | 10.492 | 4.7695 | 0.04879 | 0.3655 | 0.097879 | -0.26762 |  | 249.991 | 0.025
17 | NO | 5.5885 | 5.6055 | 0.4278 | -0.3015 | 0.099961 | 0.40146 |  | 200.487 | 0.0285
17 | NO | 0 | 0 | 0 | -6.728 | 0.10233 | 6.8303 |  | 174.995 | 0.022
18 | YES | 162.759 | 2.497 | 0 | 0.024 | 0.033234 | 0.009234 |  | 150.001 | 0.022
18 | YES | 149.998 | 2.527 | 0 | 0.042 | 0.038652 | -0.0033485 |  | 124.987 | 0.034
18 | YES | 124.998 | 2.829 | 0 | 0.025 | 0.049265 | 0.024265 |  | 100.506 | 0.037
18 | YES | 100.49 | 2.8525 | 0.00070706 | 0.0425 | 0.05967 | 0.01717 |  | 75.002 | 0.062
18 | YES | 74.983 | 3.019 | 0 | 0.04 | 0.070499 | 0.030499 |  | 50.5015 | 0.0645
18 | YES | 50.507 | 3.3895 | 0.016264 | 0.0805 | 0.080891 | 0.00039051 |  | 41.001 | 0.125
18 | YES | 41.01 | 3.516 | 0 | 0.084 | 0.084923 | 0.00092247 |  | 30.003 | 0.062
18 | YES | 29.993 | 3.606 | 0 | 0.098 | 0.0896 | -0.0084001 |  | 19.99 | 0.287
18 | YES | 19.996 | 4.116 | 0 | 0.02 | 0.093844 | 0.073844 |  | 5.504 | 0.666
18 | YES | 10.501 | 5.843 | 0.14708 | 0.071 | 0.097875 | 0.026875 |  | 2.655 | 0.019
18 | NO | 4.4875 | 7.1605 | 0.16476 | -0.3025 | 0.10043 | 0.40293 |  | 312.258 | 0.013
18 | NO | 0 | 0 | 0 | -7.117 | 0.10233 | 7.2193 |  | 299.998 | 0.023
19 | YES | 351.925 | 2.715 | 0 | 0.0049999 | -0.047077 | -0.052077 |  | 250.002 | 0.024
19 | YES | 300.001 | 2.753 | 0 | 0.042 | -0.025032 | -0.067033 |  | 200.509 | 0.0115
19 | YES | 249.991 | 2.787 | 0 | 0.025 | -0.0038006 | -0.0288 |  | 175.009 | 0.027
19 | YES | 200.487 | 2.8455 | 0.0049497 | 0.0285 | 0.017216 | -0.011284 |  | 150.004 | 0.032
19 | YES | 174.995 | 2.86 | 0 | 0.022 | 0.028039 | 0.0060393 |  | 124.986 | 0.034
19 | YES | 150.001 | 2.89 | 0 | 0.022 | 0.03865 | 0.01665 |  | 100.499 | 0.118
19 | YES | 124.987 | 2.901 | 0 | 0.034 | 0.04927 | 0.01527 |  | 75.003 | 0.061
19 | YES | 100.506 | 2.875 | 0.0084853 | 0.037 | 0.059663 | 0.022664 |  | 50.4985 | 0.097
19 | YES | 75.002 | 2.93 | 0 | 0.062 | 0.070491 | 0.0084912 |  | 40.982 | 0.043
19 | YES | 50.5015 | 3.2185 | 0.0091924 | 0.0645 | 0.080893 | 0.016393 |  | 29.991 | 0.195
19 | YES | 41.001 | 3.501 | 0 | 0.125 | 0.084926 | -0.040074 |  | 19.999 | 0.104
19 | YES | 30.003 | 3.598 | 0 | 0.062 | 0.089596 | 0.027596 |  | 2.709 | 0.043
19 | YES | 19.99 | 3.945 | 0 | 0.287 | 0.093847 | -0.19315 |  | 151.799 | 0.022
19 | NO | 10.498 | 4.91 | 0.16122 | 0.565 | 0.097876 | -0.46712 |  | 125.011 | 0.022
19 | YES | 5.504 | 6.632 | 0.11738 | 0.666 | 0.099997 | -0.566 |  | 100.985 | -0.071
19 | YES | 2.655 | 6.771 | 0 | 0.019 | 0.10121 | 0.082206 |  | 75.004 | 0.127
20 | YES | 312.258 | 2.748 | 0 | 0.013 | -0.030236 | -0.043236 |  | 50.494 | 0.037
20 | YES | 299.998 | 2.767 | 0 | 0.023 | -0.025031 | -0.048031 |  | 40.991 | 0.015
20 | YES | 250.002 | 2.822 | 0 | 0.024 | -0.0038053 | -0.027805 |  | 29.993 | 0.082
20 | YES | 200.509 | 2.8765 | 0.00070722 | 0.0115 | 0.017207 | 0.005707 |  | 20 | 0.246
20 | YES | 175.009 | 2.896 | 0 | 0.027 | 0.028033 | 0.0010331 |  | 10.5025 | 0.3975
20 | YES | 150.004 | 2.908 | 0 | 0.032 | 0.038649 | 0.0066489 |  | 5.496 | 0.1085
20 | YES | 124.986 | 2.831 | 0 | 0.034 | 0.04927 | 0.01527 |  | 2.506 | 0.086 | 0.062574 | 0.13775
20 | YES | 100.499 | 3.131 | 0.06364 | 0.118 | 0.059666 | -0.058334
20 | YES | 75.003 | 3.535 | 0 | 0.061 | 0.070491 | 0.0094907
20 | YES | 50.4985 | 3.869 | 0.0098995 | 0.097 | 0.080894 | -0.016106
20 | YES | 40.982 | 3.85 | 0 | 0.043 | 0.084934 | 0.041934
20 | YES | 29.991 | 4.062 | 0 | 0.195 | 0.089601 | -0.1054
20 | YES | 19.999 | 4.126 | 0 | 0.104 | 0.093843 | -0.010157
20 | NO | 10.503 | 5.2705 | 0.39386 | 0.9575 | 0.097874 | -0.85963
20 | NO | 5.492 | 6.44 | 0.15132 | 0.058 | 0.1 | 0.042002
20 | YES | 2.709 | 6.735 | 0 | 0.043 | 0.10118 | 0.058183
21 | YES | 151.799 | 2.835 | 0 | 0.022 | 0.037887 | 0.015887
21 | YES | 125.011 | 2.858 | 0 | 0.022 | 0.04926 | 0.02726
21 | YES | 100.985 | 2.892 | 0 | -0.071 | 0.05946 | 0.13046
21 | YES | 75.004 | 3.32 | 0 | 0.127 | 0.07049 | -0.05651
21 | YES | 50.494 | 3.788 | 0.0028284 | 0.037 | 0.080896 | 0.043896
21 | YES | 40.991 | 3.805 | 0 | 0.015 | 0.084931 | 0.069931
21 | YES | 29.993 | 3.911 | 0 | 0.082 | 0.0896 | 0.0075999
21 | YES | 20 | 4.229 | 0 | 0.246 | 0.093842 | -0.15216
21 | YES | 10.5025 | 4.6065 | 0.012021 | 0.3975 | 0.097875 | -0.29963
21 | YES | 5.496 | 4.6605 | 0.017678 | 0.1085 | 0.1 | -0.0084999
21 | YES | 2.506 | 4.74 | 0 | 0.086 | 0.10127 | 0.01527
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. OXY_SBE
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(OXY_SBE - OXY_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (OXY_SBE - OXY_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 3.4503e-05 |  | X | CALC Y
m= | 0.0001051 | 0.0011204 |  | Std Dev of Resid |  | 0.10827 |  | 1 | 0.053695
b= | 0.05359 | 0.014153 |  | Resid Sum of Squares |  | 2.9891 |  | 21 | 0.055797
# of Values Extracted = |  |  | 285
# of Values Used in Fit = |  |  | 257
File | Use | Pressure | OXY_SBE | OXY_SBE | Diff (OXY_SBE - OXY_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 173.451 | 0.0485 | 0.00070711 | 0.0045 | 0.053695 | 0.049195 |  | 1 | 0.0045
1 | YES | 150.001 | 0.053 | 0 | 0.013 | 0.053695 | 0.040695 |  | 1 | 0.013
1 | YES | 125.01 | 0.241 | 0 | 0.099 | 0.053695 | -0.045305 |  | 1 | 0.099
1 | YES | 100.504 | 1.0745 | 0.013435 | 0.2115 | 0.053695 | -0.1578 |  | 1 | 0.2115
1 | YES | 75 | 0.63 | 0 | -0.175 | 0.053695 | 0.2287 |  | 1 | -0.175
1 | YES | 50.4965 | 2.851 | 0.05374 | 0.118 | 0.053695 | -0.064305 |  | 1 | 0.118
1 | YES | 41.004 | 3.079 | 0 | 0.092 | 0.053695 | -0.038305 |  | 1 | 0.092
1 | YES | 30.003 | 3.281 | 0 | 0.14 | 0.053695 | -0.086305 |  | 1 | 0.14
1 | YES | 19.996 | 3.394 | 0 | 0.145 | 0.053695 | -0.091305 |  | 1 | 0.145
1 | NO | 4.5095 | 4.1225 | 0.16193 | 0.3435 | 0.053695 | -0.2898 |  | 1 | 0.429
1 | NO | 0 | 0 | 0 | -3.918 | 0.053695 | 3.9717 |  | 2 | 0.0084999
1 | YES | 2.477 | 6.125 | 0 | 0.429 | 0.053695 | -0.3753 |  | 2 | 0.01
2 | YES | 252.453 | 1.2085 | 0.0035355 | 0.0084999 | 0.0538 | 0.0453 |  | 2 | 0.0055001
2 | YES | 249.989 | 1.213 | 0 | 0.01 | 0.0538 | 0.0438 |  | 2 | -0.018
2 | YES | 200.506 | 1.2595 | 0.0035355 | 0.0055001 | 0.0538 | 0.0483 |  | 2 | -0.104
2 | YES | 175.221 | 1.3 | 0 | -0.018 | 0.0538 | 0.0718 |  | 2 | -0.068
2 | YES | 149.987 | 1.55 | 0 | -0.104 | 0.0538 | 0.1578 |  | 2 | 0.049
2 | YES | 125.003 | 2.066 | 0 | -0.068 | 0.0538 | 0.1218 |  | 2 | 0.102
2 | YES | 101.002 | 2.145 | 0 | 0.049 | 0.0538 | 0.0048003 |  | 2 | 0.0345
2 | YES | 74.994 | 2.421 | 0 | 0.102 | 0.0538 | -0.0482 |  | 2 | 0.322
2 | YES | 50.477 | 2.8605 | 0.00070722 | 0.0345 | 0.0538 | 0.0193 |  | 2 | 0.064
2 | YES | 41.036 | 3.277 | 0 | 0.322 | 0.0538 | -0.2682 |  | 2 | -0.147
2 | YES | 30.012 | 3.267 | 0 | 0.064 | 0.0538 | -0.0102 |  | 2 | -0.5165
2 | YES | 19.846 | 3.797 | 0 | -0.147 | 0.0538 | 0.2008 |  | 2 | -0.074
2 | YES | 10.491 | 4.1235 | 0.027577 | -0.5165 | 0.0538 | 0.5703 |  | 3 | 0.015
2 | NO | 5.4895 | 4.543 | 0.10748 | -0.227 | 0.0538 | 0.2808 |  | 3 | 0.015
2 | YES | 3.985 | 4.711 | 0 | -0.074 | 0.0538 | 0.1278 |  | 3 | 0.04
3 | YES | 219.979 | 1.242 | 0 | 0.015 | 0.053905 | 0.038905 |  | 3 | 0.066
3 | YES | 202.99 | 1.244 | 0 | 0.015 | 0.053905 | 0.038905 |  | 3 | 0.079
3 | YES | 177.016 | 1.305 | 0 | 0.04 | 0.053905 | 0.013905 |  | 3 | 0.4415
3 | YES | 150.001 | 1.592 | 0 | 0.066 | 0.053905 | -0.012095 |  | 3 | 0.214
3 | YES | 124.984 | 1.839 | 0 | 0.079 | 0.053905 | -0.025095 |  | 3 | 0.296
3 | YES | 100.496 | 2.2775 | 0.044548 | 0.4415 | 0.053905 | -0.38759 |  | 3 | 0.388
3 | YES | 77.016 | 2.313 | 0 | 0.214 | 0.053905 | -0.16009 |  | 3 | 0.294
3 | YES | 51.988 | 2.689 | 0 | 0.296 | 0.053905 | -0.24209 |  | 3 | 0.452
3 | YES | 41.027 | 3.167 | 0 | 0.388 | 0.053905 | -0.33409 |  | 3 | 0.0040002
3 | YES | 30.003 | 3.523 | 0 | 0.294 | 0.053905 | -0.24009 |  | 3 | -0.0155
3 | YES | 19.991 | 4.188 | 0 | 0.452 | 0.053905 | -0.39809 |  | 3 | 0.233
3 | YES | 10.004 | 4.277 | 0 | 0.0040002 | 0.053905 | 0.049905 |  | 4 | 0.023
3 | YES | 5.4935 | 4.2675 | 0.023335 | -0.0155 | 0.053905 | 0.069406 |  | 4 | 0.023
3 | YES | 2.595 | 4.961 | 0 | 0.233 | 0.053905 | -0.17909 |  | 4 | 0.027
4 | YES | 185.396 | 1.189 | 0.0014143 | 0.023 | 0.05401 | 0.03101 |  | 4 | 0.027
4 | YES | 174.996 | 1.197 | 0 | 0.023 | 0.05401 | 0.03101 |  | 4 | 0.0040001
4 | YES | 150.001 | 1.201 | 0 | 0.027 | 0.05401 | 0.02701 |  | 4 | -0.024
4 | YES | 125.011 | 1.211 | 0 | 0.027 | 0.05401 | 0.02701 |  | 4 | -0.099
4 | YES | 100.497 | 1.521 | 0.0070711 | 0.0040001 | 0.05401 | 0.05001 |  | 4 | 0.017
4 | YES | 76.991 | 1.846 | 0 | -0.024 | 0.05401 | 0.07801 |  | 4 | -0.06
4 | YES | 50.506 | 2.702 | 0.0042427 | -0.099 | 0.05401 | 0.15301 |  | 4 | -0.037
4 | YES | 41.009 | 2.919 | 0 | 0.017 | 0.05401 | 0.03701 |  | 4 | 0.15
4 | YES | 29.983 | 3.096 | 0 | -0.06 | 0.05401 | 0.11401 |  | 4 | 0.2525
4 | YES | 19.989 | 3.443 | 0 | -0.037 | 0.05401 | 0.09101 |  | 4 | 0.134
4 | YES | 10.4915 | 4.093 | 0.0014141 | 0.15 | 0.05401 | -0.09599 |  | 5 | 0.0185
4 | YES | 5.5005 | 4.0935 | 0.00070706 | 0.2525 | 0.05401 | -0.19849 |  | 5 | 0.039
4 | YES | 2.693 | 4.083 | 0 | 0.134 | 0.05401 | -0.07999 |  | 5 | -0.036
5 | YES | 175.173 | 1.2195 | 0.0035355 | 0.0185 | 0.054116 | 0.035615 |  | 5 | -0.201
5 | NO | 0 | 0 | 0 | -1.213 | 0.054116 | 1.2671 |  | 5 | 0.00099993
5 | YES | 124.991 | 1.291 | 0 | 0.039 | 0.054116 | 0.015116 |  | 5 | 0.026
5 | YES | 100.495 | 1.719 | 0.041012 | -0.036 | 0.054116 | 0.090115 |  | 5 | 0.035
5 | YES | 75.006 | 2.327 | 0 | -0.201 | 0.054116 | 0.25512 |  | 5 | 0.05
5 | YES | 49.502 | 2.797 | 0.0070711 | 0.00099993 | 0.054116 | 0.053116 |  | 5 | 0.036
5 | YES | 40.995 | 2.902 | 0 | 0.026 | 0.054116 | 0.028116 |  | 5 | -0.0095
5 | YES | 30.001 | 2.989 | 0 | 0.035 | 0.054116 | 0.019115 |  | 5 | 0.007
5 | YES | 19.988 | 3.048 | 0 | 0.05 | 0.054116 | 0.0041154 |  | 6 | 0
5 | YES | 10.5055 | 3.154 | 0.0014141 | 0.036 | 0.054116 | 0.018116 |  | 6 | -0.023
5 | YES | 5.5105 | 3.2195 | 0.026163 | -0.0095 | 0.054116 | 0.063616 |  | 6 | 0.035
5 | YES | 3.129 | 3.282 | 0 | 0.007 | 0.054116 | 0.047116 |  | 6 | 0.0435
6 | YES | 114 | 1.594 | 0 | 0 | 0.054221 | 0.054221 |  | 6 | 0.138
6 | YES | 100.998 | 1.82 | 0 | -0.023 | 0.054221 | 0.077221 |  | 6 | 0.078
6 | YES | 74.991 | 2.072 | 0 | 0.035 | 0.054221 | 0.019221 |  | 6 | 0.465
6 | YES | 50.4905 | 2.8865 | 0.014849 | 0.0435 | 0.054221 | 0.010721 |  | 6 | 0.2165
6 | YES | 40.992 | 3.314 | 0 | 0.138 | 0.054221 | -0.083779 |  | 6 | 0.082
6 | YES | 29.985 | 3.462 | 0 | 0.078 | 0.054221 | -0.023779 |  | 6 | 0.037
6 | YES | 19.999 | 3.885 | 0 | 0.465 | 0.054221 | -0.41078 |  | 7 | 0.0024999
6 | YES | 10.5035 | 3.9585 | 0.02192 | 0.2165 | 0.054221 | -0.16228 |  | 7 | 0.0205
6 | YES | 5.504 | 4.034 | 0.0070712 | 0.082 | 0.054221 | -0.027779 |  | 7 | -0.071
6 | YES | 2.212 | 4.047 | 0 | 0.037 | 0.054221 | 0.017221 |  | 7 | 0.0595
7 | YES | 123.143 | 1.8955 | 0.0035355 | 0.0024999 | 0.054326 | 0.051826 |  | 7 | 0.151
7 | YES | 100.495 | 2.2795 | 0.0021212 | 0.0205 | 0.054326 | 0.033826 |  | 7 | 0.239
7 | YES | 75.01 | 2.691 | 0 | -0.071 | 0.054326 | 0.12533 |  | 7 | 0.345
7 | YES | 50.5045 | 3.0885 | 0.0035356 | 0.0595 | 0.054326 | -0.0051742 |  | 7 | -0.003
7 | YES | 40.999 | 3.284 | 0 | 0.151 | 0.054326 | -0.096674 |  | 7 | -0.0195
7 | YES | 29.988 | 3.391 | 0 | 0.239 | 0.054326 | -0.18467 |  | 7 | 0.004
7 | YES | 19.998 | 3.692 | 0 | 0.345 | 0.054326 | -0.29067 |  | 8 | -0.017
7 | YES | 10.5015 | 3.731 | 0.0098995 | -0.003 | 0.054326 | 0.057326 |  | 8 | 0.057
7 | YES | 5.494 | 3.7395 | 0.0035356 | -0.0195 | 0.054326 | 0.073826 |  | 8 | -0.023
7 | YES | 2.708 | 3.792 | 0 | 0.004 | 0.054326 | 0.050326 |  | 8 | 0.044
8 | YES | 145.194 | 2.058 | 0.012728 | -0.017 | 0.054431 | 0.071431 |  | 8 | 0.0785
8 | YES | 124.991 | 2.187 | 0 | 0.057 | 0.054431 | -0.002569 |  | 8 | 0.024
8 | YES | 100.995 | 2.678 | 0 | -0.023 | 0.054431 | 0.077431 |  | 8 | 0.034
8 | YES | 75.007 | 2.997 | 0 | 0.044 | 0.054431 | 0.010431 |  | 8 | 0.101
8 | YES | 50.4985 | 3.2705 | 0.028991 | 0.0785 | 0.054431 | -0.024069 |  | 8 | 0.051
8 | YES | 41.004 | 3.369 | 0 | 0.024 | 0.054431 | 0.030431 |  | 8 | 0.04
8 | YES | 29.996 | 3.404 | 0 | 0.034 | 0.054431 | 0.020431 |  | 8 | 0.078
8 | YES | 19.988 | 3.471 | 0 | 0.101 | 0.054431 | -0.046569 |  | 9 | 0
8 | YES | 10.512 | 3.465 | 0.0028284 | 0.051 | 0.054431 | 0.0034308 |  | 9 | -0.0025001
8 | YES | 5.5 | 3.485 | 0.0056569 | 0.04 | 0.054431 | 0.014431 |  | 9 | 0.016
8 | YES | 2.202 | 3.543 | 0 | 0.078 | 0.054431 | -0.023569 |  | 9 | 0.018
9 | YES | 218.883 | 2.662 | 0 | 0 | 0.054536 | 0.054536 |  | 9 | 0.051
9 | YES | 200.508 | 2.6645 | 0.00070706 | -0.0025001 | 0.054536 | 0.057036 |  | 9 | 0.0049999
9 | YES | 175.002 | 2.687 | 0 | 0.016 | 0.054536 | 0.038536 |  | 9 | 0.103
9 | YES | 150.011 | 2.71 | 0 | 0.018 | 0.054536 | 0.036536 |  | 9 | 2.3842e-07
9 | YES | 125.005 | 2.799 | 0 | 0.051 | 0.054536 | 0.0035359 |  | 9 | -0.044
9 | YES | 101.009 | 3.01 | 0 | 0.0049999 | 0.054536 | 0.049536 |  | 9 | 0.093
9 | YES | 77.007 | 3.255 | 0 | 0.103 | 0.054536 | -0.048464 |  | 9 | 0.222
9 | YES | 50.495 | 3.36 | 0.0028284 | 2.3842e-07 | 0.054536 | 0.054536 |  | 9 | -0.0785
9 | YES | 41.01 | 3.376 | 0 | -0.044 | 0.054536 | 0.098536 |  | 10 | 0.026
9 | YES | 29.993 | 3.516 | 0 | 0.093 | 0.054536 | -0.038464 |  | 10 | 0.0315
9 | YES | 22 | 3.699 | 0 | 0.222 | 0.054536 | -0.16746 |  | 10 | 0.029
9 | YES | 11.4935 | 3.7645 | 0.030406 | -0.0785 | 0.054536 | 0.13304 |  | 10 | -0.04
9 | NO | 5.4995 | 4.353 | 0.13435 | 0.318 | 0.054536 | -0.26346 |  | 10 | 0.041
9 | NO | 2.748 | 5.394 | 0 | 1.237 | 0.054536 | -1.1825 |  | 10 | 0.0145
10 | YES | 226.669 | 2.715 | 0.0056569 | 0.026 | 0.054641 | 0.028641 |  | 10 | -0.079
10 | YES | 200.49 | 2.7145 | 0.006364 | 0.0315 | 0.054641 | 0.023141 |  | 10 | 0.034
10 | YES | 176.984 | 2.747 | 0 | 0.029 | 0.054641 | 0.025641 |  | 10 | 0.062
10 | YES | 152.993 | 2.806 | 0 | -0.04 | 0.054641 | 0.094641 |  | 10 | 0.108
10 | YES | 124.993 | 3.062 | 0 | 0.041 | 0.054641 | 0.013641 |  | 10 | 0.106
10 | YES | 100.496 | 3.1095 | 0.006364 | 0.0145 | 0.054641 | 0.040141 |  | 10 | -0.094
10 | YES | 75.007 | 3.205 | 0 | -0.079 | 0.054641 | 0.13364 |  | 10 | -0.1545
10 | YES | 51.509 | 3.39 | 0.0084852 | 0.034 | 0.054641 | 0.020641 |  | 10 | -0.218
10 | YES | 40.997 | 3.438 | 0 | 0.062 | 0.054641 | -0.0073589 |  | 11 | 0.027
10 | YES | 29.982 | 3.506 | 0 | 0.108 | 0.054641 | -0.053359 |  | 11 | 0.038
10 | YES | 19.989 | 3.584 | 0 | 0.106 | 0.054641 | -0.051359 |  | 11 | 0.003
10 | YES | 10.501 | 3.885 | 0.014142 | -0.094 | 0.054641 | 0.14864 |  | 11 | 0.0415
10 | YES | 5.493 | 3.9215 | 0.0021213 | -0.1545 | 0.054641 | 0.20914 |  | 11 | 0.028
10 | YES | 2.71 | 3.974 | 0 | -0.218 | 0.054641 | 0.27264 |  | 11 | 0.0605
11 | YES | 171.333 | 2.808 | 0.0084852 | 0.027 | 0.054746 | 0.027746 |  | 11 | 0.058
11 | YES | 149.986 | 2.942 | 0 | 0.038 | 0.054746 | 0.016746 |  | 11 | 0.007
11 | YES | 125.003 | 2.915 | 0 | 0.003 | 0.054746 | 0.051746 |  | 11 | 0.059
11 | YES | 100.512 | 3.1775 | 0.0049497 | 0.0415 | 0.054746 | 0.013246 |  | 12 | 0.01
11 | YES | 75.002 | 3.298 | 0 | 0.028 | 0.054746 | 0.026746 |  | 12 | 0.022
11 | YES | 50.4985 | 3.3765 | 0.019092 | 0.0605 | 0.054746 | -0.0057539 |  | 12 | 0.013
11 | YES | 41.003 | 3.446 | 0 | 0.058 | 0.054746 | -0.0032539 |  | 12 | 0.0165
11 | YES | 30.014 | 3.526 | 0 | 0.007 | 0.054746 | 0.047746 |  | 12 | 0.03
11 | YES | 19.997 | 3.7 | 0 | 0.059 | 0.054746 | -0.0042538 |  | 12 | 0.0079999
11 | NO | 10.4875 | 4.647 | 0.43558 | 0.451 | 0.054746 | -0.39625 |  | 12 | 0.045
11 | NO | 4.511 | 5.543 | 0.10889 | 0.048 | 0.054746 | 0.0067459 |  | 12 | 0.067
11 | NO | 0 | 0 | 0 | -6.35 | 0.054746 | 6.4047 |  | 12 | 0.025
12 | YES | 317.761 | 2.595 | 0 | 0.01 | 0.054851 | 0.044851 |  | 12 | 0.051
12 | YES | 300.002 | 2.616 | 0 | 0.022 | 0.054851 | 0.032851 |  | 12 | 0.042
12 | YES | 250.007 | 2.78 | 0 | 0.013 | 0.054851 | 0.041851 |  | 12 | 0.045
12 | YES | 200.494 | 2.7875 | 0.0049497 | 0.0165 | 0.054851 | 0.038351 |  | 12 | 0.103
12 | YES | 174.997 | 2.812 | 0 | 0.03 | 0.054851 | 0.024851 |  | 12 | 0.397
12 | YES | 150.003 | 2.81 | 0 | 0.0079999 | 0.054851 | 0.046851 |  | 12 | 0.092
12 | YES | 124.996 | 2.856 | 0 | 0.045 | 0.054851 | 0.0098515 |  | 13 | 0.01
12 | YES | 100.492 | 3.051 | 0.022627 | 0.067 | 0.054851 | -0.012149 |  | 13 | 0.004
12 | YES | 75.015 | 3.075 | 0 | 0.025 | 0.054851 | 0.029851 |  | 13 | 0.011
12 | YES | 50.5025 | 3.338 | 0.011314 | 0.051 | 0.054851 | 0.0038514 |  | 13 | -0.0049999
12 | YES | 40.99 | 3.363 | 0 | 0.042 | 0.054851 | 0.012851 |  | 13 | 0.024
12 | YES | 29.993 | 3.383 | 0 | 0.045 | 0.054851 | 0.0098515 |  | 13 | 0.011
12 | YES | 19.999 | 3.494 | 0 | 0.103 | 0.054851 | -0.048149 |  | 13 | 0.036
12 | YES | 10.504 | 4.305 | 0.039598 | 0.397 | 0.054851 | -0.34215 |  | 13 | 0.047
12 | NO | 5.4955 | 5.924 | 0.30264 | -0.49 | 0.054851 | 0.54485 |  | 13 | 0.031
12 | YES | 3.037 | 6.938 | 0 | 0.092 | 0.054851 | -0.037149 |  | 13 | 0.0145
13 | YES | 371.267 | 2.553 | 0 | 0.01 | 0.054956 | 0.044956 |  | 13 | 0.021
13 | NO | 0 | 0 | 0 | -2.6 | 0.054956 | 2.655 |  | 13 | 0.04
13 | YES | 299.998 | 2.656 | 0 | 0.004 | 0.054956 | 0.050956 |  | 13 | 0.089
13 | YES | 249.994 | 2.684 | 0 | 0.011 | 0.054956 | 0.043956 |  | 13 | 0.199
13 | YES | 200.491 | 2.707 | 0.0042427 | -0.0049999 | 0.054956 | 0.059956 |  | 14 | 0.0050001
13 | YES | 174.999 | 2.768 | 0 | 0.024 | 0.054956 | 0.030956 |  | 14 | 0.01
13 | YES | 149.988 | 2.811 | 0 | 0.011 | 0.054956 | 0.043956 |  | 14 | 0.01
13 | YES | 125.005 | 2.854 | 0 | 0.036 | 0.054956 | 0.018956 |  | 14 | 0.0215
13 | YES | 100.503 | 2.864 | 0.0084852 | 0.047 | 0.054956 | 0.0079565 |  | 14 | 0.039
13 | YES | 74.997 | 3.091 | 0 | 0.031 | 0.054956 | 0.023956 |  | 14 | 0.027
13 | YES | 50.493 | 3.2885 | 0.019092 | 0.0145 | 0.054956 | 0.040456 |  | 14 | 0.011
13 | YES | 41 | 3.338 | 0 | 0.021 | 0.054956 | 0.033956 |  | 14 | 0.0235
13 | YES | 29.983 | 3.424 | 0 | 0.04 | 0.054956 | 0.014956 |  | 14 | 0.028
13 | YES | 19.994 | 3.523 | 0 | 0.089 | 0.054956 | -0.034044 |  | 14 | 0.041
13 | NO | 10.4955 | 4.092 | 0.14991 | 0.0050001 | 0.054956 | 0.049956 |  | 14 | 0.049
13 | NO | 5.4975 | 6.1545 | 0.35143 | 0.1805 | 0.054956 | -0.12554 |  | 14 | 0.071
13 | YES | 2.704 | 6.702 | 0 | 0.199 | 0.054956 | -0.14404 |  | 14 | 0.043
14 | YES | 339.954 | 2.631 | 0 | 0.0050001 | 0.055061 | 0.050061 |  | 14 | 0.1845
14 | YES | 300.006 | 2.675 | 0 | 0.01 | 0.055061 | 0.045061 |  | 14 | -0.01
14 | YES | 249.986 | 2.717 | 0 | 0.01 | 0.055061 | 0.045061 |  | 15 | 0.027
14 | YES | 200.488 | 2.7675 | 0.0049497 | 0.0215 | 0.055061 | 0.033562 |  | 15 | 0.024
14 | YES | 174.996 | 2.822 | 0 | 0.039 | 0.055061 | 0.016061 |  | 15 | 0.023
14 | YES | 150.006 | 2.824 | 0 | 0.027 | 0.055061 | 0.028061 |  | 15 | 0.019
14 | YES | 124.991 | 2.836 | 0 | 0.011 | 0.055061 | 0.044062 |  | 15 | 0.036
14 | YES | 100.494 | 2.8485 | 0.0021213 | 0.0235 | 0.055061 | 0.031562 |  | 15 | 0.039
14 | YES | 74.992 | 2.953 | 0 | 0.028 | 0.055061 | 0.027061 |  | 15 | 0.023
14 | YES | 50.503 | 3.266 | 0.022627 | 0.041 | 0.055061 | 0.014061 |  | 15 | 0.025
14 | YES | 40.974 | 3.26 | 0 | 0.049 | 0.055061 | 0.0060615 |  | 15 | 0.049
14 | YES | 29.983 | 3.3 | 0 | 0.071 | 0.055061 | -0.015938 |  | 15 | 0.074
14 | YES | 19.992 | 3.499 | 0 | 0.043 | 0.055061 | 0.012061 |  | 15 | 0.054
14 | YES | 10.494 | 3.8205 | 0.078489 | 0.1845 | 0.055061 | -0.12944 |  | 15 | 0.038
14 | NO | 5.488 | 4.358 | 0.25314 | -0.96 | 0.055061 | 1.0151 |  | 15 | 0.404
14 | YES | 2.717 | 6.252 | 0 | -0.01 | 0.055061 | 0.065062 |  | 16 | 0.01
15 | YES | 274.923 | 2.702 | 0 | 0.027 | 0.055167 | 0.028167 |  | 16 | 0.0079999
15 | YES | 250 | 2.734 | 0 | 0.024 | 0.055167 | 0.031167 |  | 16 | 0.003
15 | YES | 200.508 | 2.828 | 0.011314 | 0.023 | 0.055167 | 0.032167 |  | 16 | 0.035
15 | YES | 175.005 | 2.845 | 0 | 0.019 | 0.055167 | 0.036166 |  | 16 | 0.0070002
15 | YES | 150.006 | 2.889 | 0 | 0.036 | 0.055167 | 0.019167 |  | 16 | 0.014
15 | YES | 125.01 | 2.855 | 0 | 0.039 | 0.055167 | 0.016167 |  | 16 | 0.038
15 | YES | 100.492 | 2.865 | 0.0070711 | 0.023 | 0.055167 | 0.032167 |  | 16 | 0.033
15 | YES | 74.989 | 2.938 | 0 | 0.025 | 0.055167 | 0.030167 |  | 16 | 0.048
15 | YES | 50.4885 | 3.227 | 0.0042427 | 0.049 | 0.055167 | 0.0061666 |  | 16 | 0.057
15 | YES | 40.989 | 3.329 | 0 | 0.074 | 0.055167 | -0.018833 |  | 16 | 0.06
15 | YES | 30.017 | 3.445 | 0 | 0.054 | 0.055167 | 0.0011667 |  | 16 | 0.08
15 | YES | 20.011 | 3.699 | 0 | 0.038 | 0.055167 | 0.017167 |  | 16 | 0.054
15 | YES | 10.5005 | 5.176 | 0.056569 | 0.404 | 0.055167 | -0.34883 |  | 17 | 0.019
15 | NO | 5.4955 | 5.942 | 0.34365 | -0.179 | 0.055167 | 0.23417 |  | 17 | 0.027
15 | NO | 0 | 0 | 0 | -6.545 | 0.055167 | 6.6002 |  | 17 | 0.042
16 | YES | 318.774 | 2.591 | 0 | 0.01 | 0.055272 | 0.045272 |  | 17 | 0.021
16 | YES | 300.012 | 2.643 | 0 | 0.0079999 | 0.055272 | 0.047272 |  | 17 | 0.0315
16 | YES | 250.023 | 2.719 | 0 | 0.003 | 0.055272 | 0.052272 |  | 17 | 0.07
16 | YES | 200.496 | 2.833 | 0.0014143 | 0.035 | 0.055272 | 0.020272 |  | 17 | 0.075
16 | YES | 174.999 | 2.854 | 0 | 0.0070002 | 0.055272 | 0.048271 |  | 17 | 0.082
16 | YES | 150.009 | 2.877 | 0 | 0.014 | 0.055272 | 0.041271 |  | 17 | 0.038
16 | YES | 125.003 | 2.892 | 0 | 0.038 | 0.055272 | 0.017272 |  | 17 | 0.269
16 | YES | 100.495 | 2.889 | 0.0028284 | 0.033 | 0.055272 | 0.022272 |  | 17 | 0.3655
16 | YES | 75.001 | 2.957 | 0 | 0.048 | 0.055272 | 0.0072716 |  | 18 | 0.024
16 | YES | 50.496 | 3.185 | 0.011314 | 0.057 | 0.055272 | -0.0017282 |  | 18 | 0.042
16 | YES | 41.014 | 3.371 | 0 | 0.06 | 0.055272 | -0.0047282 |  | 18 | 0.025
16 | YES | 29.999 | 3.507 | 0 | 0.08 | 0.055272 | -0.024728 |  | 18 | 0.0425
16 | YES | 19.987 | 3.592 | 0 | 0.054 | 0.055272 | 0.0012718 |  | 18 | 0.04
16 | NO | 10.498 | 4.1565 | 0.1393 | 0.3915 | 0.055272 | -0.33623 |  | 18 | 0.0805
16 | NO | 5.502 | 6.281 | 0.31395 | -0.115 | 0.055272 | 0.17027 |  | 18 | 0.084
16 | NO | 0 | 0 | 0 | -6.5 | 0.055272 | 6.5553 |  | 18 | 0.098
17 | YES | 188.992 | 2.788 | 0 | 0.019 | 0.055377 | 0.036377 |  | 18 | 0.02
17 | YES | 174.999 | 2.802 | 0 | 0.027 | 0.055377 | 0.028377 |  | 19 | 0.0049999
17 | YES | 149.997 | 2.891 | 0 | 0.042 | 0.055377 | 0.013377 |  | 19 | 0.042
17 | YES | 125.004 | 2.897 | 0 | 0.021 | 0.055377 | 0.034377 |  | 19 | 0.025
17 | YES | 100.494 | 2.8815 | 0.0063638 | 0.0315 | 0.055377 | 0.023877 |  | 19 | 0.0285
17 | YES | 74.992 | 2.944 | 0 | 0.07 | 0.055377 | -0.014623 |  | 19 | 0.022
17 | YES | 50.5005 | 3.286 | 0.012728 | 0.075 | 0.055377 | -0.019623 |  | 19 | 0.022
17 | YES | 41.007 | 3.389 | 0 | 0.082 | 0.055377 | -0.026623 |  | 19 | 0.034
17 | YES | 30.001 | 3.442 | 0 | 0.038 | 0.055377 | 0.017377 |  | 19 | 0.037
17 | YES | 19.991 | 3.669 | 0 | 0.269 | 0.055377 | -0.21362 |  | 19 | 0.062
17 | YES | 10.492 | 4.7695 | 0.04879 | 0.3655 | 0.055377 | -0.31012 |  | 19 | 0.0645
17 | NO | 5.5885 | 5.6055 | 0.4278 | -0.3015 | 0.055377 | 0.35688 |  | 19 | 0.125
17 | NO | 0 | 0 | 0 | -6.728 | 0.055377 | 6.7834 |  | 19 | 0.062
18 | YES | 162.759 | 2.497 | 0 | 0.024 | 0.055482 | 0.031482 |  | 19 | 0.287
18 | YES | 149.998 | 2.527 | 0 | 0.042 | 0.055482 | 0.013482 |  | 19 | 0.019
18 | YES | 124.998 | 2.829 | 0 | 0.025 | 0.055482 | 0.030482 |  | 20 | 0.013
18 | YES | 100.49 | 2.8525 | 0.00070706 | 0.0425 | 0.055482 | 0.012982 |  | 20 | 0.023
18 | YES | 74.983 | 3.019 | 0 | 0.04 | 0.055482 | 0.015482 |  | 20 | 0.024
18 | YES | 50.507 | 3.3895 | 0.016264 | 0.0805 | 0.055482 | -0.025018 |  | 20 | 0.0115
18 | YES | 41.01 | 3.516 | 0 | 0.084 | 0.055482 | -0.028518 |  | 20 | 0.027
18 | YES | 29.993 | 3.606 | 0 | 0.098 | 0.055482 | -0.042518 |  | 20 | 0.032
18 | YES | 19.996 | 4.116 | 0 | 0.02 | 0.055482 | 0.035482 |  | 20 | 0.034
18 | NO | 10.501 | 5.843 | 0.14708 | 0.071 | 0.055482 | -0.015518 |  | 20 | 0.118
18 | NO | 4.4875 | 7.1605 | 0.16476 | -0.3025 | 0.055482 | 0.35798 |  | 20 | 0.061
18 | NO | 0 | 0 | 0 | -7.117 | 0.055482 | 7.1725 |  | 20 | 0.097
19 | YES | 351.925 | 2.715 | 0 | 0.0049999 | 0.055587 | 0.050587 |  | 20 | 0.043
19 | YES | 300.001 | 2.753 | 0 | 0.042 | 0.055587 | 0.013587 |  | 20 | 0.195
19 | YES | 249.991 | 2.787 | 0 | 0.025 | 0.055587 | 0.030587 |  | 20 | 0.104
19 | YES | 200.487 | 2.8455 | 0.0049497 | 0.0285 | 0.055587 | 0.027087 |  | 20 | 0.043
19 | YES | 174.995 | 2.86 | 0 | 0.022 | 0.055587 | 0.033587 |  | 21 | 0.022
19 | YES | 150.001 | 2.89 | 0 | 0.022 | 0.055587 | 0.033587 |  | 21 | 0.022
19 | YES | 124.987 | 2.901 | 0 | 0.034 | 0.055587 | 0.021587 |  | 21 | -0.071
19 | YES | 100.506 | 2.875 | 0.0084853 | 0.037 | 0.055587 | 0.018587 |  | 21 | 0.127
19 | YES | 75.002 | 2.93 | 0 | 0.062 | 0.055587 | -0.006413 |  | 21 | 0.037
19 | YES | 50.5015 | 3.2185 | 0.0091924 | 0.0645 | 0.055587 | -0.0089129 |  | 21 | 0.015
19 | YES | 41.001 | 3.501 | 0 | 0.125 | 0.055587 | -0.069413 |  | 21 | 0.082
19 | YES | 30.003 | 3.598 | 0 | 0.062 | 0.055587 | -0.006413 |  | 21 | 0.246
19 | YES | 19.99 | 3.945 | 0 | 0.287 | 0.055587 | -0.23141 |  | 21 | 0.3975
19 | NO | 10.498 | 4.91 | 0.16122 | 0.565 | 0.055587 | -0.50941 |  | 21 | 0.1085
19 | NO | 5.504 | 6.632 | 0.11738 | 0.666 | 0.055587 | -0.61041 |  | 21 | 0.086
19 | YES | 2.655 | 6.771 | 0 | 0.019 | 0.055587 | 0.036587
20 | YES | 312.258 | 2.748 | 0 | 0.013 | 0.055692 | 0.042692
20 | YES | 299.998 | 2.767 | 0 | 0.023 | 0.055692 | 0.032692
20 | YES | 250.002 | 2.822 | 0 | 0.024 | 0.055692 | 0.031692
20 | YES | 200.509 | 2.8765 | 0.00070722 | 0.0115 | 0.055692 | 0.044192
20 | YES | 175.009 | 2.896 | 0 | 0.027 | 0.055692 | 0.028692
20 | YES | 150.004 | 2.908 | 0 | 0.032 | 0.055692 | 0.023692
20 | YES | 124.986 | 2.831 | 0 | 0.034 | 0.055692 | 0.021692
20 | YES | 100.499 | 3.131 | 0.06364 | 0.118 | 0.055692 | -0.062308
20 | YES | 75.003 | 3.535 | 0 | 0.061 | 0.055692 | -0.0053081
20 | YES | 50.4985 | 3.869 | 0.0098995 | 0.097 | 0.055692 | -0.041308
20 | YES | 40.982 | 3.85 | 0 | 0.043 | 0.055692 | 0.012692
20 | YES | 29.991 | 4.062 | 0 | 0.195 | 0.055692 | -0.13931
20 | YES | 19.999 | 4.126 | 0 | 0.104 | 0.055692 | -0.048308
20 | NO | 10.503 | 5.2705 | 0.39386 | 0.9575 | 0.055692 | -0.90181
20 | NO | 5.492 | 6.44 | 0.15132 | 0.058 | 0.055692 | -0.0023081
20 | YES | 2.709 | 6.735 | 0 | 0.043 | 0.055692 | 0.012692
21 | YES | 151.799 | 2.835 | 0 | 0.022 | 0.055797 | 0.033797
21 | YES | 125.011 | 2.858 | 0 | 0.022 | 0.055797 | 0.033797
21 | YES | 100.985 | 2.892 | 0 | -0.071 | 0.055797 | 0.1268
21 | YES | 75.004 | 3.32 | 0 | 0.127 | 0.055797 | -0.071203
21 | YES | 50.494 | 3.788 | 0.0028284 | 0.037 | 0.055797 | 0.018797
21 | YES | 40.991 | 3.805 | 0 | 0.015 | 0.055797 | 0.040797
21 | YES | 29.993 | 3.911 | 0 | 0.082 | 0.055797 | -0.026203
21 | YES | 20 | 4.229 | 0 | 0.246 | 0.055797 | -0.1902
21 | YES | 10.5025 | 4.6065 | 0.012021 | 0.3975 | 0.055797 | -0.3417
21 | YES | 5.496 | 4.6605 | 0.017678 | 0.1085 | 0.055797 | -0.052703
21 | YES | 2.506 | 4.74 | 0 | 0.086 | 0.055797 | -0.030203
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. OXY_SBE
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(OXY_SBE - OXY_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (OXY_SBE - OXY_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0051237 |  | X | CALC Y
m= | 0.0097308 | 0.02292 |  | Std Dev of Resid |  | 0.11687 |  | 2 | 0.040607
b= | 0.021146 | 0.071393 |  | Resid Sum of Squares |  | 0.47803 |  | 4 | 0.060069
# of Values Extracted = |  |  | 42
# of Values Used in Fit = |  |  | 37
File | Use | Pressure | OXY_SBE | OXY_SBE | Diff (OXY_SBE - OXY_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
2 | YES | 252.453 | 1.2085 | 0.0035355 | 0.0084999 | 0.040607 | 0.032107 |  | 2 | 0.0084999
2 | YES | 249.989 | 1.213 | 0 | 0.01 | 0.040607 | 0.030607 |  | 2 | 0.01
2 | YES | 200.506 | 1.2595 | 0.0035355 | 0.0055001 | 0.040607 | 0.035107 |  | 2 | 0.0055001
2 | YES | 175.221 | 1.3 | 0 | -0.018 | 0.040607 | 0.058607 |  | 2 | -0.018
2 | YES | 149.987 | 1.55 | 0 | -0.104 | 0.040607 | 0.14461 |  | 2 | -0.104
2 | YES | 125.003 | 2.066 | 0 | -0.068 | 0.040607 | 0.10861 |  | 2 | -0.068
2 | YES | 101.002 | 2.145 | 0 | 0.049 | 0.040607 | -0.0083929 |  | 2 | 0.049
2 | YES | 74.994 | 2.421 | 0 | 0.102 | 0.040607 | -0.061393 |  | 2 | 0.102
2 | YES | 50.477 | 2.8605 | 0.00070722 | 0.0345 | 0.040607 | 0.0061072 |  | 2 | 0.0345
2 | YES | 41.036 | 3.277 | 0 | 0.322 | 0.040607 | -0.28139 |  | 2 | 0.322
2 | YES | 30.012 | 3.267 | 0 | 0.064 | 0.040607 | -0.023393 |  | 2 | 0.064
2 | YES | 19.846 | 3.797 | 0 | -0.147 | 0.040607 | 0.18761 |  | 2 | -0.147
2 | NO | 10.491 | 4.1235 | 0.027577 | -0.5165 | 0.040607 | 0.55711 |  | 2 | -0.074
2 | NO | 5.4895 | 4.543 | 0.10748 | -0.227 | 0.040607 | 0.26761 |  | 3 | 0.015
2 | YES | 3.985 | 4.711 | 0 | -0.074 | 0.040607 | 0.11461 |  | 3 | 0.015
3 | YES | 219.979 | 1.242 | 0 | 0.015 | 0.050338 | 0.035338 |  | 3 | 0.04
3 | YES | 202.99 | 1.244 | 0 | 0.015 | 0.050338 | 0.035338 |  | 3 | 0.066
3 | YES | 177.016 | 1.305 | 0 | 0.04 | 0.050338 | 0.010338 |  | 3 | 0.079
3 | YES | 150.001 | 1.592 | 0 | 0.066 | 0.050338 | -0.015662 |  | 3 | 0.214
3 | YES | 124.984 | 1.839 | 0 | 0.079 | 0.050338 | -0.028662 |  | 3 | 0.296
3 | NO | 100.496 | 2.2775 | 0.044548 | 0.4415 | 0.050338 | -0.39116 |  | 3 | 0.294
3 | YES | 77.016 | 2.313 | 0 | 0.214 | 0.050338 | -0.16366 |  | 3 | 0.0040002
3 | YES | 51.988 | 2.689 | 0 | 0.296 | 0.050338 | -0.24566 |  | 3 | -0.0155
3 | NO | 41.027 | 3.167 | 0 | 0.388 | 0.050338 | -0.33766 |  | 3 | 0.233
3 | YES | 30.003 | 3.523 | 0 | 0.294 | 0.050338 | -0.24366 |  | 4 | 0.023
3 | NO | 19.991 | 4.188 | 0 | 0.452 | 0.050338 | -0.40166 |  | 4 | 0.023
3 | YES | 10.004 | 4.277 | 0 | 0.0040002 | 0.050338 | 0.046338 |  | 4 | 0.027
3 | YES | 5.4935 | 4.2675 | 0.023335 | -0.0155 | 0.050338 | 0.065838 |  | 4 | 0.027
3 | YES | 2.595 | 4.961 | 0 | 0.233 | 0.050338 | -0.18266 |  | 4 | 0.0040001
4 | YES | 185.396 | 1.189 | 0.0014143 | 0.023 | 0.060069 | 0.037069 |  | 4 | -0.024
4 | YES | 174.996 | 1.197 | 0 | 0.023 | 0.060069 | 0.037069 |  | 4 | -0.099
4 | YES | 150.001 | 1.201 | 0 | 0.027 | 0.060069 | 0.033069 |  | 4 | 0.017
4 | YES | 125.011 | 1.211 | 0 | 0.027 | 0.060069 | 0.033069 |  | 4 | -0.06
4 | YES | 100.497 | 1.521 | 0.0070711 | 0.0040001 | 0.060069 | 0.056069 |  | 4 | -0.037
4 | YES | 76.991 | 1.846 | 0 | -0.024 | 0.060069 | 0.084069 |  | 4 | 0.15
4 | YES | 50.506 | 2.702 | 0.0042427 | -0.099 | 0.060069 | 0.15907 |  | 4 | 0.2525
4 | YES | 41.009 | 2.919 | 0 | 0.017 | 0.060069 | 0.043069 |  | 4 | 0.134
4 | YES | 29.983 | 3.096 | 0 | -0.06 | 0.060069 | 0.12007
4 | YES | 19.989 | 3.443 | 0 | -0.037 | 0.060069 | 0.097069
4 | YES | 10.4915 | 4.093 | 0.0014141 | 0.15 | 0.060069 | -0.089931
4 | YES | 5.5005 | 4.0935 | 0.00070706 | 0.2525 | 0.060069 | -0.19243
4 | YES | 2.693 | 4.083 | 0 | 0.134 | 0.060069 | -0.073931 |  |  |  | 0.050338 | 0.11553
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. OXY_SBE
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE